## Slide 2: WP 13 : Beam Diagnostics Conceptual Specifications
Rhodri Jones (BE/BI)

## Slide 3: Introduction
Cryogenic BLMs & Radiation Hard Electronics
  Cryogenic BLMs
  Radiation hard electronics
Fast WireScanners
Insertion Region BPMs
  Cold directional couplers
  Tungsten shielded cold directional couplers
  Warm directional couplers
  High precision electronics for insertion region BPMs
Luminosity Monitors
Diagnostics for Crab Cavities
Upgrade to Synchrotron Light Monitors
  Upgrade to existing monitor
  New light source
  Halo diagnostics
Beam Gas Vertex Detector
  Final Implementation
Long-Range Beam-Beam Compensator
  Prototype
  Final Implementation
WP 13 Tasks

## Slide 4: Cryogenic BLMs
HL-LHC Cryogenic BLMs (BLMC) – Baseline (64061)
  For present BLMs additional signal from accidental losses masked by collision debris
  Proposal to put radiation detectors inside the coldmass of triplet magnets
  Dose measured then correspond to dose deposited in the coils
Performance Objectives
  Ability to detect secondary particle showers
  Work in cryogenic environment, high magnetic fields (4T) & withstand up to 20 MGy
Assumptions
  Detectors based on CVD diamond or p+-n-n+ silicon wafers (detector area ~25mm2)
  Present a negligible heat load for the cryogenic system of the inner triplet magnets
  Longitudinal layout to be determined by FLUKA/GEANT studies
Location
  Up to 6 detectors installed inside the cold mass of each triplet quadrupole
Technical Parameters
  1 semi-rigid coaxial cable per detector provides HV (up to 1000V) & extracts signal
Installation
  Monitors & cables mounted during assembly of the magnet cold mass in the cryostat
  Baseline procurement of 100 detectors (96 installed & 4 spares).
  Equipping 11T dipole & all the spare triplet magnet assemblies requires total of 150 detectors
Reliability & Availability
  If monitors break external BLM system still provides protection against damage (not quench)

## Slide 5: Fast Wirescanners
HL-LHC Fast Wirescanners (BWSF) – Baseline (64062)
  Installation of fast wirescanners capable of scanning at a speed of up to 20ms-1
Performance Objectives
  Currently system has max speed of 1 ms-1
  Limits total intensity scanned at injection to ~2.7×1013 (~200 nominal bunches).
  20ms-1 allows average beam size measurement with full injected physics beam
  Goals:
    Accuracy of 5 mm on beam width, i.e. error of 5% for nominal emittance beams
    Large dynamic range acquisition system to improve reliability in operation
    Increase MTBF compared to existing system (no bellows)
Assumptions
  Scanners based on prototype currently being tested in the SPS
  Acquisition system based on diamond detector located downstream
  Design modifications adequate for machine impedance & RF heating
Location
  Beam 1 : 2 H & 2 V scanners located next to the existing scanners in 5R4
  Beam 2 : 2 H & 2 V scanners located next to the existing scanners in 5L4
Technical Parameters & Installation
  Rotative scanner of length 500mm adapted to 80mm warm beam pipe aperture

## Slide 6: Tungsten Shielded Cryogenic BPMs
HL-LHC Tungsten Shielded Cryogenic BPMs (BPMSQT) – Baseline (64063)
  Measure position of both beams in both planes
  Essential for maintaining stable orbit at IP
  Could be used for continuous luminosity optimisation
Performance Objectives
  Beam position resolution for each beam of 1mm
  Fill to fill reproducibility of 10mm
Assumptions
  Stripline BPM, tungsten shielding required, 45° stripline orientation & one design fits all
Location
  In interconnect between Q2a & Q2b, Q2b and Q3a, Q3a and CP
Technical Parameters
  Aperture adapted to beam screen
  Minimum length 220mm
Constraints
  Longitudinal location from IP (1.87 + N x 3.743)m where N = integer (where possible)
Installation
  12 installed + 3 spares
  8 cryogenic coaxial cables per BPM with feedthrough location on cryostat to be determined
  Additional Rack for electronics in UA/UJ each side of LSS with fibre connections to surface

## Slide 7: Cryogenic BPMs
HL-LHC Cryogenic BPMs (BPMSQ) – Baseline (64063)
  Measure position of both beams in both planes
  Essential for maintaining stable orbit at IP
  Could be used for continuous luminosity optimisation
Performance Objectives
  Beam position resolution for each beam of 1mm
  Fill to fill reproducibility of 10mm
Assumptions
  Stripline BPMs, no tungsten shielding required, 90° stripline orientation & one design fits all
Location
  In front of the Q2a
  Before and after the D1 magnet
Technical Parameters
  Aperture adapted to beam screen
  Minimum length 220mm
Constraints
  Longitudinal location from IP (1.87 + N x 3.743)m where N = integer (where possible)
Installation
  12 installed + 3 spares
  8 cryogenic coaxial cables per BPM with feedthrough location on cryostat to be determined

## Slide 8: Warm BPMs
HL-LHC Warm BPMs (BPMSQW) – Baseline (64063)
  Measure position of both beams in both planes
  Essential for maintaining stable orbit at IP
  Could be used for continuous luminosity optimisation
Performance Objectives
  Beam position resolution for each beam of 1mm
  Fill to fill reproducibility of 10mm
Assumptions
  Stripline BPMs, no tungsten shielding required, 90° stripline orientation & one design fits all
Location
  In front of the Q1a
Technical Parameters
  Aperture adapted to beam screen
  Minimum length 285mm
Constraints
  Longitudinal location from IP (1.87 + N x 3.743)m where N = integer (where possible)
Installation
  4 installed + 2 spares
  8 semi-rigid coaxial cables per BPM connecting to 8 standard coaxial cables

## Slide 9: Upgrade of Existing Synchrotron Light Monitors
HL-LHC Synchrotron Light Monitor (BSR) – Baseline (64066)
  Construction of an evacuated optical light path
    from BSRTM.5L4.B2 to UA43
    from BSRTM.5R4.B1 to UA47
  Construction of an associated optical hutch in UA43 and UA47
Performance Objectives
  Installation of highly expensive & non radiation resistant optical detection equipment
    Fast cameras for bunch to bunch diagnostics
    High dynamic range cameras for halo diagnostics
    Streak cameras for intra-bunch diagnostics
  Must be located in a low radiation area
  Easier set-up and alignment
Technical & Operational Parameters
  Optical light path consisting of:
    100mm diameter evacuated pipe
    Remotely controlled steering mirrors
    Distance as short as possible between extraction mirror and optical hutch
  Optical hutch of at least 10m2 with related services
  Possible need for additional holes between RA43/47 and UA43/47

## Slide 10: New Light Extraction System for Synchrotron Light Diagnostics
HL-LHC Synchrotron Light Monitor (BSR) – Baseline (64066)
  Construction of a new light extraction system for synchrotron light diagnostics in LSS4
  Extract light from D4 with in-vacuum mirror located between D4 and D3 on each incoming beam
  Extracted light sent via optical light path from tunnel to new optical hutch in UA
Performance Objectives
  Optical table for existing synchrotron light telescope is full to capacity
    Need for additional synchrotron light source
  Fulfil HL-LHC requirements: beam halo, intra-bunch measurements & interlocked abort gap
Assumptions
  Measurements only required above 2TeV using edge or central radiation from the D4
  The possibility to integrate a light extraction system in this region
Technical & Operational Parameters
  Low impedance extraction mirror design to minimise RF heating
  Optical light path consisting of:
    100mm diameter evacuated pipe
    Remotely controlled steering mirrors
    Distance as short as possible between extraction mirror and optical hutch
  Optical hutch of at least 10m2 with related services
  Possible need for additional holes between RA43/47 and UA43/47

## Slide 11: Beam Gas Vertex Detector
HL-LHC Beam Gas Vertex Detector (BGV) – Baseline (64067)
  One BGV detector per beam for non-invasive beam size measurements
  Foreseen for installation between Q6 and Q7 in LSS4 left and right
  Space reservation : LHC-BGV-EC-0001 (EDMS 1282994)
Performance Objectives
  Bunch width measurements with a 5 % resolution within 1 minute
  Beam width measurements with an absolute accuracy of 2 % within 1 minute
  Would allow meaningful measurements during the energy ramp
  Could measure beam halo, longitudinal density & longitudinal beam profile
Assumptions
  Based on prototype currently being installed
    Optimised design for HL-LHC
    New detectors & acquisition electronics
  Final HL-LHC optics allows installation within reserved zone in LSS4

## Slide 12: Beam Gas Vertex Detector
Technical & Operational Parameters
  Operates with neon gas injection at pressure of 6×10–8 mbar
  Pressure outside tank area should remain below 6×10–9 mbar
  Aperture at detector should be as small as possible to improve accuracy
  Total length ~8m per BGV station
  ~100 cables and 2 additional racks per BGV station
Further Assumptions
    Design adequate for machine impedance & RF heating
    Gas pressure not severely detrimental to radiation dose in the neighbouring arc sections