## Slide 1: Measuring the magnetic axis of quadrupoles by a stretched wire
Domenico Caiazza
4.2.2015

## Slide 2: Outline
My role in PACMAN
Magnetic measurements and measurement problem
Classical stretched-wire method
Vibrating-wire method
Preliminary results and comparison
Criticalities about the vibrating wire system
Conclusions

## Slide 3
WP4 Beam Instrumentation
M. Wendt
WP3 Precision mech. & stabilization
M. Modena
WP2 Magnetic Measurements
S. Russenschuck
WP1 Metrology & Alignment
H. Mainaud Durand
Domenico
Claude
Solomon
Vasileios
Giordana
Iordan
Peter
David
Silvia
Natalia
“Stretched-wire systems for the magnetic measurement of small-aperture magnets”

## Slide 4: CLIC - Compact Linear Collider
Metrological challenges imposed by CLIC
  Small beam size and impact section
    submicron range
  High-precision in focusing the beam
    500 nm horizontal and 5 nm vertical normalized beam emittance
  Tight specifications on components’ alignment
    14-17 μm error budget over 200 m
A two beams, 50 km, e+/e- collider
QUAD
BPM
ACCEL. Struct.

## Slide 5: Magnetic measurements
4000 quadrupoles in the main linac
Magnetic measurements for PACMAN
Main beam quad
Final focus quad
Magnetic Measurements Section
Stretched wires
Rotating coils

## Slide 6: Measurement problem
Upgrade of LHC
(Linear Accelerator 4)
Future accelerators
(Compact Linear Collider)
Existing large accelerators
(Large Hadron Collider)
Ø 50 mm
Ø 20 mm
Ø 8 mm
2008
2015
2020
2.5 x
2.5 x
Limited access for Hall probe
The smaller the radius coils,
the higher the uncertainty

## Slide 7: Wire-based transducers
A single stretched wire
Magnetic measurements
field strength and direction
field harmonics
magnetic axis
longitudinal field profile

## Slide 8: Solution 1: Stretched-wire method
Metrological reference for integrated field strength, direction and magnetic axis of LHC magnets
Magnetic flux measurement
J. Di Marco et al., “Field alignment of quadrupole magnets for the LHC interaction Regions”. IEEE Transactions on Applied Superconductivity, 2000.

## Slide 9: Solution 2: Vibrating-wire method
Plucked string
Feeding the wire by alternating current
(Lorentz Force)
Wire in magnetic field
Measure wire motion
amplitude X and Y components
Relate motion to
magnetic field properties
A. Temnykh. “Vibrating wire field-measuring technique”. Nuclear Instruments and Methods in Physics Research, 1997.

## Slide 10: Mathematical model
Linearity
Plane motion
Uniform and constant tension
Small deflections
Constant length
Uniform mass distribution

## Slide 11: Measurement system design
Working frequencies for max sensitivity
Resonances
Driving current frequency tuning

## Slide 12: Vibrating-wire axis measurement
A zero-finding problem
QUAD
A
B
QUAD
A
B
Center finding (horizontal and vertical)
  First resonance
  Co-directional scanning
Tilt finding (pitch and yaw)
  Second resonance
  Counter-directional scanning
P. Arpaia, C. Petrone, S. Russenschuck, L. Walckiers, “Vibrating-wire measurement method for centering and alignment of solenoids”, IOP Journal of Instrumentation, Vol. 8, Nov., 2013.
Z. Wolf. “A Vibrating Wire System For Quadrupole Fiducialization ”, Tech. rep. LCLS-TN-05-1. SLAC National Accelerator Laboratory, Menlo Park, California, USA, 2005.
Vibration amplitude

## Slide 13: Experimental setup
P. Arpaia, M. Buzio, J. G. Perez, C. Petrone, S. Russenschuck, L. Walckiers.
“Measuring field multipoles in accelerator magnets with small-apertures by an oscillating wire moved on a circular trajectory”, IOP JINST - Journal of Instrumentation, 2012.
Sensors: phototransistor SharpTM GP1S094HCZ0F
Current generator: Keithley® 6351
Common marble support for magnets and stages
System architecture
Configuration
Wire alloy | - | Cu-Be
Wire mass density | ρ | 1.1·10-4 Kg·m-1
Wire length | L | 1870 mm
Reference magnet
LHC-LI-QS

## Slide 14: Comparison stretched-vibrating: Preliminary results
Precision
  Compatibility
  30 times better for the vibrating wire (3-4 μm)
Sensitivity
  Vibrating wire is suitable at low integral gradient (< 0.2 T)
Background fields
  Stretched wire not sensitive
  Compensation needed for vibrating wire
Method 1: classical stretched wire center
x0 [mm] | sx [mm] | y0 [mm] | sy [mm]
-0.270 | 0.136 | 0.123 | 0.098
x0 [mm] | sx [mm] | y0 [mm] | sy [mm]
-0.272 | 0.003 | 0.086 | 0.004
Method 2: vibrating wire center

## Slide 15: Background field compensation (1/2)
Solution 1: displace the quadrupole at L/4
A
B
Z. Wolf. “A Vibrating Wire System For Quadrupole Fiducialization ”, Tech. rep. LCLS-TN-05-1. SLAC National Accelerator Laboratory, Menlo Park, California, USA, 2005
Solution 2: measure at different gradients
A
B
Works if there is not a constant dipole in the magnet (not depending on Im)

## Slide 16
Magnetic center as a function of the magnet gradient
Hyperbolic model
Actual center
Background field effect
Fit to the hyperbolic model
Background field compensation (2/2)

## Slide 17: Reference frame
Fiducial marker localization

## Slide 18: DBQ axis measurement
Magnetic axis measurement + fiducial markers localization
  In collaboration with EN-MEF-SU (Large Scale Metrology Section)
Both center and tilt were measured by the vibrating wire
Axis determination with +3 μm horizontal and +4 μm vertical precision
Chart: Magnet current
| 82.5 | 40 | 20 | 16 | 12 | 8 | 4
x [μm] | 0 | 1 | 9.5 | 11.8 | 15.7 | 23.5 | 47
y [μm] | 0 | -0.5 | -1 | -2.2 | -4.5 | -9 | -21
Magnetic center as a function of the magnet current
CLIC DBQ (12 T integrated gradient) on the fiducialization bench with vibrating wire system
M. Duquenne et al., “Determination of the magnetic axis of a CLIC drive beam quadrupole with respect to external alignment targets using a combination of wps, cmm and laser tracker measurements”. Proceedings of IPAC2014, Dresden, 2014.

## Slide 19
Background fields & Plane motion
No magnet on the measurement station
Background fields
  2% alteration of first harmonic
Fringe field from equipment (tensioning system)
  High-order modes amplified
Elliptically polarized trajectory
  in resonance condition
Predicted in:
J. A. Elliott. “Intrinsic nonlinear effects in vibrating strings”, American Journal of Physics, 1989

## Slide 20
Around resonance
  Non-constant oscillation amplitude!!!
  Effect depending on the excitation frequency: minimal in resonance condition (5%)
Steady-state modulation
Possible reasons
  Non constant length and/or tension
  Non ideal clamping (friction on the supports)
  Excluded: coupling with ground vibrations
Side band components in the frequency spectrum

## Slide 21: Conclusions
Two methods for magnetic axis measurement
Preliminary test on a reference quadrupole
  Comparison of precision and sensitivity
  Influence of background field
Outlook
  Uncertainty analysis of the vibrating-wire system
  Design and commissioning of the PACMAN system
  Measurement on the CLIC MBQ magnet
  Comparison with PCB rotating coil

## Slide 22: Thank you for your attention
Measuring the magnetic axis of quadrupoles by a stretched wire
PACMAN Workshop 2015

## Slide 23
SPARES

## Slide 24: Method 1: preliminary results
Horizontal offset with respect to the magnetic center
Measured gradient as a function of magnet current
+300 μm uncertainty for low-gradient (< 0.2 T/m·m)
Average horizontal and vertical center offset
m
x0 [mm] | St. Dev. [mm] | y0 [mm] | St. Dev. [mm]
-0.270 | 0.136 | 0.123 | 0.098

## Slide 25: Method 2: preliminary results
Average horizontal and vertical center offset
Linear regression of wire oscillation amplitude
x0 [mm] | St. Dev. [mm] | y0 [mm] | St. Dev. [μm]
-0.272 | 0.003 | 0.086 | 0.004

## Slide 26: Results
Reconstruction error 3% of the field peak
Repeatability 2%
  RMS difference
Bandwidth limitation
Uncertainty sources

## Slide 27: Overtones
overtones
main tone
Overtone amplitude from 2% to 7% of the main tone
  Depending on system configuration
Overtones not contained in the current excitation signal
Nonlinearity!

## Slide 28: Optical sensor characterization
Keyence ® CCD sensors
  Wide linear domain (6 mm)
  Worse SNR
Sharp® phototransistors (currently employed)
  Linear domain 40 μm
  Low price
  High sensitivity
Total Harmonic Distortion
  No harmonic distortion by Sharp®
Overtones are not artifacts from nonlinear sensor response