## Slide 1
E-TCT measurements with laser beam directed parallel to strips
Igor Mandić1, Vladimir Cindro1, Andrej Gorišek1, Gregor Kramberger1, Marko Mikuž1,2, Marko Zavrtanik1
1Jožef Stefan Institute, Ljubljana, Slovenia
2 Faculty of Mathematics and Physics, University of Ljubljana, Slovenia

## Slide 2
Introduction
in usual E-TCT measurements focused laser beam is directed to the side of the
detector perpendicular to the direction of strips
- IR laser (λ = 1064 nm), attenuation length in Si ~1 mm
- charge is generated below many strips
- carriers drifting to the neighbor strips also induce current on the readout electrode
 effective weighting field modified:
  instead of peaked weighting field of a strip detector the effective weighting field
is approximately constant Ew = 1/D as in pad detectors
  good for measurements of velocity profile:
  charge collection profile of irradiated detector is not as in detector application
if laser beam is directed parallel to strips all carriers (in ideal case) generated along laser
beam drift in the same electric and weighting field -> true charge collection profiles could be
measured
  in reality laser beam divergence smears the measurement
in this presentation some aspects of this measurement technique will be presented

## Slide 3
Weighting field
strongly peaked at the readout strip
Electric field
Simulation
KDetSim by G. Kramberger
described in:
G. Kramberger et al., IEEE Trans.Nucl.Sci., Vol 49, (2002), 1717.
G. Kramberger et al., NIM A 457, pp-550-557,2001.
G. Kramberger et al., 2014 JINST 9 P10016
P-type strip detector, pitch 75 um, strip width 20 um
Neff = 4.76e12 cm-3 (Vfd ~ 350 V)
(approx. Hamamatsu ATLAS12)
Bias = 400 V

## Slide 4
Laser
Readout electrode
…
+
+
+
+
current pulse on readout strip in E-TCT is induced also by charge drifting under neighbor strips
 measured pulse is the sum of all contributions
Simulation
standard E-TCT

## Slide 5
Readout strip
Laser
pulses different because of carriers drifting to other strips  reason for effective Ew = 1/D in E-TCT
charge (integral of the pulse) same for blue and red pulse before irradiation
 in irradiated detector carriers do not drift over whole weighting field because of trapping
 charge is not the same for the two cases
Blue: all charge
Red: only charge under strip
Simulation

## Slide 6
laser beam in the direction of strips:
 “real” strip weighting field
 “real” charge collection profiles in irradiated strip detectors could be measured
Problem: laser beam divergence
Simulation
scan this area with laser beam :
Readout strip
Detector depth
Laser beam spot
Beam direction

## Slide 7
- Gaussian beam spot size:
 initial beam spot radius w0 = 4 µm  at z = 2 mm radius already w = 48 µm
 absorption length ~ 1 mm, contribution gets smaller as beam spot grows
Gaussian beam

## Slide 8
Simulation
compare collected charge - time integral of the induced current pulse
(at depth y = 50 µm) with and without beam divergence
Blue: no divergence
Black: beam divergence
readout strip centre
more than half pitch (35 µm) from the strip centre large effect of beam divergence
strip
strip

## Slide 9
Simulation
below readout strip, less than half pitch from the strip centre, differences not dramatic
further away carriers released by diverged beam drifting to the readout strip dominate the pulses
compare induced current pulses with and without beam divergence

## Slide 10
E-TCT measurements with laser beam parallel to strips can be done in the area under strip
 not too sensitive to beam divergence
compare standard E-TCT and parallel with strips
 E-TCT parallel with strip: information about charge collection profile in irradiated strip detectors
 standard E-TCT: weighting field not realistic, overestimated contribution of charge from larger
detector depths
Simulation:
Trapping: multiply induced current pulses with exp(-t/τeff) and integrate to get the charge:
Warning: electric field here is not realistic and is the same for both trapping time
Strips
Back plane
Back plane
Strips
Charge under readout strip (arb.)
Charge under readout strip (arb.)
depth
depth

## Slide 11
Laser light
HV filter
HV filter
HV
Scope
Laser light
Measurements
Atlas12 mini strip detector
n-in-p
pitch 75 µm
Vfd ~ 370 V
Laser light

## Slide 12
Measurements
charge before irradiation, Bias = 100 V (not fully depleted)
Depth
Charge collection profile:
slice under the readout strip
readout strip centre
strip
strip
Lego version
Back plane
Charge
Beam direction

## Slide 13
Measurements
 charge collection profiles under strip before irradiation similar
 beam spread influences also resolution in y direction
Not depleted
(100 V)
Fully depleted
(400 V)
compare charge collection profiles measured with standard E-TCT and
with laser parallel to strip:
Strips
Back plane
Strips
Back plane

## Slide 14
Measurements
velocity profiles and pulses different:
Velocity profile
Pulses at y = 250 µm deep (near backplane)
Ew different:

## Slide 15
Measurements
irradiated detector Φ = 5e14 n/cm2
normalize to same charge at y = 0 µm for standard and parallel (along strip)
Strips
Back plane
less charge than with standard E-TCT measured near back plane

## Slide 16
Measurements
1e15, 500 V
1e15, 800 V
2e15, 800 V
2e15, 500 V
higher fluences

## Slide 17
Conclusions
E-TCT with laser light parallel to strips can be used to better estimate
charge collection profiles in irradiated strip detectors compared to standard E-TCT
 standard E-TCT overestimates contribution of the back side of the detector
 standard E-TCT good for velocity profiles
because of divergence of laser beam in the detector only limited spatial resolution can be reached
Future work:
try with laser light with shorter penetration depth to get better resolution
measure charge collection profiles with detectors irradiated with 26 GeV protons or pions
where strong double peak field was observed with standard e-TCT:
Example: standard E-TCT charge collection profile
measured with ATLAS07 detector irradiated
with PSI pions to Φeq = 1.6e15 n/cm2
M. Milovanović, 18thRD50 meeting, Liverpool 2011