## Slide 1: Post-mortem from the LEIR start-up Could any issue have been caught during hardware/software commissioning or dry-runs?
Eva Barbara Holzer for the LEIR team
Maria Elena Angoletta, Jerome Axensalva, Michael Andreas Bodendorfer, Steen Jensen, Django Manglunki, Sergio Pasinelli
BI Technical Board
CERN

## Slide 2: LEIR start-up
The cold check-out was reduced to less than one week, from the planned 1 month, because of the accumulated delays from the hardware test
Consequently debugging happened with beam
  Particularly challenging because of the new ion species (Ar)
Controls issues mitigated by presence of excellent controls coordinator
LEIR team appreciates this initiative of a post-mortem
In general, start-up for beam instrumentation was very successful

## Slide 3: BPM
Bunched beam orbit: 16 BPMs H+V
Orbit measurement works
  Not yet completely tested (gains, polarity ,…):
  data not yet used for corrections
Questions:
  How are the BPMs calibrated / re-calibration?  Answer from Lars Soby: calibrated before start-up and several times per year by injecting calibration signal to readout electronics
  Can they measure turn-by-turn?  Answer from Lars Soby: No

## Slide 4: Schottky transversal and longitudinal)
Only used longitudinal for the moment and for un-bunched beam
Only the pick-ups are BI, but not the Agilent network analyzers
Outcome of the meeting 8 September:
  Aim: new BI signal analysis and post processing, which can be operated from the CCC by 2016
  use an adapted version of the ELENA system
  Looking forward to the new instrument
Timing signal by mistake removed, because an old front-end was removed from the control room
  Could have been found during a dry-run

## Slide 5: Current Transformers
5 transformers in injection, extraction and transfer line to PS
2 in the ring:
  ER.MTR12
    Slow transformer, bunched and un-bunched during the cycle
  ER.MTRF12 / BCTFT: (semi) fast transformer
    Accumulation of the injected beam over 70 turns
    Inversion of signal due to mix-up with the cables
    Only detected with beam
    Caused delays, because a negative signal normally means beam losses upstream (e- shower) due to optics problem.

## Slide 6
Screens
  Improved during LS1: material changed to Alumine (Al2O3), filter wheels added
  Used extensively; very happy with the changes
Pepper pot
  Movement works, but application SW is missing, will try to get an existing application
SEM grids in ETL line
  Improved during LS1 – now it measures also the injected beam (as well as the extracted beam)
  Same device – two applications
  Very happy with the improvement
  A problem with the movement was caught during the dry-run

## Slide 7: BGI
Tested and confirmed working by Alexandre Frassier
Specialist device, only operated from the LEIR control room (switching on the HV)
The only transverse profile measurement  very important for operation
Question:
  Still planned to be made operational by non-specialist from the CCC with a copy of the AD?
  Can a FESA class be made operation safe for non-experts? A first version of such a FESA call exist.

## Slide 8
Tune and chromaticity
  BBQ hardware
  Improvement during LS1 by RF and OP: Qmeter application was adapted for LEIR
  Works very well, very happy with the new application
Tomoscop using wall current monitor
  System belongs to RF, working fine
Electron cooler
  Worked immediately, very happy
  Propose to schedule an MD to measure the cooling performance for Ar with the BI experts for IPM and e-cooler