## Slide 1: Non-Destructive Laser-based Emittance Measurement at LINAC4
Project Summary
T. Hofmann, E. Bravin, U. Raich, F. Roncarolo, F. Zocca (CERN)
G. Boorman, A. Bosco, S. Gibson, K. Kruchinin (RHUL)
B. Schmauss (FAU)
27.3.2015
Thomas Hofmann
CERN – Beam Instrumentation (BE/BI/PM)
thomas.hofmann@cern.ch

## Slide 2: Contents
Motivation
Concept
Emittance Measurements at 3 MeV and 12 MeV
Summary
Motivation

## Slide 3: High Luminosity LHC Project
Luminosity: L = N2 * fb/(4π σx σy)
How to maximize the integrated luminosity?
  High brightness beams
  Minimize downtimes ( -> Non-destructive and reliable instruments!)

## Slide 4: CERN Accelerator overview
LINAC 4

## Slide 5: LINAC 4 overview
160 MeV
100 MeV
50 MeV
Transfer line to PSB
Laser Emittance Meter
Particles: H- ions
Top Energy: 160 MeV
Pulsed: 400µs / 40 mA
Length: 80 m
Instrument tests at:
12 MeV / 3 MeV
100 MeV / 50 MeV

## Slide 6: Transverse emittance
Transverse Phasespace
  Position (x,y)
  Angle (x‘,y‘)
  ε2 = <xi2> * <xi‘2> - <xixi‘>2
Conventional Techniques
  3 Profile method
  Quadrupol Scan
  Pepperpot
  Slit & Grid
Destructive / invasive
Angle x`
Position x
Angle y`
Position y

## Slide 7: Contents
LINAC4 Background
Concept of Non-destructive Emittance Meter
Challenges
Emittance Measurements at 3 MeV and 12 MeV
Design for final location at 160 MeV
Summary
Image Concept: Animation shot
Concept

## Slide 8: Concept of Non-destructive Emittance Meter
H-
Bending Magnet
H- beam
Laser
1064 nm
Notes: ICT: new LHC transformer to replace the FBCT in order to solve the beam position dependency issue

## Slide 9: Concept of Non-destructive Emittance Meter
H- beam
H0 beamlet
< 1 ppm
Principle:
H-
Bending Magnet
H- beam
Laser
1064 nm
H- + γ
H0 + e-
Notes: ICT: new LHC transformer to replace the FBCT in order to solve the beam position dependency issue

## Slide 10: Concept of Non-destructive Emittance Meter
H0 Detector
H- beam
H0 beamlet
< 1 ppm
beamlet
after drift
Principle:
H-
H0
Bending Magnet
H- beam
Laser
1064 nm
H- + γ
H0 + e-
Notes: ICT: new LHC transformer to replace the FBCT in order to solve the beam position dependency issue

## Slide 11: Concept of Non-destructive Emittance Meter
H0 Detector
H- beam
H0 beamlet
< 1 ppm
beamlet
after drift
Principle:
H-
H0
Bending Magnet
x-scan
H- beam
Laser
1064 nm
H- + γ
H0 + e-
x’
x
Scan result:
Phasespace
Notes: ICT: new LHC transformer to replace the FBCT in order to solve the beam position dependency issue

## Slide 12: Concept of Non-destructive Emittance Meter
H0 Detector
H- beam
H0 beamlet
< 1 ppm
beamlet
after drift
Principle:
H-
H0
Bending Magnet
x-scan
y-scan
H- beam
Laser
1064 nm
Scan result:
Phasespace
H- + γ
H0 + e-
x’
x
y
y’
Notes: ICT: new LHC transformer to replace the FBCT in order to solve the beam position dependency issue

## Slide 13: State of the Art
SNS: Operational Multi-Station Profile Measurement
  Challenge: Laser-Delivery!
  250 m complex free space delivery line due to radiation environment!
Y. Liu et al., Proc. PAC 2011

## Slide 14: Challenge: Laser-Delivery
Laser | MOPA Fiberlaser (used at LINAC4) | Nd-YAG (used at SNS)
Energy per pulse | 0.1 mJ | 50 mJ
Diameter | ~ 150 µm | ~ 200 µm
Pstrip @ 160 MeV for H- crossing laser | 0.1 % | > 99 %
tpulse | 3…300 ns | 5…10 ns
Repetition rate | 10...500 kHz | < 20 Hz
Beam transport | Fiber | Free Space
Comparison of Laser Systems

## Slide 15
Full screen: Measurements:
Me? measuring something
Prototype Design & Tests

## Slide 16: 3 / 12 MeV Prototype Test Setup
T. Hofmann et al., Proc IBIC 2014

## Slide 17: Laser Focus / Scan / Diagnostics - Setup
10m LMA fiber

## Slide 18: Diamond Strip Detector

## Slide 19: Diamond Raw Signal & Signal Analysis
LINAC4 H- current
H0 particles originating from laserpulses
Background from H0, produced by residual gas collisions
-> Diamond capable to detect H0 beamlets with high linearity and sensitivity

## Slide 20: 3 / 12 MeV Prototype Test Setup
T. Hofmann et al., Proc IBIC 2014

## Slide 21: Results – Beam Profile and Phasespace
Comparison with Slit & Grid
  Beam Profile
  Phasespace

## Slide 22: Summary – Non-destructive emittance meter
What did we achieve?
Prototype design
  Simple fiber-based laser delivery
  H0 Detection based on a diamond detector
Prototype test at 3 and 12 MeV H- beam
  Comparison with Slit&Grid as reference shows excellent agreement
Proof of Principle for Non-Destructive Emittance Meter!
  CERN approved installation for operational use at LINAC4
What is to be done?
Design for final location at 160 MeV
  Longer Fiber delivery -> Non-linear effects?
  High resolution diamond detector -> Readout scheme!
Test of Non-destructive Profile Meter at 50 & 100 MeV beam

## Slide 23: Thanks for your attention!
This project has received funding from the European Union’s Seventh Framework Programme for research, technological development and demonstration under grant agreement no 289191.

## Slide 24: Spare Slides

## Slide 25: Photo Neutralisation
Ref: J.T. Broad and W.P. Reinhardt, Phys. Rev. A14 (6) (1976) 2159
Binding Energy: 0.75 eV
Laser Stripping
Requirements for Laser:
  M2 < 2 for diameter < 200 µm
  PPeak > 1 kW
Residual Gas Stripping Background
  Stripping probability rises with gas pressure
  Stripping probability decreases with particle energy
H-
H2
H0
e-
Beam pipe
    l = length;
;

## Slide 26: LINAC4 & Laser Timing
400 µs - Macropulse
1.2 s
2.84 ns
24 Laserpulses per 400 µs Macropulse
LINAC4
  1.2 s Period length
  400 µs Marcropulse
  352 MHz RF-System
MOPA Fiberlaser
  60 kHz Master Oszillator
  110 ns Pulsewidth
  Modulation via Power Amplifier
17 µs
110 ns
Laser on

## Slide 27: 3 MeV Results - Comparison with Slit & Grid
Beam Profile
Phasespace
Ref: F. Roncarolo et al., Transverse profile and emittance measurements with a laser stripping system during the CERN LINAC4 commissioning at 3 and 12 MeV, Proc. LINAC14, Geneva, Switzerland

## Slide 28: Emittance Value at 12 MeV
Slit & Grid
Emittance:
0.235 mm mrad
Laser Diamond
Emittance:
0.299 mm mrad

## Slide 29: Concept Profile Meter
Laser
H- beam
Electrons

## Slide 30: Detector Possibilities for Electron Collector
Type | Faraday cup | Diamond detector | Silicon PIN Diode | Electron Multiplier
Sensitivity | 1 | 10^4 | 10^5 | >10^6
Rad. Hardness | +++ | ++ | + | +++
Time constant | High! | ps | 16ns | 2ns
Homogeneity | +++ | - | +++ | +++
Saturation | No limit | No limit | 2…3mA | 10mA
Vacuum | good | good | diffcult | good

## Slide 31: Radiation environment at LINAC4 160 MeV area
Dose over lifetime of LINAC4 (red: path of fiber)
-> ca. 50Gy will be radiation load of fiber over whole lifetime

## Slide 32: Simulation – Electron Collection

## Slide 33: Beam at 160 MeV