## Slide 1: Non-invasive Profile Measurement for the CLIC Drive Beam
Adam Jeff
CERN & University of Liverpool

## Slide 2: Contents
adam.jeff@cern.ch
Curtain Gas Jet
  Lab tests at the Cockcroft Institute
Pencil Gas Jet
  Focusing element in production, tests at C.I. soon
Electron-Beam Scanner
  Simulations ongoing (O. Antoniou)
Off-axis Synchrotron Radiation
  Simulations completed

## Slide 3
adam.jeff@cern.ch The CLIC Drive Beam 1 / 24
Beam Energy | to 2.4 GeV
Beam Current | 4.2 A
Pulse Length | 140 μs

## Slide 4
adam.jeff@cern.ch Principle of the gas jet monitor 2 / 24
Generate thin atom gas curtain,
Ionize atoms with primary particle beam,
Extract ions via electric field,
Monitor on MCP, P screen.
Y. Hashimoto et al., Proc. Part. Acc. Conf., Chicago (2001)

## Slide 5
adam.jeff@cern.ch Test Setup @ Cockcroft Institute 3 / 24
Thanks to M. Putignano

## Slide 6
adam.jeff@cern.ch Test Setup @ Cockcroft Institute 4 / 24

## Slide 7
adam.jeff@cern.ch Test Setup @ Cockcroft Institute 5 / 24
Gas Jet
e- beam

## Slide 8
e-
Residual Gas Ions
Gas Jet
adam.jeff@cern.ch Test Setup @ Cockcroft Institute 6 / 24
E
Gas Jet Ions

## Slide 9
adam.jeff@cern.ch Test Setup @ Cockcroft Institute 5 / 24

## Slide 10
adam.jeff@cern.ch Test Setup @ Cockcroft Institute 7 / 24
Gas density in jet cross-section
Vertical position (mm)
Horizontal position (mm)
e-

## Slide 11
adam.jeff@cern.ch Test Setup @ Cockcroft Institute 8 / 24

## Slide 12
adam.jeff@cern.ch Vacuum Effects 9 / 24
V. Tzoganis, A. Jeff, C. Welsch , “Gas dynamics considerations in a non-invasive profile monitor for charged particle beams”,
Vacuum (2014), dx.doi.org/10.1016/j.vacuum.2014.07.009

## Slide 13
adam.jeff@cern.ch Pulsed Jet 10 / 24
Valve Control
Chamber Pressure
Jet Brightness

## Slide 14
adam.jeff@cern.ch Gas Jet Halo Monitor 11 / 24
Shaping of gas curtain possible through sole manipulation of Pressure and Temperature of Gas Reservoir.
1 order of magnitude density difference between core and side strands
M. Putignano, C.P. Welsch, “Numerical study on the generation of a planar supersonic gas-jet”, NIM A 667 (2012)

## Slide 15
adam.jeff@cern.ch Gas Jet Scanner 12 / 24
Planar jet suitable for large beams only
  Difficult to make the jet thin enough for CLIC Drive Beam
  Space Charge also a problem
Solution: Gas Jet Scanner
  Generate a thin pencil jet and scan it through the beam
  Like a wire scanner but non-interceptive
  Still collect ions but position not important: not affected by space charge
  Need a way to generate a thin jet…

## Slide 16
adam.jeff@cern.ch Matter-Wave Focusing 13 / 24
T. Reisinger, S. Eder, M.M. Greve, H.I. Smith, B. Holst, “Free-standing silicon nitride zoneplates for neutral-helium microscopy”, Microelectronic Engineering 87 (2010)
Matter-Wave Focusing for a Thin Neutral Jet
Smallest zone 100 nm
FWHM at focus 2μm !

## Slide 17
adam.jeff@cern.ch Fresnel Zone Plates 14 / 24
The path difference between each successive light ring is equal to 1 wavelength (at the focal point) constructive interference.
Each zone is equal in area
Focal spot size is roughly the width of the narrowest (outer) zone
Compared to traditional lens: no spherical aberration, large chromatic aberration
DeBroglie wavelength ≈ 0.05 nm for room temperature Helium
Focal length of zone plate
Resolution ≈ width of smallest zone (ignoring chromatic effects)
radius of outer zone
width of outer zone

## Slide 18
adam.jeff@cern.ch Photon Sieves 15 / 24
Photon Sieve replaces clear zones of an FZP with a series of holes
Apodised Photon Sieve reduces higher order diffraction, increases central maximum
+ Sharper focusing
- Lower transmission
+ Easier to manufacture
L. Kipp et al, “Sharper images by focusing soft X-rays with photon sieves”, Nature 414 (2001)
Notes: In reality many more rings

## Slide 19
adam.jeff@cern.ch ZEMAX Simulations 16 / 24

## Slide 20
adam.jeff@cern.ch The Atomic Sieve 17 / 24
The Atomic Sieve
Notes: Largest holes diameter 1.5 micron
Smallest 80 nm

## Slide 21: Gas Jet - Conclusions
adam.jeff@cern.ch Conclusions 18 / 24
The test setup at the Cockcroft Institute has demonstrated reliable gas jet operation, and can be used for profile measurement in both continuous and pulsed mode.
Thanks to efficient dumping of the jet and differential pumping in the jet generator, the effect on the beam vacuum system is small.
For measurement of smaller beams with intense space charge, a new gas jet scanner is proposed.
A focusing method based on the deBroglie wavelength of the neutral gas atoms will be tested.
The ‘atomic sieve’ is under production and will be tested later this year.

## Slide 22: Electron Beam Scanner
adam.jeff@cern.ch Electron Beam Scanner 19 / 24
Already used at SNS, but could be difficult for short bunches
Technical student Orfeas Antoniou is investigating
Electrons are shot perpendicularly to the main beam
The deflection of the secondary beam is caught by a monitor
Take the derivative of the maxima to compute the beam Y-profile

## Slide 23: Electron Beam Scanner
adam.jeff@cern.ch Electron Beam Scanner 20 / 24
Plots from O. Antoniou
Theoretical ellipse
CST simulation
Probe beam draws an ellipse if the main beam bunch length is small

## Slide 24: Electron Beam Scanner
adam.jeff@cern.ch Electron Beam Scanner 21 / 24
So far:
  Used a mathematical model to see expected results
  Modeled the scanner in CST
    Investigated how parameter changes affect the results
    Investigated if bunches affect each other
Next steps:
  Better simulation time
  More accurate results

## Slide 25: Off-Axis Undulator Radiation
adam.jeff@cern.ch Undulator Radiation 21 / 24
Synchrotron Radiation is always a nice option for non-invasive diagnostics
Comes ‘for free’ in the combiner rings, turn-arounds
Not so easy in the linac

## Slide 26: Off-Axis Undulator Radiation
adam.jeff@cern.ch Undulator Radiation 23 / 24
Simulations with Synchrotron Radiation Workshop (SRW)
Optical propagation with ZEMAX
All wavelengths
500nm only

## Slide 27: Off-Axis Undulator Radiation
adam.jeff@cern.ch Undulator Radiation 24 / 24
Beam Energy | Undulator Period (cm) | K | Light yield (photons / e-)
300 MeV | 10 | 0.7 | 2 x 10-3
500 MeV | 12 | 1.3 | 4.4 x 10-3
1 GeV | 14.5 | 1.6 | 2.8 x 10-3
2.4 GeV | 15.7 | 1.7 | 7 x 10-4
Undulator period and K parameter to produce visible SR at 2.5 mrad opening angle
Undulator period and strength have to be tuned depending on electron energy
A realistic solution can be found for any energy > 300 MeV

## Slide 28: Thank you for your Attention

## Slide 30: Evaluating the beam profile
Monitor the deflection of the scanning beam
Plot the maxima and compute the derivative