## Slide 1: Glass GEM Charge-up studies
Patrik Thuiner

## Slide 2: Setup
Cu x-ray
tube
GEM
Al shielding
Fe-55 source
Area irradiated by Fe-55 source
Area irradiated by Cu x-ray tube
Collimator

## Slide 3: Rate
Rate of Cu x-ray increasing with x-ray HV current
Rate of Fe-55 stable
Peak position should be stable
  Need to play with binning!
Chart: Rate
| 0.6 | 1 | 3 | 6 | 9 | 12 | 15 | 18 | 21 | 24 | 27 | 30
Fe-55 source and Cu x-ray | 8157 | 9163 | 14529 | 22725 | 30351 | 38132 | 45030 | 51998 | 58916 | 65587 | 72196 | 78513
Fe-55 source | 6778 | 6773 | 6760 | 6790 | 6636 | 6667 | 6470 | 6476 | 6142 | 6125 | 6779
Chart: Peak position
| 0 | 0.6 | 1 | 3 | 6 | 9 | 12 | 15 | 18 | 21 | 24 | 27 | 30
Fe-55 source | 815 | 830 | 820 | 820
Fe-55 source and Cu x-ray | 1080 | 1050 | 1130 | 1280 | 1420 | 1500 | 1550 | 1590 | 1620 | 1670 | 1660 | 1660

## Slide 4: Peak position
Movement of peak position comparable to older measurements with collimated beam
  Rise slower due to lower gain (1.5k instead of 3.5k)
Chart: Normalized peak position Cu x-ray
| 1643.82 | 2739.7 | 8219.0999999999985 | 16438.199999999997 | 24657.3 | 32876.399999999994 | 41095.5 | 49314.6 | 57533.7 | 65752.799999999988 | 73971.899999999994 | 82191
Normalized Peak Position X | 0.64671 | 0.62874 | 0.67665 | 0.76647 | 0.8503 | 0.8982 | 0.92814 | 0.9521 | 0.97006 | 1 | 0.99401 | 0.99401
Normalized Peak Position Y | 0.62264 | 0.59434 | 0.63679 | 0.73585 | 0.82075 | 0.88679 | 0.92453 | 0.95283 | 0.9717 | 1 | 1 | 1
Chart: Normalized Gain vs Rate
| 8584.1999999999989 | 16543.8 | 24503.399999999998 | 32463 | 40422.600000000006 | 48382.2 | 56341.799999999996 | 64301.4 | 72261 | 80220.600000000006
Low Gain Collimated Beam (peak position) | 0.80933 | 0.87885 | 0.93049 | 0.95929 | 0.97319 | 0.98113 | 0.98808 | 0.99503 | 1 | 1
Low Gain Collimated Beam (Anode Current) | 0.72178 | 0.84759 | 0.89166 | 0.91412 | 0.95597 | 0.97057 | 0.98105 | 0.98639 | 0.99733 | 1

## Slide 5: Charge sharing
Charge sharing between x- and y-strips
Charge sharing decreases with higher rates!
Chart: Peak position
| 0.25684687499999997 | 0.42807812499999992 | 1.2842343749999998 | 2.5684687499999996 | 3.8527031249999997 | 5.1369374999999993 | 6.4211718749999998 | 7.7054062499999993 | 8.9896406249999998 | 10.273874999999999 | 11.558109374999999 | 12.84234375
Peak Position X [chn] | 1080 | 1050 | 1130 | 1280 | 1420 | 1500 | 1550 | 1590 | 1620 | 1670 | 1660 | 1660
Peak Position Y [chn] | 660 | 630 | 675 | 780 | 870 | 940 | 980 | 1010 | 1030 | 1060 | 1060 | 1060
Chart: 
| 0.25684687499999997 | 0.42807812499999992 | 1.2842343749999998 | 2.5684687499999996 | 3.8527031249999997 | 5.1369374999999993 | 6.4211718749999998 | 7.7054062499999993 | 8.9896406249999998 | 10.273874999999999 | 11.558109374999999 | 12.84234375
Charge sharing | 1.6364 | 1.6667 | 1.6741 | 1.641 | 1.6322 | 1.5957 | 1.5816 | 1.5743 | 1.5728 | 1.5755 | 1.566 | 1.566

## Slide 6: Rate and Flux estimation
Estimated rate used for calculation of flux
Hit map shows hits of 2x2 strips, ~ 800x800 μm²
Chart: Rate without Fe-55 source
| 0.6 | 1 | 3 | 6 | 9 | 12 | 15 | 18 | 21 | 24 | 27 | 30
Rate w/o Fe-55 [Hz] | 1310.4402 | 2338.2959 | 8053.7319 | 16553.8073 | 24691.7912 | 32454.198 | 39962.5637 | 47575.5933 | 54468.4636 | 61664.8705 | 68454.4933 | 75028.1362
 | 2338.2959 | 8053.7319 | 16553.8073 | 24691.7912
Chart: Extrapolated flux
| 0.6 | 1 | 3 | 6 | 9 | 12 | 15 | 18 | 21 | 24 | 27 | 30
Extrapolated Flux w/o Fe-55 offset [MHz/cm²] | 0.25685 | 0.42808 | 1.2842 | 2.5685 | 3.8527 | 5.1369 | 6.4212 | 7.7054 | 8.9896 | 10.2739 | 11.5581 | 12.8423

## Slide 7: Flux estimation
Chart: Current over extrapolated flux
| 1.2842343749999998 | 2.5684687499999996 | 3.8527031249999997 | 5.1369374999999993 | 6.4211718749999998 | 7.7054062499999993 | 8.9896406249999998 | 10.273874999999999 | 11.558109374999999 | 12.84234375
Drift Current [nA] | 0.15 | 0.3 | 0.9 | 1.15 | 1.4 | 1.8 | 2.1 | 2.4 | 2.75 | 3.6
Estimated Anode Current [nA] | 1.9792 | 3.9583 | 5.9375 | 7.9166 | 9.8958 | 11.875 | 13.8541 | 15.8333 | 17.8124 | 19.7916

## Slide 8: Next
Cu x-ray
tube
Uncollimated
Beam
GEM
Al shielding
Fe-55 source
Area irradiated by Fe-55 source
Area irradiated by Cu x-ray tube

## Slide 9: Next step
Remove collimator
Get data from areas with high flux and low flux
  Area with high flux should behave as with collimated beam
  Area with low flux should behave like measurements with low flux before
Area irradiated by Fe-55 source
Area irradiated by Cu x-ray tube

## Slide 10: Thank you!
Questions?