## Slide 1: Instrument Modeling Response - FEC response
Yuki Shimizu
SEUC/JAXA

## Slide 2: Contents
Calibration of detector response
Proposals how to introduce FEC response to simulation data

## Slide 3: Detector response calibration
Detector components (scintillators, photosensors and amplifiers etc.) have each characteristics such as yield, gain and linearity. While these characteristics have been investigated during trial production, those of the flight model itself must be measured by flight model tests and introduced into our simulation in order to improve the accuracy.
The pulse height corresponding to the MIP energy deposit (muon) and the response function of FEC measured in the ground tests are available.
Scintillator
Photosensor
Amplifier
ADC
Photons
(Photoelectrons)
charge
voltage
light yield
gain
gain
linearity
saturation
linearity
deposit
energy
digital
data
muon measurement (MIP calibration)
FEC
electric performance test (response curve)

## Slide 4: Electrical performance test for FECs
The response curves of FECs has been measured in functional tests.
The output of FECs is shown as a function of the input charge.
(Further analysis is necessary to fix parameters as told yesterday.)
Pulse
generator
Capacitor
C
Q=CV
V
TASC-FEC (X side)
 | Photo detector | Shower particle range | Output charge range
CHD | PMT | 1 – 1400MIP | 500fC - 700pC
IMC | 64 anode | 1 – 2500MIP | 8fC - 20pC
 | dynode | 1 – 400MIP | 40fC – 8pC
TASC | PMT | 1 – 1800MIP | 400fC – 700pC
 | APD | 1 – 5000MIP | 2fC - 10pC
 | PD | 103 – 5×106MIP | 2fC – 10pC
high gain
low gain
×～30

## Slide 5: Proposal to introduce FEC response to simulation
Taking into account the detector response to muons, the deposit energy that is calculated in simulation shall be converted to the charge linearly. The charge shall be digitized with employing the response curve. Saturation of ADC shall be also taken into account (usually 215-1).The charge could be converted to pseud-energy with a nominal factor (mean MIP energy / corresponding charge) (e.g. 22MeV / 400fC for PWO-PMTs of TASC, depending on the PMT gain).The pseud-energy [MeV] and digitized output [ADU] shall be introduced to the unified data format. Details must be discussed among M&S team.
high gain
low gain
response curve function
saturation
Such conversion is also necessary for generation of simulated level0 and 1 data.
Should be discussed with DHA team.