## Slide 1: Calibrations and tests carried out in manufacturing
Yuki Shimizu
SEUC/JAXA
2014. Oct. 14th

## Slide 2: Contents
Manufacturing process
Test/calibration and some preliminary results
  FEC electric performance test
  Functional test for detector assemblies
  Thermal vacuum test for detector assemblies
  Functional test for system assembly
Current status and future plan of the flight model test
Calibration and characteristics parameters of detector components

## Slide 3: Manufacturing process
Scintillator
FEC
MDC
Detector Assembly
System Assembly
To launch site
Structure
Detector structure
Function test
Cosmic ray
UV pulsed laser test
Thermal vacuum test
Cosmic ray
Function test
Cosmic ray
Acoustic test
EM compatibility test
Thermal vacuum test
Cosmic ray
Alignment measurement
Function test
Signal check
Electrical performance (FEC)
Cosmic-ray simulated pulse test
(function of FEC and MDC)
Alignment measurement
Discharge resistance test (IMC)

## Slide 4: Electrical performance test for FECs
As a part of functional test of FECs, the following electrical properties has been measured. Instead of photosensors, pulsed charge was input to each channel.
Pedestal
Onboard test pulse (fixed charge)
Linearity of readout (dynamic range)
Discriminator response
Pulse
generator
Capacitor
C
Q=CV
V
TASC-FEC (X side)
 | Photo detector | Shower particle range | Output charge range
CHD | PMT | 1 – 1400MIP | 500fC - 700pC
IMC | 64 anode | 1 – 2500MIP | 8fC - 20pC
 | dynode | 1 – 400MIP | 40fC – 8pC
TASC | PMT | 1 – 1800MIP | 400fC – 700pC
 | APD | 1 – 5000MIP | 2fC - 10pC
 | PD | 103 – 5×106MIP | 2fC – 10pC
Charge
injection
IMC-FEC (single layer)
VA chip ASSY

## Slide 5: CHD-FEC
Measurement range
- 500fC～700pC
Relative errors from the best fitting lines are less than 2% above 600fC.
The curves tend to fall to the right. One more parameter could be introduced.
Fitting line

## Slide 6: TASC-FEC-X1 (PMT)
high gain
low gain
Measurement range - 400fC～700pC
Range of use
Range of use
Relative errors from the best fitting lines are less than 2% in the range of use.
The curves of the low gain tend to fall to the right (similar with CHD).
×～30
overlap
overlap

## Slide 7: TASC-FEC (PD/APD)
Measurement range
APD - 2fC～15pC
PD - 2fC～15pC
Range of use
Range of use
Range of use
Range of use
overlap
overlap
overlap
overlap
Relative errors from the best fitting lines are mostly less than 2% above 10fC.
The maximum measured points (12pC) fall comparing with other points.
Below 10fC, errors of some channels are large due to noise.
⇒Further investigation is necessary.
high gain
low gain
×～30

## Slide 8: Functional test for detector assemblies
Detector
FEC
Power supply,
Data transfer etc.
During construction, functional tests were conducted for every two layers.
Electric functional test
  Pedestal
  Onboard test pulse
Muon measurement
  MIP calibration
  Discriminator response
UV pulsed laser test (PD/APD)
  Response check for PDs
  Dynamic range of PD/APDs
  ⇒ Ozawa and Asaoka’s talk
Trigger circuit
(NIM/VME)
Plastic scintillator
CHD/IMC detector assembly
TASC detector assembly
IMC FEC
TASC FEC
CHD FEC
Trigger signal
Test equipment

## Slide 9
Detector with
black plastic sheets
(light shielding)
Detector assembly
NIM and VME system
(for trigger scintillators)

## Slide 10: Setup of cosmic-ray measurement
Trigger
scintillator
Detector
assembly

## Slide 11: Muon deposit energy distribution
CHD
X, Paddle #6
TASC (PMT)
X-1, log #9
X-1, fiber #225
IMC
These results have been used to adjust the PMT gain of each channel.
The gain of all the channels should be even in each layer equipped with LD circuits.
All events
Tracking selection

## Slide 12: Discriminator response
All events
Events with a LD flag(single)
Discriminator efficiency
100% efficiency
IMC-Y4 (Dynode)

## Slide 13: Thermal vacuum test for detector assemblies
HV-BOX
(spare)
MDC
(FEM)
TASC
IMC/CHD
Thermal vacuum chamber
The detector components of CAL were subjected to thermal vacuum test which simulated space environment for the following purposes:
Verify thermal design of IMC/CHD and TASC detectors
Confirm discharge resistance of CHD and TASC
Validate operation of the instruments between 5 and 45℃.
coolant
Power for FECs were supplied by MDC-FEM (function equivalent model).
HV were supplied by a spare of HV-BOX (provided by IC).
During the test, detectors were set to cosmic-ray measurement except power-off or functional test.

## Slide 14: Experimental setup
IMC/CHD
TASC
MDC
HV-BOX
Thermal vacuum chamber in IHI Aerospace
CHD/IMC and TASC were separately placed and operated.

## Slide 15: Temperature profile
～+45℃
～+5℃
Muon measurement
Function test
Function test
Function test
Function test

## Slide 16: Examples of cosmic-ray data
Example of muon track (IMC)
Channel Number
Counts [ADU]
layer #5
layer #6
layer #7
layer #8
layer #4
layer #3
layer #2
layer #1
Example of cosmic ray shower (TASC)
PMT gain adjustment (IMC dynode)

## Slide 17: System assembly test
coolant
20℃
Low rate
data
Middle rate
data
CALET system assembly
HV-BOX
MDC
After assembling of all the components into the main structure, function test was carried out with supplying power to all the components including CAL and CGBM.
CAL
Cosmic ray measurement (single track setting)
CGBM
Background spectrum measurement (PH, GRB)

## Slide 18: Screenshot of data monitor

## Slide 19: Examples of CAL events
Muon event
Showering event
Gain correction of each channel in the analysis has not been completed.
Further analysis is on-going.
log( pulse height )

## Slide 20: Current status and future plan
Last week, Initial functional test of the system assembly was finished in IHI Aerospace.
CALET PFM will be transferred to Tsukuba in this week.
In Tsukuba, environmental tests are going to be conducted (EMC, acoustic, thermal vacuum tests).
During these tests, cosmic-ray measurements are expected to be carried out as parts of functional tests.
After the environmental tests, final functional test is planed to be performed. Then, CALET PFM will be transferred to the launch site (Tanegashima Space Center).
13mΦ space chamber (Tsukuba)
Acoustic test for CALET STM
EMC test room (Tsukuba)

## Slide 21: Calibration and characteristics parameters of components
List of parameters obtained from component level tests
Response curve function of FECs
Mean or peak pulse heights of muon energy deposit distribution
RMS noise (including electronics and leakage current)
Mean pulse height of onboard test pulses
Response function of discriminators
Efficiency curves of discriminators
These parameters will be distributed to international collaboration after further analysis.

## Slide 23: IMC-FEC
Input charge [pC]
Output [ADU]

## Slide 24: CHD
Offset
7/16 15:30 (high temp.)
7/17 19:30 (low temp.)
7/18 20:00 (low temp.)
7/19 10:40 (high temp.)
7/19 16:40 (room temp.)
Noise
Test pulse
160fC

## Slide 25: IMC
7/16 15:30 (high temp.)
7/17 19:30 (low temp.)
7/18 20:00 (low temp.)
7/19 10:40 (high temp.)
7/19 16:40 (room temp.)
Offset
Noise
Test
pulse
2fC

## Slide 26: TASC (APD)
Offset
7/16 15:30 (high temp.)
7/17 19:30 (low temp.)
7/18 20:00 (low temp.)
7/19 10:40 (high temp.)
7/19 16:40 (room temp.)
Noise
Test
pulse
1fC