## Slide 1: CALET Instrument Characterization Parameters for Flight Data Analysis
T. Gregory Guzik Louisiana State University
CALET TIM October 2014 (Japan)

## Slide 2: Instrument Characterization
Part of initial document concerning on-orbit calibrations of CALET (July 2014)
We assume that during development of the CALET FM the responsible company(s) are also collecting data characterizing the flight instrument detectors and readout electronics
These data are very important for the flight data analysis, to assess changes in the instrument over time and as input to a “CALET Flight Instrument Model”
Fundamental questions include:
  What instrument characterization data will be released?
  What is the schedule for releasing these data to international collaborators?

## Slide 3: Detector Data
Attenuation lengths for all CHD, IMC and TASC flight scintillators.
  Japan Response: Not available
  Next step: Fold into detector response map
Response maps for all detectors in flight configuration along the full length of the scintillator.
  Japan Response: Not available
  Next step: To be done on-orbit or with muon data
Physical location position map of all flight scintillators.
  Japan Response: Available by 3D measurements
  Next step: Require more information about this procedure. Also need date for data release.

## Slide 4: Thermal Characteristics
Thermal coefficients for all CHD, IMC and TASC flight scintillators plus thermal coefficients for all corresponding detector readouts from light sensor through ADC.
  Japan Response: No specific measurement for this kind of data.
  Next step: Perhaps we can get away with just one set of thermal coefficients from the scintillator through the ADC. However, this will likely be difficult to calibrate on-orbit.
Physical location position map of all thermal sensors.
  Japan Response: Available
  Next step: Need date for when these data will be released
I assume that calibrations for all individual housekeeping sensors (i.e. temperature, voltage, current, etc.) have been performed as part of the pre-launch work and these data will be made available.
  Japan Response: Since we will convert them into physical value in Level1, I think you don't need them for house keeping data (or at least their priority is low).
  Next step: Sensor calibrations used in the Level0 to Level1 conversion should be made available.

## Slide 5: Readout Electronics
High statistics pedestal data for all 7,932 readout channels.
  At least 2,000 (external ?) triggers (preferably ~5,000 triggers) of the instrument with “Zero Suppression” turned off.
  Data in Level 1 format with both COM and CAL records so that the full state of the instrument (e.g. temperature, voltages, etc.) is recorded.
  Provides baseline for the individual pedestal distribution shape, mean and characteristic width that can be compared with much lower statistics pedestal distributions measured during flight.
  Japan Response: Should be obtained during FM test
  Next Step: Need date for when these data will be released.
Detailed data on the amplifier gain for all 7,932 readout channels.
  Data file, obtained during ground testing, that contains ADC value versus test pulse setting covering the full ADC range (i.e. 0 to 65,535).
  Data provide the basic understanding of amplifier linearity, relative gain and ADC behavior necessary to compare with flight amplifier measurements.
  Japan Response: Test pulse only for one charge is available during flight . The linearity obtained at company is useful.
  Next Step: Is ground test data as described here planned to be available & released?

## Slide 6: Voltage Dependencies
PMT anode output versus high voltage setting for every flight CHD, IMC and TASC flight PMT.
  Japan Response: Available
  Next Step: Need date for when these data will be released.
APD output versus bias voltage setting for every TASC flight APD.
  Japan Response: Will be available
  Next Step: Need date for when these data will be released.
PD output versus bias voltage setting for every TASC flight PD.
  Japan Response: PD has only one bias voltage, and is obtained
  Next Step: Need date for when these data will be released.
Ground level muon data with instrument in flight configuration.
  Data in Level 1 format with both COM and CAL records so that the full state of the instrument is recorded.
  Japan Response: No comment
  Next Step: Need date for when these data will be released.