## Slide 1: Performance tests in Tsukuba SC and results (“PI calibration”)
T. Sakamoto, R. Inoue, Y. Kawakubo,
S. Terasawa, I. Takahashi (AGU)
on behalf of CGBM team

## Slide 2
Goal:
Collect the fundamental data using the flight detectors for the development of the energy response function.
Date:
2014/01/14 – 01/28 Tsukuba Space Center
Test contents:
2D scan of the surface of the detectors by the collimated X-rays and gamma-rays (HXM).
Energy linearity and resolution (HXM/SGM)
Angular response (HXM/SGM)
GRB trigger test (SGM)

## Slide 3: 2D scan of HXM
CGBM
試験ラック
X線発生装置
X-Y ステージ
X-Y stage
X-ray Generator
Collimator of X-ray Generator
(φ 1.5 mm)
Collimator of RI
(φ 3 mm)
Lead block
RI

## Slide 4: Scanning points
Fe@6.4keV
Y@14.9keV
Na@511keV
5mm pitch, total 149 points
X-ray generator
5mm pitch total 197 points
X-ray generator
6mm pitch, Total 149 points
Radio isotope (22Na)
Blue circle : PMT’s diameter 60.96 mm
Red circle : Crystal surface diameter 66.04 mm
Green circle : Crystal bottom diameter 71.76 mm
Al
Al
LaBr3(Ce)
PMT
X[mm]
X[mm]
X[mm]
Y[mm]
Y[mm]
Y[mm]
Scanning order

## Slide 5: HXM: 6.4 keV (Fe) 2D map
Gain map
Count rate map
Energy resolution map
HXM
1
HXM
2
HXM1: No count-rate at the upper part (Y > 0) of the crystal
HXM2: Donuts shape feature in the gain map
HXM1 & 2: Count-rate is not uniform across the crystal surface

## Slide 6: HXM: 6.4 keV spectrum
HXM1: 6.4 keV spectrum at X=0
Red: Y=0, green: Y=-20, blue: Y=-30, light blue: Y=20, pink: Y=30
HXM2: 6.4 keV spectrum at X=0
6.4keV
6.4keV
Strength of 6.4 keV line of HXM1 decrease at off-center points. Overall strength of 6.4 keV line is weaker for HXM1 comparing to that of HXM2.
HXM2 doesn’t show a strong position dependence.
HXM1
HXM2

## Slide 7: HXM: 511 keV (22Na) 2D map
Gain map
Count rate map
Energy resolution map
HXM
1
HXM
2
HXM1 & 2: Count-rate is uniform up to the diameter of the crystal surface (66 mm).
HXM2: Both gain and energy resolution show a donuts shape variation at the surface.

## Slide 8: HXM: 511 keV spectrum
HXM1: 511 keV spectrum at X=0
Red: Y=0, green: Y=20, blue: Y=30, light blue: Y=-20, pink: Y=-30
HXM2: 511 keV spectrum at X=0
・No big difference in the count rate between HXM１ and HXM2.
・Gain shifts at the lower channel range from center to edge of the surface.
・HXM2 shows the highest gain at Y=±10 mm position.
HXM1
HXM2

## Slide 9: HXM: Linearity and Energy resolution (High gain)
HXM1
HXM2
-0.52 ± 0.11
-0.44 ± 0.03

## Slide 10: HXM: Linearity and Energy resolution (Low gain)
HXM1
HXM2
-0.34 ± 0.01
-0.29 ± 0.01

## Slide 11: SGM: Linearity and Energy resolution
High gain
Low gain
0.51 ± 0.07
0.52 ± 0.05

## Slide 12: HXM: Angular response
HXM1
HXM2

## Slide 13: SGM: Angular response

## Slide 14: SGM: GRB trigger test
 | Test | Result
Test #1 | High SNR: should trigger | Y
Test #2 | Low SNR: should not trigger | N
Test #3 | Short and high SNR: should trigger | Y
Disk
RI
(Test #1)
(Test #2)
(Test #3)
7σ
13σ
5.5σ
Time since trigger [s]

## Slide 15: Possible Explanation of HXM1 issue
LaBr3 itself (somehow) becomes an X-ray absorber
  10-25 mm thick LaBr3 can explain the result.
It is very important to monitor the low energy performance of HXM1.
HXM1: 6.4 keV Count-rate map

## Slide 16: Summary
Successfully collected a fundamental data to construct the energy response function of CGBM
Low energy efficiency issue for HXM1
Proposing future test plan
Irradiate 57Co to HXM1 to investigate whether the low energy performance degrades as a function of time or not
Irradiate 109Cd and 22Na at the several positions

## Slide 18: Measurement of Spare HXM (HXM0)
Setup
2-inch PMT
HXM0
57Co
Black sheet
Date: April 22-26, 2014

## Slide 19: Measurement of Spare HXM (HXM0)
6.4keV
14.4keV
HXM0
HXM1
HXM2
PI Cal
PI Cal
No correction
Absorption by the air
Corrected value
Relative intensity to 122 keV

## Slide 20: HXM: 14.9 keV (Y) 2D map
Gain map
Count rate map
Energy resolution map
HXM
1
HXM
2
HXM1: Count rate of HXM1 decreases at Y>0 surface.
HXM2: Donuts shape feature in the gain map (similar to the 6.4 keV map)

## Slide 21: HXM: 14.9 keV spectrum
HXM1: 14.9 keV spectrum at X=0
Red: Y=0, green: Y=-20, blue: Y=-30, light blue: Y=20, pink: Y=30
HXM2: 14. 9 keV spectrum at X=0
14.9keV
14.9keV
Strength of 14.9 keV line of HXM1 decrease at off-center points. Overall strength of 14.9 keV line is weaker for HXM1 comparing to that of HXM2.
HXM2 doesn’t show a strong position dependence.