## Slide 1: LHC BLM system: Status of changes during LS1+
Christos Zamantzas on behalf of the BLM team.
BI Technical Board December 2014

## Slide 2: Overview
Hardware Changes
  Tunnel installation & Acquisition electronics
  Surface installation & Processing electronics
Firmware/Server Changes
  BLETC and BLECF
Application Changes
  Updates, additions and wishlist
Verification & Commissioning
  RSM, Dry-Runs, Checkout, Beam Commissioning
Injection Interlock Inhibit
  Solution agreed and planning

## Slide 3: LHC BLM system: Hardware Changes
Listing main actions to be performed or developed in the system

## Slide 4: Modifications in the tunnel installation
Dismantle all detectors in the ARC and DS (~2500) and
Dismantle 70% of the detectors in LSS (~1000)
  in order to allow intervention in the interconnections
Exchange 40 multiwire cables with the NES18 type
  Reduce noise on 240 detectors
  Continuation of campaign to exchange all noisy installations
Relocate 816 detectors
  New supports, cable extensions, cable trays needed
  Improve protection
  see Eduardo Nebot Del Busto’s presentation for the reasons
ECRs in preparation
New locations of the detectors





## Slide 5: Modifications in the tunnel installation
Exchange 360 acquisition crate backplanes
  Improve stability, add features
Modify 309 signal distribution boxes
  Allow reset of the complete sector
Installation/Connection to WorldFIP
  Straight Section needs additional electronics
  Allow remote access per card
Add everywhere High Voltage divider boxes
  Increase accuracy
Modify 40 High Voltage distribution boxes
  Add suppressor diodes and resistors
  Only in areas with high losses
ECR in preparation
For more info see Ewald Effinger’s presentation in 73th MPP
Improved backplane




to do

## Slide 6: Modifications in the ACQ electronics
Modify all BLECF modules (~700)
  Change limit for the HV level detection flag
    currently too restrictive for proper use by the SIS.
    change detection level from 1370V to 950V
  Change laser settings of the GOH
    increase transmission power to give more margin of attenuation
    add ~ 30% more power

cancelled

## Slide 7: Modifications in the surface installation
Replacement of all racks
  Add temperature regulated racks
  Remove and disconnect complete installation (crates, cables, power etc)
  Reconnection/replacement of all fibre patchcords (~1600 fibres)
Modifications for apt interventions
  Replacement of all rear connection cables
  Allow direct access to power supplies and rear rack connections
Add two new crates for possible future Injection Interlock Inhibit implementation:
  Install additional racks in P2 and P8
  Separate problematic monitors for injection
  see also Wolfgang Bartmann’s presentation




## Slide 8: Modifications in the TC electronics
Maintenance of all processing modules (~400)
  Fix or Replace ~ 20% of mezzanines
  Clean-up of the optical adaptors and connectors
Shuffle optical links
  Expect to improve availability by removing common mode failure
  80% of optical link errors and failures were in LK: 1-2
E2000-APC input connections to four optical fibres.
Optical Link 1 input is less reliable
Accumulation of dust on the connectors



## Slide 9: Expert Name Convention Change
BLMQI.06R7.B1E30_MQMxxx(_S)
BLM location or feature:
Q – Quadrupole
B – Dipole (Bending)
T – Collimator (Target)
E – Extra
D – Dump Line
P – Direct Dump (Protection)
C – Cryo
I – Injection lines (TI2,TI8,TT41)*
X – eXperiments
2 – Coupled detectors
M – Mobile detectors
1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 Official name
Position:
01-34 - number of cell
L or R - side from IP
1-8 - IP
xxxx - distance from beginning of TD [dm]
Beam relation:
B1 – Beam 1
B2 – Beam 2
B0 – Centre
Optional:
_S – second coupled monitor
(counting from beam direction)
_T – vertical slice test monitor
_DUMP – inside beam dump
Position on the element:
for B1/2 from beam direction:
10+i at the entrance
20+i at the middle
30+i at exit
for B0 from IP:
10 – first
20 – second
Observed elements:
- Up to two elements (e.g. MQ, MQXA or MBA.A-MBB)
- Function (batt, hv)
- Distance from IP6 [m] _ Change of TD diameter [mm]
- Distance from beginning of Dump Core [m]
Detector type:
I – IC
S – SEM
L – LIC
F – FIC
D – Diamond
S – Silicon
B – ACEM (Bunch meas.)
C – Channel Check
Transverse position:
E external
I internal
C central
T top
B bottom
L left to beam direction
R right to beam direction

## Slide 10: Database Updates
Completed:
  900 channels have been updated in MTF, LAYOUT and LSA
  21,000 new Logging variables created
  2,000 Logging variables renamed
On-going actions:
  Data on 300 channels to be updated
  Slots for 30 new channels to be added
  Rename of 10,000 Logging variables

## Slide 11: Incident in IP6: 6R
Ewald

## Slide 12: Incident in IP6: 6L
Ewald

## Slide 13: 2014/12/02: Status of the installation
All crates operational
3,216 channels complete, 687 channels with issue
  280 ch. from S45 (cards not yet installed)
  95 ch. from IP6 (cards in repair)

## Slide 14: LHC BLM system: Firmware & SERVER Changes

## Slide 15: Current State of Firmware/FESA
All functionality present in Run 1 is again operational with the new CPU for both BLETC and BLECS modules.
Tested:
  Logging (measurements and statuses), XPOC, PM, IQC
  Read/write of threshold values and system parameters
  Online MCS check
  Sanity Checks

## Slide 16: Modification of the BLECS firmware
Combiner and Survey firmware additional modifications:
  Higher priority:
  Improve regular automatic system checks
    block increase of input offset
    add timeout for check completion
  Improve connectivity check
    Increase range of values allowed
  Improve Energy value reception and logging
    fast changes of energy values are propagated to the BLETC modules but not logged
  Lower priority:
  Add better compatibility with new CPUs
    new VMEbus core with MBLT mode
    new memory map optimized for block transfers
  Preparation for the “Injection Interlock Inhibit” feature

## Slide 17: Modification of the BLETC firmware
Processing Module firmware additional modifications:
  Higher priority:
  Add better compatibility with new CPUs
    new VMEbus core with MBLT mode
    new memory map optimized for block transfers
  Improve PM and UFO Buster buffers
    increase size to 43,690 samples per channel
  Improve XPOC buffer
    split data and triggers per beam
  Lower priority:
  Improve data for Collimation Beam Based Alignment
    increase data rate (now at 12.5 Hz)
FESA server will need to be updated accordingly

## Slide 18: LHC BLM system: Software Changes

## Slide 19: Threshold Management Application
View on parameterization and correction functions
Blazej & Stephen

## Slide 20: Threshold Management Application
View and Plotting of Threshold values
Blazej & Stephen

## Slide 21: Threshold Management Application
Comparison of thresholds for different families and different dates
Blazej & Stephen

## Slide 22: Threshold Management Application
Application ready wrt specifications
  GUI implemented
  DB middle-tier connection
To Do:
  Improve plots (SW)
  Fixing bugs and data in DB (CO)
  Computation algorithms missing (BL & CO)
  Extensive testing of the application and DB logic of threshold values calculation (BL & SW)
Blazej & Stephen

## Slide 23: Other Applications
Moved to Maven and some corrections:
  System Parameters application
  Threshold Values application modifications
  Connectivity Check results viewer
  Master table viewer
Expert Applications are all operational
  Next year to be moved also to Maven
To correct:
  System Status application (via Concentrator)
  Mobile Monitor Setup application
Blazej & Stephen

## Slide 24: Application Improvement wishlist
System Parameters application (BE/SW)
  automatic fill of serials when a card is exchanged
  step increase of connectivity check limits
System Status application additions (BE/OP)
  give global overview
  provide system error cause
System Management application (BE/OP and BE/CO)
  Generate new monitors/parameters, Drive settings to electronics, Initiate/Abort checks
  provide a clear procedure for actions to be taken
  avoid errors due to complexity of the generic applications (generation, trim etc)
Monitor Factor application (BE/OP)
  show history
  compare two points in time
  roll back of changes (very useful for MDs, etc.)

## Slide 25: LHC BLM system: Verification & Commissioning

## Slide 26: Issue Tracking
On-going actions:
  Check all optical connections
  Check behaviour of each card
  Check behaviour of each channel
  Check all interlock connections
To start in 2015:
  Update all serial numbers
  Calculate new Connectivity check limits

## Slide 27: Radiation Source Measurements
Last two sectors pending:
  S45: W49 installation, W50 measurements
  S34: W51 measurements
In 2015, selective measurements in areas with many issues
Slava

## Slide 28: Dry-Runs
Completed already 3; several issues identified and corrected
W50 dry-run (last) to verify:
  triggered acquisitions (IQC, XPOC, PM) fixed displays,
  acquisition of continuous data not blocked when trigger arrives
IQC display during Dry-Run, November 6-7 2014

## Slide 29: System Checkout
When the installation is complete and all operational: 2 major and time consuming checks
Exhaustive threshold triggering
  automatic test in all the LHC crates
  30-50 hours
Large scale HV modulation tests
  each test consists of 200 x modulation checks to verify the integrity of the monitors and connections,
  Results are used to calculate accurate detection limits
  3-5 nights

## Slide 30: Commissioning with Beam
Tests to perform with beam:
Test transmission time USER_PERMIT
  Measure system latency between beam loss and interlock
MPS functionality of the BLM
  All crates and channels can provoke an interlock
Interface of direct BLMs with the LBDS
  Verify the system can provoke an interlock
Performed together with FMCM by OP:
Threshold and reaction time at 450GeV
Threshold and reaction time at physics energy

## Slide 31: LHC BLM system: Injection Interlock Inhibit
Details on the modifications planned and the new functionality

## Slide 32: Modifications: Installation
Move relevant detectors to separate acquisition cards
  See ECR for injection BLMs for the list of monitors involved.
  Re-cabling of signal and power distribution is necessary for some.
  Note: we asked initially the current grouping of monitors to not be broken. This was later agreed to be an unrealistic demand.
Add two new processing crates
  Connect to the standard interlock daisy chain between crates
  Will be the last in the chain to avoid inhibiting other crates’ interlocks
Add new processing modules and connect fibres.
Update the MTF, LAYOUT and LSA databases with this configuration.
All to be completed during LS1

## Slide 33: Configuration in SR2 & SR8
BLECS
1
U
BLETC 1
BLETC 16
BLETC 2
U
M
M
UA
M
UB
MA
MB
U
Crate 1
Termination
U
M
BLECS
2
U
BLETC 1
BLETC 16
BLETC 2
U
M
M
UA
M
UB
MA
MB
U
Crate 2
Termination
U
M
BLECS*
4
U
BLETC 1
BLETC 16
BLETC 2
U
M
M
UA
M
UB
MA
MB
U
Crate 4 (I)
Termination
U
M
Maskable beam permit
Unmaskable beam permit
Unmaskable Beam Info
Maskable Beam Info
Legend:
Interlock daisy chain from BLETC 1 to BLECS
Interlock daisy-chain from BLECS 4 to CIBUs
CIBU Maskable
BIS network
CIBU UnMaskable
Configuration
in SR2 & SR8
UBI
MBI
Termination

## Slide 34: Modifications : Hardware
Modify the BOBM (BST master) configuration to forward a timing event related to the injection
  There are several to choose from, e.g. HIX.FW-CT, HIX.W100-CT, HIX.W20-CT, HIX.AMC-CT.
Modify the BOBR (BST receiver) configuration to distribute the timing event in the backplane of the crate.
  A pulse is broadcasted to all cards through a dedicated line of the VME64x P0 connection whenever the event is received
Modify the BLECS firmware to force the Beam Permit line for a fixed period of time when it receives the injection signal and certain other conditions are satisfied.
  Persistent FESA settings per crate will (in the future can become part of the MCS parameters):
    Activate/Disactivate the inhibit functionality.
    Define the time the beam permit is forced to TRUE
  Inhibit of the output will happen only under certain conditions:
    during injection,
    energy is 450 GeV,
    …

## Slide 35: Interlock Inhibit Functionality
BLECS
1
U
BLETC 1
BLETC 16
BLETC 2
U
M
M
UA
M
UB
MA
MB
U
Crate 1
Termination
U
M
BLECS*
4
U
BLETC 1
BLETC 16
BLETC 2
U
M
M
UA
M
UB
MA
MB
U
Crate 4 (I)
Termination
U
M
Maskable beam permit
Unmaskable beam permit
Unmaskable Beam Info
Maskable Beam Info
Legend:
Interlock daisy chain from BLETC 1 to BLECS
Interlock daisy-chain from BLECS 4 to CIBUs
CIBU Maskable
BIS network
CIBU UnMaskable
UBI
MBI
Termination
Injection
Event
BOBR
Crate 2
Crate 3
Termination

## Slide 36: Additional information
The modified firmware, at this time, cannot affect the total reliability of the system but it can affect its availability, i.e. false dumps.
Measurement or threshold values will not be modified at any time, i.e. no blinding.
Interlocks by channels over threshold values will always be logged.
  For all channels and all integration periods at 1 Hz
  Though, it will be not possible to know when exactly or how many times inside the 1 s the interlock occurred.

## Slide 37: Deployment Plan
At TS1 deploy special BLECS firmware on new crates
  separate crates have been installed
  detector distribution (+use of LICs) has been done
  new functionality only on the two new injection crates
  injection signal via the BOBR (via backplane)
During 2015 maintain two BLECS firmware
  evaluate the need of the inhibit functionality
  discover the optimal settings and safeguards required (deadtime, max. repeat, checks etc.)
  modify further firmware and deploy as necessary (MPP has accepted reduced reliability for these two crates)
At WTS 2015, depending on outcomes,
  decide if functionality is needed
  deploy common firmware to all BLECS modules
  move FESA settings to DB parameters with MCS

## Slide 38: Summary
Most of the system has been completely re-installed
All system modules have been modified and underwent maintenance
Primarily the changes aim to reduce failures and errors
  Optical links
  HV power supply
  Sanity checks
Additional changes aim to improve
  Remote control of cards and crates
  Data collection and distribution
  Improve ability to maintain the installation
Complete re-commission will be necessary
  Rad-source measurements
  MPS checklist

## Slide 39: Thank you

## Slide 40: LHC BLM System Overview