## Slide 1
HTS Magnet Testing

## Slide 2: Over view
Feather Zero & Two reminder
High temperature margin
Testing at high temperature 60K in gas.
Magnetic field Harmonics.
Magnet Quench calculations.
Test requirements

## Slide 3: Feather-M0 and Feather-M2
Five Tesla
HTS Research
Insert-Magnet

## Slide 4: Feather Zero

## Slide 5: Feather Zero parts

## Slide 6: 10 kA Temperature margin
Large temperature margin can we use this for quench detection?
30.5 `K
70 % SS
55.5 K
Temp. sensor CCS

## Slide 7: Position of short sample in coil
16.9 Tesla
5.0 Tesla

## Slide 8: Ic v Test temperature
Chart: Ic v Test Temp Feather 2
| 4.2 | 10 | 20 | 30 | 40 | 50 | 60 | 77
Min. Ic Self Field | 9610 | 8310 | 6190 | 4360 | 2850 | 1750 | 952 | 0
Min. Ic 13T External Par. Field | 8320 | 5630 | 2800 | 1660 | 983 | 250 | 0 | 0
In the post
Chart: Feather Zero Ic. f(temp)
| 4.2 | 10 | 20 | 30 | 40 | 50 | 60 | 77
Min. Ic Self Field | 11200 | 10300 | 8640 | 7020 | 5530 | 4110 | 2870 | 1130
Min. Ic 10T External Par. Field | 11200 | 9380 | 6250 | 3940 | 2550 | 1240 | 57.6 | 0
Ic 80% of SF IC | 8960 | 8240 | 6912 | 5616 | 4424 | 3288 | 2296 | 904
16.9 T

## Slide 9
Feather 2
Iron OD ~ 240mm

## Slide 10: Quench Protection with Copper ICEE Loops
Without Loops
1ms
10ms
95 K
Induced Current After Rampdown
With Loops

## Slide 11: Review of Feather 2 components

## Slide 12: Ic versus transverse stress
No degradation up to 253 MPa
Degradation at higher pressures is irreversible
Chart: Transverse stress [MPa]
| 4.0999999999999996 | 21.4 | 4.0999999999999996 | 30.3 | 12.6 | 39.300000000000011 | 48.4 | 57.6 | 66.900000000000006 | 76 | 85.3 | 94.5 | 103.8 | 113.5 | 122.8 | 132.1 | 141.30000000000001 | 150.69999999999999 | 160 | 169.2 | 178.5 | 187.9 | 197.2 | 206.4 | 215.8 | 225 | 234.3 | 252.9 | 262.10000000000002 | 85.6 | 271.39999999999992 | 85.6 | 280.7 | 85.5 | 299.2 | 85.5 | 308.60000000000002 | 85.6 | 317.8 | 85.5 | 326.89999999999992 | 85.5
 | 2070 | 2066 | 2051 | 2063 | 2061 | 2064 | 2062 | 2076 | 2070 | 2072 | 2061 | 2071 | 2071 | 2061 | 2070 | 2071 | 2060 | 2068 | 2072 | 2078 | 2065 | 2079 | 2079 | 2064 | 2075 | 2075 | 2076 | 2079 | 1991 | 1984 | 1976 | 1975 | 1980 | 1970 | 1989 | 1965 | 1978 | 1969 | 1548 | 1507 | 1529 | 1521
T = 4.2 K
B = 10.5 T
Feather 2
Working range 150 to 200

## Slide 13: Quench detection
Voltage taps (low current)
Temperature sensors (high current)
Pickup coils high current (Feather zero trial)
Acoustics
Thanks to Maxim Marchevsky at LBL

## Slide 14
Erkki Härö
Finite Element Model using anisotropic thermal conductivities
Will we be able to see this in the cable ?
Feather-M0 Quench
Behaviour 4.5K f(current)

## Slide 15
88 ms
202 ms
7.2 s
Quench simulation results for Feather-M0 with different operation temperatures:
4.2 K, 30 K and 77 K.
Operation currents with 4.2 K, 30 K and 77 K were 8960 A, 5616 A and 904 A,
respectively.
Operation current was chosen to be 80 % of the short sample Ic value.
20 ms
50 ms
2.5 s
4.2K
30K
77K
If we can see it in the Roebel cable and coil
Feather M0 Quench
Behaviour f(test temp)

## Slide 16: Non-Linear Current Sharing Model
Current sharing refers to the current distribution between the matrix material and the superconductor provides Vnl(I)
In the model three options are available:
  Pure power law (not useable for quench analysis because only valid below or just above Tcs)
  Linear transition between Tcs and Tc
  Superconducting power law element in parallel with resistive (matrix) element (implicit equation)
Angular Dependent Critical Current
Current Sharing
Isc
Isc
Inc
Inc
T [K]
I [A]

## Slide 17: Feather–M2 Harmonics
Due to limitations of the BemFem no Iron pole (yet)
Simultaneous ramp of Fresca2 and insert over 20 minutes (1200 s)
Top deck generates more shielding currents due to field angle
Harmonics at 2/3 aperture less than 10 units
Keeping alignment during ramping could reduce the field error by a great deal
Top
Deck
Central
Deck
Central
Deck
Top
Deck

## Slide 18: Quench Detection Race
Voltage Taps
Pick up Coils
Temp sensors
Field Sensors
Acoustics
FPGA -DAQ

## Slide 19: Test station spec.
  Feather‘s Current ~10 kA
  Dump resistor.
  10 kA semiconductor Fast switch 1 or 2 mS
  Variable test temperature ~80 to 4K in steps! (temp stability, current leads opp.,
  Instrumentation. ( 224 channel FBGA DAQ)
  Temp sensors, voltage taps, pick up coils.
  Feather 2 dynamic field measurements. 40mm
  Feather 2 in Iron yoke od ~ 240 mm
  Feather 0 in Iron yoke od ~ 120mm

## Slide 20: HTS Quench DAQ
7 modules: 16 differential analog inputs each
+/- 200 mV till +/- 10 V input range
16 bits
7.8 kS/s per channel
1 module: 32 digital outputs
5 V TTL
7 µs response time
Also available: high speed module
4 differential analog inputs
16 bits
1 MS/s, simultaneous sampling
Processor:
667 MHz dual-core ARM
512 MB RAM
1 GB storage
NI Linux Real-Time OS
FPGA:
Xilinx Artix-7
2 M cells
A similar system is used by the EL group to capture voltage transients on the electrical network caused by EDF switching, thunderstorms and internal load changes.
30 cm
Received at CERN 14 March
NI-CompactRIO
224 total

## Slide 21: Cross section 5 T Stand alone 98 mm outer diameter, to be able to fit inside Fresca-2 without touching. Standalone we may fit a simple magnetic yoke
This is one of the many possible designs we are Thinking about: Placing IRON yoke to Exaggerate the change in field angle to assess angular dependence
Iron OD ~ 240mm

## Slide 22: Possible test ideas after stand alone
KEK D1
due late 2015
Fresca 2
Inner coil test 2015