## Slide 1
DETECTOR TECHNOLOGIES
Lecture 3: Semi-conductors
- Generalities
- Material and types
- Evolution

## Slide 2
Semiconductors : generalities
Solid-States band structures :
Valence band : e– bond atoms together
Conduction band : e– can freely jump from an atom to another
At T ≠ 0 K
Electrons may acquire enough energy
to pass the band gap… Thermal condution

## Slide 3
Semiconductors : generalities
Not the same !
(Thermal excitation + phonons)

## Slide 4
Semiconductors : generalities
Energy loss by a charged particle : Bethe-Bloch
Standard :
Energy loss :
electrons – holes pairs created
(NOT electrons – ions…)
If Field (even natural)
electrons migration
Electrical pulse
Information
Too simple !

## Slide 5
Semiconductors : generalities
One MIP in Silicon at 300°K
Energy loss : dE / dx ≈ 388 Ev/µm
Ionisation Energy : 3.62 eV
e – holes pairs created : 107/µm
For 300 µm : 3.2 10 4 pairs created
Free charge carriers in the same volume :
≈ 4.5 10 9
Signal is lost !
Solution : Depletion of the detector
- Doping
- Blocking contacts
Depletion : removing the maximum possible thermally excitable electrons

## Slide 6
Semiconductors : generalities : the Fano factor
Number of e – h pairs is a statistical process :
Number of e-h pairs : N = E loss / E ionization
If excitations are independants , they obey to a Poisson statistic
with a standard deviation
variance :
Fano factor :
variance / mean of the process (should be 1 for a perfect Poisson distribution)
Si | 0.115
Ge | 0.13
GaAs | 0.10
Diamond | 0.08
Fano factor related to energy resolution :
A Fano factor < 1 means that the energy resolution would be better than
Theoretically expected…

## Slide 7
Semiconductors : generalities : p and n types
dopants :
Boron, Arsenic, Phosphorous, Gallium

## Slide 8
Semiconductors : generalities : p and n types
Typically : doping level for a Silicon Detector : 10 12 atoms / cm 3
Doping us usually done by ion implantation.

## Slide 9
Semiconductors : generalities : junction detectors
A p-n junction is formed when a
single crystal of semiconductor is doped with acceptors on one side and donors on the other

## Slide 10
Semiconductors : generalities : junction detectors

## Slide 11
Semiconductors : generalities : reverse biasing scheme
The p – n zones will be used for contact
and to block (Blocking Contacts) the undesired noise

## Slide 12
Semiconductors : generalities : building
DC Coupling Silicon detector

## Slide 13
Semiconductors : generalities : building
AC Coupling Silicon detector

## Slide 14
Semiconductors : generalities : building
AC coupled Si detectors create 2 electrical circuits :
- Read-out circuit to the amplifier (AC current)
- Biasing circuit (DC current)

## Slide 15
AC Coupling Silicon detector : bias voltage system
Semiconductors : generalities : building

## Slide 16
Most commonly scheme AC + poly S-bias resistor
Semiconductors : Si detectors designs

## Slide 17
Semiconductors : Si detectors designs
CMS design
ATLAS design

## Slide 18
Semiconductors : Si detectors designs

## Slide 19
Semiconductors : Radiation Damage
Two types of radiation damage :
 Bulk (Crystal) damage due to Non Ionizing Energy Loss (NIEL) - displacement damage, built up of crystal defects –
  Change of effective doping concentration (higher depletion voltage, under- depletion) Increase of leakage current (increase of noise, thermal runaway) Increase of charge carrier trapping (loss of charge)
 Surface damage due to Ionizing Energy Loss (IEL) - accumulation of positive in the oxide (SiO2) and the Si/SiO2 interface – affects: interstrip capacitance (noise factor), breakdown behavior, …
Impact on detector performance (depending on detector type and geometry and readout electronics!) Signal/noise ratio is the quantity to watch
 Sensors can fail from radiation damage !

## Slide 20
Semiconductors : Effect of radiations
Loss of collected charges
(new 300 µm Silicon ≈ 24 000 e- for 1 MIP)
Trapping is characterized by an effective trapping time eff for electrons and holes:
where

## Slide 21
Semiconductors : Effect of radiations
Increase of Leakage current

## Slide 22
Semiconductors : Effect of radiations
Change in depletion voltage and type inversion
before inversion
after inversion
n+
p+
n+
p+
Innermost layers
should still work after

## Slide 23
Semiconductors : 2-dimensional detectors
Double Sided Silicon Detectors (DSSD) Not much in use…

## Slide 24
Semiconductors : 2-dimensional detectors
Stereo Modules

## Slide 25
Semiconductors : Performance

## Slide 26
Semiconductors : Pixels Detectors
Pixel sizes :
ATLAS : 50 µm x 400 µm
CMS : 100 µm x 150 µm
ALICE : 50 µm x 425 µm

## Slide 27
Semiconductors : Pixels Detectors
CONNECTION BY BUMP BONDING

## Slide 28
Semiconductors : Pixels Detectors
Pitch : 50 µm
(wire bonding typically 200µm)

## Slide 29
CMS ≈ 215 m2
ATLAS ≈ 61 m2
LHCb ≈ 12.5 m2
ALICE ≈ 1.5 m2
CDF ≈ 3.5 m2
NA 11
DELPHI
Semiconductors : Silicon history
CDF

## Slide 30
Semiconductors : CMS Silicon Detector

## Slide 31
Semiconductors : CMS Silicon Detector

## Slide 32
Semiconductors : ATLAS Silicon Detector

## Slide 33
Semiconductors : ATLAS Silicon Detector

## Slide 34
Semiconductors : CMS and ATLAS Silicon Detector

## Slide 35
Semiconductors : Challenges and Evolutions
Main Challenge : The LHC at High Luminosity (2024 ?)
More tracks : Occupancy increases - Less resolution
More Flux : Radiation (bulk) damage

## Slide 36
Reduce the Occupancy : Increase the granularity
Mini-strips sensors (reduce lenght from 10 cm to 5 cm)
- Increases the number of channels
- Increases the cost
- Increases the power to be dissipated
Reduce the matérial : Thin Si sensors
- Reduce the Charges Collected
Reduce the number of layers
- Reduce the overall Tracker efficiency
300 µm
150 µm
Semiconductors : Challenges and Evolutions

## Slide 37
Change the material : Oxygenated Silicon
HE detectrors : FZ (Float Zone) Crystal - High resistivity > 3-4 kΩcm
- O2 contens < 50 10 16
New Materials : DOFZ : O2 doped FZ Silicon (Oxydation of wafer at high temperature)
MCZ (Magnetic Czochralki) - Less resistivity ≈ 1.5 kΩcm
- O2 contens > 5 10 17
EPITAXIAL growth : Chemical Vapor Deposition on CZ substrate
Oxygen concentration in DOFZ
EPI layer
CZ substrate
Oxygen concentration in Epitaxial
Semiconductors : Challenges and Evolutions

## Slide 38
24 GeV/c proton irradiation
Standard FZ silicon
  type inversion at ~ 21013 p/cm2
  strong Neff increase at high fluence
Oxygenated FZ (DOFZ)
  type inversion at ~ 21013 p/cm2
  reduced Neff increase at high fluence
CZ silicon and MCZ silicon
  no type inversion in the overall fluence range (verified by TCT measurements) (verified for CZ silicon by TCT measurements, preliminary result for MCZ silicon)  donor generation overcompensates acceptor generation in high fluence range
Common to all materials (after hadron irradiation):
  reverse current increase
  increase of trapping (electrons and holes) within ~ 20%
Semiconductors : Challenges and Evolutions

## Slide 39
Semiconductors : Challenges and Evolutions
MAPS (Monolithic Active Sensor) or CMOS (Complementary Metal Oxide Semiconductor)

## Slide 40
Semiconductors : Challenges and Evolutions
3D Silicon Detectors
Manufacturing challenge
Electrodes : dead zones
Efficiency vs fluence

## Slide 41
Semiconductors : diamonds
Diamond detectors
 | CVD Diamond | Si
Z | 6 | 14
Energy Gap | 5,5 eV | 1,21 to 1,1eV
Resistivity | 1013 – 1016 Ωcm | 105 – 106 Ωcm
Breakdown | 107 V/cm | 3.105 V/cm
Mobility (electrons) | 2000 cm2/V/s | 1350 cm2/V/s
Mobility (holes) | 1600 cm2/V/s | 480 cm2/V/s
Displacement Energy (e-) | 43 eV/atom | 13 à 20 eV/atom
Pairs Creation | 13 eV | 3.6 eV
Charge Collection Distance | 250 m | 100 m ?
Mean signal (MIP) | 3600 e- / m | 8900 e- / m
Dielectric Constant | 5.5 | 10 à 12
Thermal Conductivity (W/m·K) | 1600 - 2000 | 150
Diamond is better than Silicon
Does not need any doping
Better radiation hardness
Better thermal conductivity
Better speed (1psec vs 1 nsec)
Light insensitive
Multi-metalization possible
(test and physics)
But :
3 times less signal for MIPs (3.6 / 13)
Difficult to manufacture
Expensive
Diamond is not understood (at the moment)
2 forms :
Polycristalline
Wafer 6 inches
Monocrystalline
max : 4 x 4 mm 2

## Slide 42
Semiconductors : Diamonds
CCD : CHARGE COLLECTION DISTANCE
Energy loss (Bethe-Bloch 1932)
Charge transportation (Hecht 1932)
d = le + lh = (me te +mh th ) E
l : mean drift distance (mean free path of the carrier)
m : mobility
t : lifetime
E : applied electric field
CHARGE COLLECTION DISTANCE :
Si (mono) = 100m
Si (amorphe) =10 µm
Diam = 0(100µm)

## Slide 43
Semiconductors : Diamonds
CCD : MEASUREMENT
Collected charge:
Charge Collection Distance :
Pairs / MIP : 3600 / 100 µ diamant)
1 V/µ
CCD : Measurement of diamond quality (pCVD ou sCVD)
Détecteur de particules :
Nélectrons (MIP) ~ 10 000
Epaisseur (Xo…) ~ 300 µm
CCD ~ 280 µm

## Slide 44
Semiconductors : Diamonds
Diamant ?
Natural diamond: Lots of impurities
Lots of defects
Diamant HPHT (High Pressure – High Temperature) 1940
P > 50 000 bars
T > 2000 °C
Monocrystal
Dimension (few mm2)
Impurities
Diamant CVD (Carbon Vapor Deposition) 1980
Plasma CH4 – H2 (+X…)
P  0,1 bar
T  1000 °K
Slow deposition on substrate
1 – 50 µm / heure

## Slide 45
Semiconductors : Diamonds
CVD polycristallin (pCVD)
Croissance sur substrat indifférent (quasiment)
Processus lent (1 µm /heure)
DIfférenciation des tailles de cristaux
Processus industriel (couches minces)
Substrat
Zone de Croissance
Gros cristaux
CCD OK
1 mm = 1000 heures !!

## Slide 46
Semiconductors : Diamonds
CVD monocristallin (sCVD)
Croissance exclusivement sur monocristal (racine HPHT)
Très faibles dimensions (typique 4 x 4 mm2)
Processus rapide (4 – 7 µm /heure)
Processus peu industrialisé
Cristal
Graine

## Slide 47
Semiconductors : Diamonds
JM Brom LAPP 22 juin 2012
Finishing :
pCVD : Nucleation (small grains) supppresion)
sCVD : Seed suppression
Métalisation (cf bonding)
charactérisation (CCD…)
Cleaning
Re- métalisation
Monocristal CVD
Racine HPHT

## Slide 48
Semiconductors : Diamonds
CVD monocristallin (sCVD)
Croissance exclusivement sur monocristal (racine HPHT)
Très faibles dimensions (typique 4 x 4 mm2)
Processus rapide (4 – 7 µm /heure)
Processus peu industrialisé
Cristal
Graine

## Slide 49
Semiconductors : Diamonds
JM Brom LAPP 22 juin 2012
APPLICATIONS FOR HEP EXPERIMENTS
Beam Conditions Monitors
Beam Loss Monitors
BaBar
CDF
ATLAS – CMS - LHCb
MAPS vertical integration
(Heat Dissipation)
Since 2007
Principe :
Continuous current measurement
with field E ~ 1 V / µ

## Slide 50
Semiconductors : Diamonds
CDF at Fermilab

## Slide 51
Semiconductors : Challenges and Evolutions
Still plenty of questions about diamond:
Some diamond are Schottky diodes (start to be conductive)
Chart: 
| -500 | -400 | -300 | -200 | -100 | 100 | 200 | 300 | 400 | 500
AT2011-19-2 Side 1 | 2.3e-09 | 1.9e-09 | 1.35e-09 | 9.25e-10 | 4.7e-10 | 4.3e-10 | 8.6e-10 | 1.3e-09 | 1.8e-09 | 2.25e-09
AT2011-19-2 Side 2 | 2.2e-09 | 1.9e-09 | 1.4e-09 | 7e-10 | 5e-10 | 5e-10 | 9e-10 | 1.4e-09 | 1.9e-09 | 2.2e-09
I(V) curve : leakage current (2 faces)
from - 500V to + 500V
Chart: 
| -500 | -400 | -300 | -200 | -100 | 100 | 200 | 300 | 400 | 500
MM7-3 Side 1 | 2.7e-09 | 2.5e-09 | 1.5e-09 | 8.8e-10 | 4.7e-10 | 4.5e-10
MM7-3 Side 2 | 2e-09 | 1.5e-09 | 1.1e-09 | 8e-10 | 3.21e-10 | 4.8e-10 | 8.4e-10 | 1.1e-09 | 1.5e-09
On one side, diamond starts to be conductive
at + 100 V
?

## Slide 52
Semiconductors : Challenges and Evolutions
Still plenty of questions about diamond:
Influence of surface finishing
Metal ( Au) by Evap : 14 000 e- Metal (Al) by Plasma : 16 000 e-
Number of electrons measured on the same diamond with 2 different metalization
?