## Slide 1: Cryogenic Current Comparator (CCC) Beam intensity monitor for the Antiproton Decelerator
Miguel Fernandes
International Conference on Accelerator Optimization
07/10/2015

## Slide 2: Presentation outline
Motivation and specifications for new intensity monitor
CCC functioning overview and design considerations
First beam measurements
Current limitations (monitor + cryogenics)
Summary
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 3: Antiproton Decelerator at CERN
AD beam parameters | 
Beta | 0.97 … 0.11
Cycle | ~85 s
N particles | (5 … 1) x 107
Current | (12 … 0.1) μA
International Conference on Accelerator Optimization
Miguel Fernandes
ELENA

## Slide 4: Motivation for new monitor in AD
  DCCT:
  Insufficient resolution (1µA) for the low current regime (low β,N)
  Fast BCTs:
  Limited to bunched phases
  Schottky monitor:
  Un-bunched:
    time resolution of
    accuracy error > 10%
  Bunched:
    time resolution of
    accuracy error of <10%
  Bunch length dependency
  Complex calibration process
Schottky Intensity Measurement
10
20
30
40
50
60
70
80
90
time [s]
Coasting beam
Bunched beam
0
1
2
3
4
N [x 107 pbar]
0.0
0.2
0.4
0.6
1.0
β [c]
0.8
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 5: Specifications for new intensity monitor
Requirements for a new current/intensity monitor:
Current resolution: 10 nA
Intensity resolution: 5 x 105 charges
Bandwidth: DC - 1 kHz
Requirements for the cryostat
“Zero-boil off” using a pulse tube cryocooler as He reliquefier unit
Long term operation
Collaboration partners:
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 6: CCC functioning overview
Magnetic shield:
  Suppresses all field components except azimuthal beam component
Pickup coil:
  Soft ferromagnetic material with high-permeability concentrates flux
Flux transformer:
  Couples magnetic flux (down to DC) to SQUID
SQUID + Electronic readout:
  Superconducting QUantum Interference Devices
  Measures the magnetic field induced in the SQUID’s input coil
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 7: Monitor and cryogenic system
CCC magnetic shield and ferromagnetic core fabricated by collaboration partners
New coupling circuit was designed using commercially available SQUID system
New cryostat was designed and fabricated
  Low-level of mechanical vibration
  Low heat in leak
  Capacity for 50 liters liquid Helium
  Diameter of ~ 0.8m
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 8: SQUID devices
Superconducting/cold parts
SQUID’s are very sensitive magnetometers
But periodic transfer function limits its dynamic range
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 9: SQUID devices
Constant Working Point
SQUID’s are very sensitive magnetometers
But periodic transfer function limits its dynamic range
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 10: SQUID devices
SQUID’s are very sensitive magnetometers
But periodic transfer function limits its dynamic range
Necessary to limit slew-rate of coupled signal to avoid flux-jumps
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 11: SQUID devices
SQUID’s are very sensitive magnetometers
But periodic transfer function limits its dynamic range
Necessary to limit slew-rate of coupled signal to avoid flux-jumps
    AD injection: !!
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 12: Limiting slew-rate of input signal
Implement signal filtering to decrease slew-rate
RC filter in coupling circuit
RF-bypass in ceramic at beam pipe
Cryostat vaccum vessel
Cryostat thermal shield
Ibeam
Imirror
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 13: Laboratory measurement
RF-bypass (CRF = 100 µF)
RF-bypass + Coupling Circuit (CRF = 100 µF)
Laboratory measurement (CRF = 125 µF)
Time response to a current signal identical to AD beam injection
Simulated signal using CST Wakefield simulation and linear system treatment
Laboratory measurement of wire carrying current through beam pipe
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 14: Measurements with no beam
50Hz
150Hz
250Hz
Frequency-domain
Magnetic cycle [a.u.]
SQUID measurement [V]
Time-domain signals
σ(Ibeam) = 219 nA
Current resolution is mainly limited by noise/perturbations at 50Hz + harmonics
Offset jump around moment of injection due to discharge of bunch rotation cavities
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 15: Beam current measurements
First measurement during beam setup
After signal filtering it’s possible to obtain resolution ~30 nA
Good example of how this measurement can speed up setup of AD beam (currently takes ~3 weeks)
Raw signal (Ts = 0.4ms)
Filtered signal (Navg = 100)
σ(Ibeam) = 280 nA
σ(Ibeam) = 27 nA
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 16: Beam intensity measurements
Nominal injection:
Ninj = 3.7 x 107
σ(N) = 4.8 x 104
CCC
SCHOTTKY
Momentum
σ(N) = 4.3 x 105
Coasting beam
Bunched beam
Intensity obtained by normalizing beam current with velocity
Noise of current measurement is amplified for low-β.
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 17: Beam intensity measurements
Low intensity injection:
Ninj = 0.61 x 107
CCC
SCHOTTKY
Intensity obtained by normalizing beam current with velocity
Noise of current measurement is amplified for low-β.
Momentum
Ibeam = 139 nA
Coasting beam
Bunched beam
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 18: Performance limitations
RF-bypass
Needed for reducing signal slew-rate
opens path for perturbation currents flowing in beam pipe
Cryostat vaccum vessel
Cryostat thermal shield
Could be responsible for:
Noise at 50Hz+harmonics
Offset jump due to discharge of rotation cavities
Ibeam
Ipert.
Imirror
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 19: Intensity measurement limitations
RF-bypass
Needed for reducing signal slew-rate
opens path for perturbation currents flowing in beam pipe
Cryostat vaccum vessel
Cryostat thermal shield
Could be responsible for:
Noise at 50Hz+harmonics
Offset jump due to discharge of rotation cavities
Proposed solution:
Move RF-bypass from ceramic gap to gap in thermal shield
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 20: Cryogenic performance limitations
  Cryogenic system is not able to keep a constant level of liquid Helium
  Current investigations indicate this is due to excessive heat-load in thermal shield of: 12.1 W ( while design value is 7 W)
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 21: Summary
CCC monitor adapted to AD-ring was implemented and installed, and first beam measurements were taken
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 22: Summary
CCC monitor adapted to AD-ring was implemented and installed, and first beam measurements were taken
Measurement of beam current with 30nA resolution (after filtering) was demonstrated
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 23: Summary
CCC monitor adapted to AD-ring was implemented and installed, and first beam measurements were taken
Measurement of beam current with 30nA resolution (after filtering) was demonstrated
First time ever CCC monitor is used on a synchrotron with a “fast” bunched beam
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 24: Summary
CCC monitor adapted to AD-ring was implemented and installed, and first beam measurements were taken
Measurement of beam current with 30nA resolution (after filtering) was demonstrated
First time ever CCC monitor is used on a synchrotron with a “fast” bunched beam
Observed resolution limitations should be due to perturbation currents flowing in the beam pipe
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 25: Summary
CCC monitor adapted to AD-ring was implemented and installed, and first beam measurements were taken
Measurement of beam current with 30nA resolution (after filtering) was demonstrated
First time ever CCC monitor is used on a synchrotron with a “fast” bunched beam
Observed resolution limitations should be due to perturbation currents flowing in the beam pipe
Cryostat needs to be periodically refilled due to excessive heat-load in thermal radiation shield
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 26: Summary
CCC monitor adapted to AD-ring was implemented and installed, and first beam measurements were taken
Measurement of beam current with 30nA resolution (after filtering) was demonstrated
First time ever CCC monitor is used on a synchrotron with a “fast” bunched beam
Observed resolution limitations should be due to perturbation currents flowing in the beam pipe
Cryostat needs to be periodically refilled due to excessive heat-load in thermal radiation shield
Corrective measures for these limitations are to be implemented during next year end technical shutdown
International Conference on Accelerator Optimization
Miguel Fernandes

## Slide 27: Acknowledgements
J. Tan; T. Koettig; A. Lees; J. Brachet; D. Lombard; P. Odier;
J. Belleman; E. Oponowicz; M. Ludwig
C. Welsch
F. Kurian; T. Schiwckert; T. Sieber; H. Reeg
R. Geithner; R. Neubert
Thank you !

## Slide 28: EXTRA SLIDES

## Slide 29: Magnetic shielding
Number of meanders is the dominant factor to total attenuation
Magnetic field:
  Earth: ;
  Signal:
Coupling strength to magnetic core of magnetic field from beam is much higher than for other modes
10 meander
0.5mm gap
-106 dB
14 meander
1mm gap
-127 dB
14 meander
0.5mm gap
-133 dB

## Slide 30: Stability limits
Ibeam
w
Working region for our system
Allowed flux jump rate: <1/hour
FLL – Flux-Locked Loop

## Slide 31: BCCCA functioning overview
Cryostat vaccum vessel
Cryostat thermal shield
Ipert.

## Slide 32: Frequency response
─ R = 0.50Ω; C = 1µF
─ R = 0.50Ω; C = 10µF
─ R = 0.23Ω; C = 1µF
─ R = 0.23Ω; C = 10µF
+ measurement with
(R = 0.23Ω; C = 10µF)
Coupling circuit theoretical response
Coupling circuit + cryostat measurement

## Slide 33: Calibration