## Slide 1
RF cavity induced sensitivity limitations on Beam Loss Monitors
CLIC / CTF3 experimental verification meeting
09 March 2015
M. Kastriotou

## Slide 2
2 linacs (~24 km each)
~70000 accelerating structures per linac!
Compact Linear Collider (CLIC)

## Slide 3
Radiation detectors positioned around the machine, detecting secondary showers generated when particles hit the vacuum chamber (Beam losses)
A system based on Ionization Chambers satisfies the requirements of Machine Protection –(CLIC Conceptual Design Report)
  But requires a large number of BLMs ( > 40000)
Optical fibre based BLMs (OBLMs)
  Pure Silica core multimode fibres as Cherenkov radiators. Both ends of fibre coupled to photon-sensors
  Potential of covering full beam line in a cost-effective way.
Fast BLMs (diagnostics oriented):
  Diamonds (1 ns): LHC for bunch by bunch beam loss monitoring
Beam Loss Monitor technology
Photon sensor
Photon sensor
c
2c/3
2c/3

## Slide 4
Normal conducting
24 or 26 discs
Copper
100 MV/m
Max RF Breakdown rate 3∙10-7 BD/(pulse∙m)
Image: Helsinki Institute of Physics
CLIC accelerating structure

## Slide 5: Dark current
Electron field emission [1]:
  The emission of electrons induced by an electrostatic field
  Emission current determined by E
  Captured by the electric field → accelerated!
[1] Dark current simulation for the CLIC T18 high gradient structure - Z.Li et al., WE5PFP046
[2] An investigation on the field emitted electrons in travelling wave accelerating structures – G. Bienvenu et al., NIM. A320 (1992) 1-8
Capture field for LEP Injector Linac (LIL) type structure (10.5 GHz)[2]
63 MV/m
105 MV/m

## Slide 6: RF Breakdown (CLIC CDR)
Complex phenomenon, beginning to be relatively understood
Sites of the structure with enhanced electron field emission
  Contaminants
  Geometric field enhancement features
  Crystal defects
Ionisation cascade → Electric arc → plasma formation
Plasma expands transversely
Electrical current of hundreds of amperes → magnetic fields which deflect the beam!
Courtesy of : Anité Pérez Fontenla

## Slide 7
Installation at the dogleg experiment
Localised
  pCVD diamond
  40 dB Broadband Current Amplifier
OBLM – optimised for dark current measurements
    Ø 900 μm fibre + SiPM
  Custom-made low-pass voltage filter and crate modules
OBLM
Diamond
125 cm
2.5 cm

## Slide 8
Students' meeting 09 March 2015
8
Experimental data : Dark current

## Slide 9
Experimental data : RF Breakdown
Less statistics

## Slide 10
10
Calculations (I)
Estimation of photons giving signal (downstream of the fiber only) taking into account:
Readout circuit design
SiPM Gain for the given HV
RF
of SiPM

## Slide 11
11
Calculations (II)
Estimation of photons detected by the SiPM per electron crossing the fiber taking into account:
Light attenuation in the optical fiber
Angular dependency of photon yield
Wavelength distribution of photon yield (~1/λ2)
MPPC efficiency dependence on wavelength
θ
e-
Distribution of photons detected per relativistic electron crossing the fiber

## Slide 12
Electrons detected (dark current)
BLM signal higher with higher input power → detection of dark current losses
Photon equivalent of the signal
Electrons crossing the fiber

## Slide 13: Input Power and BLM signal history (dark current)

## Slide 14
Electrons detected (RF Breakdowns)
BLM signal higher with higher input power → detection of dark current losses
Photon equivalent of the signal
Electrons crossing the fiber

## Slide 15: Input Power and BLM signal history (RF Breakdown)

## Slide 16: FLUKA simulations: Geometry

## Slide 17: Simulations performance
Disc 20
100 MeV/m →1 MeV/cm
20 MeV
θ
Case 1
15 MeV
Case 2

## Slide 18: Electron fluence on the fiber per electron lost (top fiber)
Top fiber
Energy (GeV)
Electrons fluence per lost particle

## Slide 19: Electron fluence for 20 MeV loss
Disk 24

## Slide 20
Conclusions
OBLM system sensitive to dark current losses, without beam presence
  Higher Power → Higher energy gradient → Higher energy dark current → Higher Signals
  Sensitivity limitations
OBLM system sensitive to RF Breakdown losses, without beam presence
  Sensitivity limitations
RF cavity induced background losses results in sensitivity limitations in Beam Loss Monitors
BUT ALSO
Beam Loss Monitors could also be utilised as a Beam diagnostics tool

## Slide 21
Same calculation for data with Beam
Diamond detector results (new position!)
New data with the OBLM (no feedback resistor → ability to monitor higher signals without saturation)
More refined simulations
Thank you!
Next Steps

## Slide 22
Backup slides

## Slide 23
Students' meeting, 09 March 2015
23
Calculations

## Slide 24
incoming
transmitted
incoming
transmitted
reflected
reflected
[1] Advances in the understanding of the physical processes of vaccuum breakdown – W.Wuensch, CLIC-Note-1025