## Slide 1: A new transparent XY-MicroMegas neutron beam profiler: first beam profiles at n_TOF
M. Diakaki1, E. Berthoumieux1, T. Papaevangelou1, F. Gunsing1,2, E. Dupont1,
M. Kebbiri1, S. Anvar1, P. Sizun1, E. Monmarthe1,
D. Desforge1, D. Jourde1, M. Sabarte-Gilarte2,
T. Geralis3, T. Daphni4, E. Ferrer-Ribas1, L. Tassan-Got5, L. Audouin5,
J. Heyse6, P. Schillebeeckx6, C. Paradela6, F. Belloni6
1 IRFU / CEA-Saclay, France.
2 CERN, Geneva, Switzerland.
3 INPP / NCSR « Demokritos », Athens, Greece.
4 University of Zaragoza, Spain.
5 IPN, Université d’ Orsay, France.
6 IRMM, Geel, Belgium.

## Slide 2: outline
XYMGAS with segmented mesh + anode development.
Application: quasi-online neutron beam profiler + flux monitor.
First profiles at the n_TOF facility at CERN : Analysis and Results

## Slide 3: Segmented MESH microbulk development
anode strips
mesh
strips
« A low mass microbulk with real XY structure »,Th. Geralis, RD51 Common Fund Project
Mass minimisation
(5um Cu - 50um kapton - 5um Cu).
Large surface detectors.
High radiopurity.
mesh
anode
Drift region
Segmented mesh microbulk:
1) No extra layers
2) Production simplification
3) Real XY structure

## Slide 4: Prototype evolution
First batch:
  Problems during etching
  due to holes topology.
  Many strips in short circuit.
Third batch
  Holes  60/50 μm
  Gaps reduced to 35 μm
  Energy resolution OK!
Second batch
  Etching OK with the new
  topology
  All detectors working
  Bad energy resolution due to
  large gaps (~150 μm)
Drift HV
The first detectors produced:
58 x 59 strips on a 6 x 6cm2 area (1mm thickness)
Mesh hole:~ 60μm / Pitch: 100 μm.
Notes: Challenge: etching of the mesh strips without completely removing the material in between the mesh strips.

## Slide 5: NEUTRON beam MONITOR + profiler
Accurate neutron cross section measurements require:
Neutron fluence/Beam interception factor
    Number/fraction of neutrons hitting the area covered by the sample.
Shape of the beam profile
Beam optics misalignment => Beam fluence variations.
For non-monoenergetic neutron sources:
=>Dependence of profile on the neutron energy
pulsed
neutron source
TOF  neutron energy
sample
Neutron production point
Neutron interaction point
Neutron flight path
sample
sample
n_TOF facility (CERN)
(thermal-GeV)
GELINA (IRMM)
(1meV-20MeV)
NFS (GANIL)
Requirements:
Quasi-online neutron flux + beam profiler as well
Minimal perturbation of the neutron beam / Minimal induced background
Stay permanently in the beam

## Slide 6: Principle of NEUTRON BEAM profileR + MONITOR OPERATION
The longest electron
drift time corresponds
to the point of the
interaction of the neutron
with the target.
The neutron flux will be simultaneously extracted from the SUM SIGNAL.
mesh
anode
neutron beam
consecutive strips
Strips 35/36/37/38/39/40
Strips 45/44/43/42

## Slide 7: ELECTRONICS
*GET, General electronics for TPC, ANR proposal / GET-QA-000-0005, AGET Data Sheet.
Challenge:
No global trigger signal => AGET electronics* + Reduced CoBo configuration
Self triggering mode / timing difference between strips.
64 analog channels /chip.
Auto trigger: discriminator and threshold
Multiplicity signal: analog OR of 6discriminators
Address of the hitted channels
SCA readout mode (all/hitted/selected channels)
Max sampling rate: 100 MHz.
16 peaking time values: 50 ns-1us.
4 charge ranges/channel: 120fC/ 240fC/
1pC/ 10 pC.

## Slide 8: PREVIOUS front-end CARDS SETUP
Front end electronics: protection card, able to
take the signal from the strips (mesh+anode)
distribute the HV (mesh)
take the sum signal (recorded with conventional electronics).
anode
mesh
Problems: missing strips, wrong routing, noise from low voltage supply module
AGET chips
Acquisition

## Slide 9: NEW FRONT END CARDS
MESH STRIPS
ANODE STRIPS
AGET CHIPS
Strip signals+Sum signal
HV splitting + strip signals
DETECTOR
Protection Diodes
Protection Diodes
Sum signal to preamp+n_TOF DAQ
HV in

## Slide 10: Detector performance WITH Xrays
No stripped microbulk
at good transparency: ~11%.
(P= 1 atm)
Ar escape
Peak
Resolution 13.5%
If all strips shortcutted:
strip no
Peak position
Gain variations within 7-8%
(not at the edges=>Detector issue)
X-ray source : 57Fe
Resolution 15%
Notes: Dependence of the resolution on the
threshold applied (Voltage threshold+
multiplicity threshold)=>
multiplicity<3 problem.

## Slide 11: The n_tof facility (cern)
Neutron beam characteristics:
1. White neutron beam (0.025 eV-1 GeV).
2. Neutron energy defined by the Time-Of-Flight.
3. High instantaneous flux.
4. 7ns width proton pulse every > 1s
5. Excellent energy resolution
PS Protons
(20 GeV/c)
EAR-2
~20 m flight path
~200 m flight path

## Slide 12: MOUNTING THE DETECTOR in ear-2 (1)
Mesh strips
anode
strips
Proper grounding + shielding essential.
6LiF target,
91.8μg/cm2
(EXTP 2013-104-01)

## Slide 13
AsAd card very close to
beam.
Vm=310 V / Vd =750 V.
Through n_TOF DAQ:
Sum signal
(neutron flux)
Trigger signal
(with conventional preamps)
MOUNTING THE DETECTOR in ear-2 (2)
AsAd card
Mesh strips (X)
Anode
Strips (Y)
Neutron
beam
SUM signal

## Slide 14
Initial pulses
Smoothed pulses
Derivative pulses
MESH
ANODE
SIGNAL ANALYSIS
Peak useful parameters are stored (Amplitude, Peak position, TOT etc)
Event useful information is stored (Time of event, multiplicities etc)
CONDITION:
If both mesh
and anode had
at least one strip
with signal.

## Slide 15: SOME RESULTS
Interval time between
consecutive events:
gamma-flash
(multiplicity)x(10us)
2) Total Amplitude histogram (sum of strip amplitudes for each event):
Good energy resolution,
(Similar for mesh+anode)
Pile-up.
Low energy
deposition
events =>
to be cleaned.
ANODE
MESH

## Slide 16: EVENT SELECTION CRITERIA FOR BEAM PROFILE (1)
From each event + for each dimension we calculate the time difference between first and last strip hit (Δt)
1
2
3
4
Δt
Criterion 2: 0.5 < tot amp ratio<1.5
Final total amplitude hist much cleaner:
mesh
Δt for mesh events
Criterion 1:
Δt <= (drift space)/(drift velocity)

## Slide 17: EVENT SELECTION CRITERIA FOR BEAM PROFILE (2)
Criterion 3: strips hit have to be
consecutive=> pile-up events
rejected
Point of interaction (strip X,strip Y)
1 2 3 4
Criterion 3: Only alphas

## Slide 18: Towards the Beam profile
1) Taking the last strip that gave signal:
2) Particle track fitting:
Δt
strip no
Better spatial resolution

## Slide 19: OUTLOOK –FUTURE PERSPECTIVES
XYMGAS neutron beam profiler in operational mode.
The electronics were improved / Dead time
Will stay permanently in-beam at n_TOF.
Further improvements in the design of the detector for the next production,
based on the tests performed is foreseen.
Complete characterisation of the detector at the nuclear reactor Orphee,
CEA-Saclay (spatial resolution etc).
Challenging physics measurements (neutron induced charged particle reactions,
angular distributions etc) will be investigated with this detector.
…………...Thank you…………….

## Slide 20: EXTRA SLIDES

## Slide 21: Reconstructed mask image x-rays
Micromesh
1 cm
X-ray
Cu mask

## Slide 22
1) Anode square pads connected through 2 extra layers.
2) Anode pads/strips connected through 1 extra layer.
Previous 2D ubulks