## Slide 1: Status of HIE-ISOLDE Commissioning
Walter Venturini Delsolaro
on behalf of the HIE ISOLDE Commissioning team
ISCC meeting, 30 June 2015

## Slide 2: Outline
HIE ISOLDE roadmap for 2015
HEBT hardware commissioning
  XT00
  XT01
  XT02
Status of CM1 commissioning

## Slide 3: Components to be commissioned
XT01
XT02
XT00
CM1
Cryogenics distribution
Cryogenics facilities

## Slide 4: HIE ISOLDE roadmap in 2015

## Slide 5: HW commissioning of XT00 and XT01
3 Steerers still to commission inside tunnel (will be done together with XT02)
XT02 campaign starting end of this week

## Slide 6: IR pictures from 13 hrs heat run

## Slide 7: CM1
QP2 (C1.1)
QP3 (C1.2)
QS4 (C1.3)
QS1 (C1.4)
QS3 (C1.5)
Thermal shield top plate
HIE ISOLDE Cryomodule 1
SC solenoid

## Slide 8: CM 1 main commissioning steps
Interlock tests
Slow pump down
RF, Instrumentation, ELQA tests before cool down
Low Level RF tests
Cool down
RF conditioning above Tc
RF tests at 4.5 K
SC solenoid test
Survey and Alignment
Heat load measurements
Thermal cycles

## Slide 9: CM1 test planning
Ongoing: RF tests at 4.5 K
Original planning: 9 weeks + 2 weeks contingency
This week: access for installation of LDB1 and SC solenoid test preparation
~ 1 week behind schedule after 5 weeks from start

## Slide 10: Cryogenics synoptic

## Slide 11: Constraints on temperature gradients during cool down
300K
70K
ΔTINLET-OUTLET > 50K
L. Valdarno
 | ΔTmax [K]
Thermal Shield | 50
Supporting Frame | 40
Frame and Helium reservoir | 100

## Slide 12: Cool down history (thermal shield)

## Slide 13: Cool down history (cavities)

## Slide 14: Vacuum
J. Somoza & A. Harrison
Today: < 10-10 mbar!

## Slide 15
    Monitoring positions of the active elements during cool down of thermal shield and supporting frame
1 mm
G. Kautzmann
All thermal contractions of cool-down have taken place
No cold leaks!
Active elements ended up ~ 1 mm above the beam line

## Slide 16: RF systems
RF systems worked well from day 1:
Controls operational for RF veto, power, couplers and tuners
LLRF: Cavities start up and power on Self Excited Loops
Measurement sequences operational
Calibrations checks ongoing

## Slide 17: Low field multipacting

## Slide 18: Low field multipacting cleared

## Slide 19: High field (1.5 MV/m) multipacting

## Slide 20: Lorentz detuning

## Slide 21: Concerns
CAV5 coupler axis: anomalous friction since before cool down. It was manually set to operational BW, no showstopper, requiring attention
At high forward power, observed drift in resonance frequency and field. Investigation in progress
Vibrations from primary vacuum pump microphonics
LVDT (position sensors) calibration, for cold alignment
Cryogenics: stability of thermal shield temperature and helium pressure.

## Slide 22: Summary (I)
So far on track on the roadmap presented at the previous ISCC
All Hardware is installed (least LDB1)
HEBT “hardware commissioned”: XT00 and XT01.
XT02 will be done in the next two weeks
Cryomodule 1 was successfully cooled down, with no cold leaks found at the end of the thermal contraction phase
CM vacuum is exceptionally good (< 10-10 mbar)
Cold alignment will be required (~ 1 mm down)

## Slide 23: Summary (II)
SCRF cavities: all multipacting levels are conditioned
All cavities tuned close to Linac frequency: tuning range OK.
CAV1-CAV4: no field emission up to 5 MV/m, CAV5 tbc
Preliminary measurement of microphonics very encouraging
Working out precise calibrations
Issue with stability at large (10 Hz) BW: being addressed
Delicate steps ahead: solenoid training, unison powering, cold alignment
Beam commissioning planning and modus operandi to be worked out in details