## Slide 1: Development of thin films for superconducting RF cavities in ASTeC
R. Valizadeh and O.B. Malyshev
on behalf of collaboration team
ASTeC Vacuum Science Group,
STFC Daresbury Laboratory, UK

## Slide 2: Motivation
The aim is to develop the PVD and CVD coating technologies of superconducting materials for RF cavities and apply it on the RF cavities
Objectives:
  a systematic study of
    Deposition parameters
    Film morphology, structure, chemistry
    AC and DC superconductivity characteristics such as Tc, Bc, RRR, etc.
    RF evaluation of samples
    Cavity deposition and test

## Slide 3: UHV PVD facility
Bakeable
Load-lock chamber
DC and RF bias
Three planar concentric targets with the variable distance to the substrate: 10-15 cm
Substrate rotation
Ion beam assist,
20  Ts  950 ºC
Differential RGA pumping to analyse the sputter gas.
HiPIMS and Pulsed DC,
Notes: Substrate can be set to continuous rotation and uniformly heated from room temperature to 950 ⁰C.

## Slide 4: PECVD/ALD deposition
Base pressure:
  1.5 x 10-5 mbar at 120 ºC
Gas flows:
  Argon, Max 5 l/min, 200 sccm MFC
  Hydrogen, Max 1 l/min, 100 sccm MFC
Heater tested up to 700 ºC (could go 950 ºC)

## Slide 5: ALD rig
ALD valves:
tested with
Arduino control,
switching <1 ms
Plasma waveguide
Reactor chamber
Sample holder
SAES gas purifiers
Notes: Home made circuit to interface Arduino and valves
Other metal-organic precursor will be used for NbN

## Slide 6: Multi-Probe UHV XPS, AES, AFM, STM, LEED and ISS

## Slide 7: New surface analysis rigs: Auger microscope
Spatial distribution of the elements (Auger maps or analysis in lines, points and areas) Secondary electron images with spatial resolution down to 200 nm.
XPS facilities with energy resolution ~ 0.9 eV.
AES and XPS depth profiling.

## Slide 8: New surface analysis rigs: SNMS
Secondary neutral mass spectrometry
Isotope elemental depth profiling.
Good sensitivity for hydrogen using Hydrides

## Slide 9: XPS analysis
C 1s
O 1s
Nb 3d5/2
Nb 3d3/2
Nb 3p3/2
Nb 3p1/2
Nb 3s
As received
Ar+ bombarded
x 10
2
10
20
30
40
50
60
70
80
CPS
600
500
400
300
200
100
0
Binding Energy (eV)

## Slide 10: Film Morphology: XRD
Nb grain sizes within our films:
  18 -73 nm
This is similar in size to the
grains produced in other studies.
The films are polycrystalline
and highly textured with the prefered direction of growth depending on the substrate orientation and deposition condition
RRR values shows no dependence to a preferred orientation
Notes: Typical XRD spectrum
Preffered orientation

## Slide 11: Film Morphology: EBSD
Only one sample has been analysed using EBSD
EBSD data shows grains larger than 18 nm present in the sample, the largest of which are of the order 250 nm across their longest axis. The larger grains are surrounded by a matrix of smaller grains of a similar size to those described by XRD.
Notes: Typical XRD spectrum

## Slide 12: Film Morphology: SEM
samples deposited with a 50-V bias
samples deposited without a bias
Chart: Power (W)
| 100 | 200 | 400 | 500 | 600
0 V | 2.29 | 4.77 | 9.56 | 15.9 | 18.5
50 V | 1.16 | 1.4 | 5.66 | 6.7 | 15
Notes: Columnar: Low grain size and low RRR
Dense: Larger grain and larger RRR

## Slide 13: Superconductivity evaluation
RRR measurements
  have been performed using a purpose built cryostat housing a four point probe.
DC SQUID measurements
  The measurement gives values for both the first and second critical fields, Hc1 and Hc2.

## Slide 14: Superconductivity evaluation: DC SQUID
5 samples studied for B perpendicular to the sample surface
2 samples studied for B parallel to the sample surface
A typical DC magnetic susceptibility measurement with the sample parallel to the magnetic field.
Some results are quite unexpectable
Notes: In all cases there is an increased RRR when a biased substrate has been used relative to unbiased although increasing the bias further does not always result with increasing RRR

## Slide 15
Hc1
Hc1
Hc1
Hc2
Hc1
Hc1
Hc2
Hc2
Hc2
Hc2
Hsmp
Hsmp
Hsmp
Hsmp
Hsmp
Superconductivity evaluation: DC SQUID

## Slide 16: RRR vs Hc2 fim deposited on Silicon
HC2 increases with RRR at T = 6 K

## Slide 17: Pill BOX Cavity: Aluminium mock up RF cavity test
Mock up 3.9 GHz pill box cavity has been fabricated to validate the simulation results obtained with Microwave Studio.
Samples:
  a 100-mm diam. Nb disk
  a 100-mm diam. copper disk with thin films of Nb deposited on the surface.

## Slide 18: LHe CAVITY CRYOSTAT
Liquid Helium Chamber
Liquid Nitrogen Chamber
Outer Vacuum Chamber
Radiation Shield
Sample Chamber
Cradle Assembly

## Slide 19: Magnetic field penetration measurement experiment suggseted by A.Gurevich
MAGNET
VAPOUR COOLED
MAGNET CURRENT LEADS
SAMPLE TUBE HOLDER
SAMPLE TUBE
RADIATION BAFFLE SET
LEVEL PROBLE
SAMPLE TUBE POSITIONER
FIELD PROBES 1 AND 2
WIRING PORT 1
WIRING PORT 2
GHE EXHAUST
LHE TRANSFER PORT
SAMPLE TUBE SLIDER
MAGNET SLIDER
4.2K
LHE CRYOSTAT
CRYOSTAT TOP PLATE
MAGNET SUPPORT
SAMPLE SUPPORT
1T Mini magnet

## Slide 20: PVD SUMMARY of RESULTS
Largest grains of 73 nm deposited by HiPIMS on MgO
Increasing gain size expected to increase Hc1
Hc1 has ranged from 0.1 to 0.18 T on Cu and Si substrate
Hc2 for films deposited by HiPIMS has been approx. 0.4 T and lower than for DC sputtered films
Low Hc2 expected of large grain films and is approaching values expected of bulk Nb (approx. 0.27 T)
More stable pinning for films deposited onto Cu substrate when compared to Si substrate
Films deposited onto Si show large flux jumps
RRR approximately doubles for films deposited onto Cu when compared to Si however no correlation between RRR and Hc2
Films with higher RRR showed fewer flux jumps

## Slide 21: CVD deposition of Nb on Cu
Tc = 9.3K
RRR = 31
Nb(v) Cl
Deposited at 650C

## Slide 22: CVD Nb – Electron diffraction

## Slide 23: SEM and TEM of PECVD of Nb

## Slide 24: Conclusions
Quality of the film (morphology, RRR, Hc2) depends on
  deposition parameters such as
    Substrate temperature,
    Ion/atom arrival ratio,
    Substrate bias,
    Plasma generation at the target
      DC pulsed or DC
      HiPIMS
    Substrate crystallography
    For CVD/ALD precursor and deposition Temperature

## Slide 25: Conclusions (2)
Sample evaluation
  RF pill-box cavity is the best, however
    High Cost of manufacturing and cost of LHe,
    Time consuming
  AC and DC susceptibility:
    Not direct for RF (Hc1 and Hc2) and a cost of LHe
    Quicker
  TC and RRR is an initial of evaluation
    RRR > 10-73, but does not corrollate with RF
    Quick, cheap
  Surface analysis to determine:
    Dense or columnar
    Grain size,
    Composition and impurity
    Defects density

## Slide 26: Future plans
HiPIMS sample deposition for
  Nb, NbN, NbCN, Nb3Ge MgB2
Plasma ALD deposition
    Nb, NbN, NbCN, MgB2
Continuing RRR measurements
AC and DC magnetisation measurements
Magnetic field penetration SCI multilayer layer
RF pill-box test facility
  Testing with a bulk Nb disk (July 2015)
  Sample measurements (August 2015)
3D coating (2016)

## Slide 27: The UK’s SRF collaboration team
Dr. R. Valizadeh, Dr. O.B. Malyshev,
A. Hannah, D.O. Malyshev,
S. Pattalwar, J. Herbert
A. Wheelhouse, P. McIntosh
Dr. G. Stenning
S. Wilde, Dr. B. Chesca
P. Pizzol, Prof. P. Chalker
F. Lockwood-Estrin,
Prof. J. Bradley