## Slide 1: CMS activities: 2015 – end LS2
Austin Ball, CERN.
with thanks to many who provided input.

## Slide 2: Introduction
CMS work in this period falls into 3 categories
Completion of the upgrade started in LS1 (which we refer to as “Phase 1”)
designed for 2 x 1034, 500fb-1
(in aspects needing Technical Sector support 85% completed)
Consolidation and upgrade of detectors to meet HL-LHC performance and timescales
(which we refer to as “Phase 2”)
(includes consolidation, upgrade or end-of-life replacement of associated infrastructure
e.g. detector construction, test & maintenance facilities & services like cooling, electricity)
Maintenance of the existing apparatus and its supporting systems (typically 50% of all tasks)
1

## Slide 3: CMS Upgrades for HL-LHC
Target ~ 30 month duration of LS3 can only be met by:
--> pre-empting large amount of infrastructure & prep work to Run 2/3 and LS2
--> advancing some detector upgrades (endcap m) to LS2 as well
10 year programme,
integrated with:
Ongoing operation,
Maintenance
Phase 1 completion
Budget scale approval
sought at RRB next month
Most detector upgrade
or revision (eg new
Tracker, Endcap Calorimeter)
will happen in LS3
2
Technical Proposal: https://cds.cern.ch/record/2020886?ln=en

## Slide 4: CMS activity 2015-2018
YETS 2015-16: Change in HF electronics (delivery permitting)
Lower HF’s Snagging from LS1: eg YE4+ opening test, RE4+ repair, smoke det.
Open YE4+ Magnet Cryo: valve change on intermediate cryostat (UXC)
(Yoke closed) cold box cleaning, PORS replacement.
Detector & services maintenance (the earlier the better) with yoke closed Detector Cooling: supervision upgrade + sectioning of rack circuits.
EYETS 2016-17: Replace pixel tracker with Phase 1 pixel tracker (CO2 cooled – PH-DT)
Full opening of Install “demonstrator” for additional muon station GE1/1 (GEMs)
Endcaps Source scans of endcap Hadron Calorimeter (check rad damage).
Check options for barrel RPC major revision in LS2
Demonstrator for new yoke opening system
Minimum essential detector/services maintenance (the earlier the better).
Install prototype cooled freewheel thyristor on M1 magnet (N.Area H2)
During Run2: Replace PM54 elevator (outside technical stops)
(2016-2018) Construct surface “battery barn”: for UPS upgrades
YETS 2017-18: Install 2nd UXC crane
Detector closed Install cooled freewheel thyristor for magnet.
3

## Slide 5: EYETS in UXC (16 weeks + lab closure)
Needs: Rigging, Survey, Transport, BP & Vacuum, FSU, Safety Coord at LS1 levels.
Early services maintenance to avoid impeding commissioning
4

## Slide 6: CMS activity LS2
Full opening: Install revised opening system
Many logistic Replace all beampipe except central section & test RP shielding
configurations Upgrade readout of endcap muon stations ME2/1,3/1,4/1 (UXC,SX5) Install new muon station GE1/1 (transport from Prevessin, fibre readout)
Finish Hadron Calorimeter Phase 1 upgrade: HF electronics
HB/HE phototransducers
Possible replacement of HE megatiles (depending on YETS 16-17 meas)
VAX/TAS region prep for HL-LHC revision in LS3
3-year detector maintenance (40-50% of total activity)
Major infrastructure revisions: Underground HVAC control, control room HVAC.
Chilled water (increase capacity)
Water production & distribution control system
Pneumatic control compressed air system
DAQ UPS (1.5MVA) DSS/RSS upgrade?
3-year services maintenance as LS1
Surface buildings Construction/completion of SX5 extension
Complete SCX control room consolidation
5

## Slide 7: LS2 Opening Sequence ~ 12 weeks + 3 contingency
Contingency for possible work on opening system
Needs: Rigging, Survey, Transport, BP & Vacuum, FSU, Safety Coord at LS1 levels.
6

## Slide 8: LS2 Closing Sequence ~ 29 weeks + 5 contingency
Contingency for possible LS3 prep work on TAS/VAX etc
Needs: Rigging, Survey, Transport, BP & Vacuum, FSU, Safety Coord at LS1 levels.
7

## Slide 9: Detector maintenance and upgrade
About 70 weeks available: - comfortable for standard programme
- quite tight if major calorimetry work (HE megatile replacement)
Cycle through barrel wheel access
Limited access to endcaps
Cycle through endcap disk access
Barrel closed
Needs: Rigging,Survey,Transport (incl GE1/1 Prevessin to P5), FSU, Saf. Coord at LS1 levels.
8

## Slide 10: Magnet power and cryogenics
Maintaining solenoid performance at 3.8T over the next 2 decades is imperative.
Effective lifetime prolongation requires early action
LS1 : improvements focussed mainly on the surface plant:
redundancy in the compressor system (hot spares)
powering redundantly from the machine and general services networks
but the coldbox risks have clearly been underestimated (see last 5 mo)
Future programme, up to end LS2 : as presently defined, focuses on:
-simulation and measurement to understand/mitigate the ageing processes
-control system upgrade (LS2), eg auto-restart of hot spare compressor
-eliminating slow discharges due to transient power converter faults
(cooled freewheel thyristor – YETS 16-17 and LS2)
Add more to cover risks over extended lifetime:
-eg reliability/redundancy of coldbox, current leads etc (see next slide)
-redundancy from P5 HL-LHC cryo by LS3 (requires cryo transfer line)
make provision during HL-LHC LS2 works at P5.
9

## Slide 11: Magnet Cryogenics
recent events show that magnet cryo system remains a point of great vulnerability for CMS
(many thanks again to EN,TE and PH for their extraordinary recent efforts)
bottom up risk (re-)analysis for all magnet systems to be done (Jim Strait)
Recent additions to Magnet project: (TE-CMS)
Sept 2015: Replace warm He transfer lines
Construct 2 spare Ads80K tanks, pre-filled
& pre-flushed to remove carbon dust.
YETS: Re-design & replace the
2015-16 primary oil recovery system. (PORS)
Deep clean the Cold Box and thermal screen
Replace damaged Cryo valve in UXC dewar
Proposed addition:
LS2/3:Backup cryo-line from new HL-LHC cryo
Other options:
Possible replacement of CB or its components
Possible construction of spare current leads
10

## Slide 12: Magnet :reliability & minimal on-off cycles
On-off cycles are bad for the magnet eg
(Target: 5/year, no FD. Will exceed in 2015)
RRR predicted to reduce with on-off cycles;
eventually (250 cycles?) will not sustain 3.8T
Measured RRR evolution not as simulated
but recovery due to LS1 warm up reversed.
Phase 2 upgrade (delayed from Phase 1)
“Cooled Freewheel thyristor”
draft tech spec/costing written with TE
Immune to power converter faults (power, cooling)
Field decays with time constant ~25 hours.
Prototype on M1 magnet in 2016,
Implement in CMS during LS2
UPS/diesel
protection of
FWT cooling
vital
Maintaining magnet performance at 3.8T over the next 2 decades is imperative
.
11

## Slide 13: Beampipe
LS1 improvements: physics performance motivated :
Central pipe (Be) reduced to minimum feasible radius (pix layer 1 closer to i/p)
LS2: Three motivations for change
(i) LHC: mandatory replacement of TAS and CT2 to match higher aperture of quads
(ii) ALARA: replace all SS beampipes with Al 2219 alloy (factor ≥ 5 x less activation)
(iii) Possible decreased cone angle (h =5.54) to allow calorimetry extension.
Engineering/simulation jointly with TE-VSC (initially for h =5.54): Order in 2016 !!!
Plan: Change all pipes (ex central) to Al 2219 in LS2 (reduced dose-rates at LS3 kick-off).
Caveats: TAS change in LS2 shown not to be feasible: new schedule challenge for end LS3
Can we maintain current (new) central pipe radius < HL-LHC TAS aperture?
If not, need adapter module at 3.1m compatible with present & future pipes
The beampipe today
h =4.9 cone
12

## Slide 14: Beampipe: existing & high h geometries
Examining feasibility of the high h geometry (CMS +TE-VSC)
main issues: feasibility of manufacture & coating, mechanical & vacuum stability
must be ordered in 2016 for delivery in 2018
13

## Slide 15: P5 Infra: UXC cranes
2nd crane bridge: (based on LS1 experience)
Single crane limits parallel work streams at –z and +z
--> install 2nd bridge with 5 ton hook
-old carriage still able to transfer 20t load thru ± 18m.
--> nacelle linked to crane bridge desirable
(safer than suspended cage- needs permit for each use)
Prelim design (EN-HE) & integration (CMS) study done
(includes visit and analysis by manufacturer)
Aiming for installation in YETS 2017-18 if possible – net benefit in LS2.
otherwise at LS2 start
20t + 5t
5t only
14

## Slide 16: P5 Infra: Cooling system
1.0 MW
0.2 MW
0.85 MW
2.6 MW
Chilled Water 6o
Mixed Water 13o
LS1: SCX -->1 MW for HLT farm - done
Detector cooling has no spare capacity
For HL-LHC:
--> need 1.5 MW more chilled water capacity
0.4MW (Tracker + Pixel)
0.6MW (ECAL + HCAL lower temp)
(include 0.15MW option to surface)
0.5MW (provision for HGC)
Mixed water capacity considered adequate
(especially since/if some dets move to chilled)
Needed near start of LS3 (test/commissioning)
Implement (or set up implementation) in LS2?
key meeting last week with EN-CV to confirm details:
to check : cooling tower capacity adequate?
: no special end-of -lifetime issues to continue through HL-LHC era?
15

## Slide 17: CMS Detector Cooling: present and future
16

## Slide 18: P5 Infra: Electrical system
LS1: 4 x 500kVA detector UPS
done n+1 redundancy (+ 1 spare ready)
1200kVA used
300kVA margin (eg for rad dam)
Pre-LS2: Add 3 x 500kVA for comp farm
n+1 redundancy
needed asap for reliability and
lower mortality fault recovery
Move 80kVA control room system
out of control room.
LS2/3: Go to 5 x 500 kVA detector UPS
(maintain n +1 redundancy)
to supply Phase 2 detectors
Discussions with EN-EL ongoing
To Check: mods to P5 elec system to support Phase 2 infrastructure (cooling)
existing UPS is obsolete and may need total replacement – no provision yet.
may require new “barn” even Pre-LS2.
- already obsolete
- in ventilation building
- battery safety issues
(being reviewed with HSE)
New system
may require
new isolated
“barn”.
CMS specific changes: extension of UPS (glitch immunity/graceful shutdown)
17

## Slide 19: P5 Infra: other common systems
Detector Gas systems (muon detectors):
infrastructure to support new endcap muon systems (GE1/1 and ME0)
end-of-lifetime replacement of general diagnostics & some infrastructure
gas purity analysers, programmable logic controllers the
entire network using pressurized nitrogen for pneumatic gas valve control.
consequences of CERN policy to reduce greenhouse gas emissions 80% by LS3.
recuperate (though not necessarily to re-use) the vented gas (notably RPC’s)
validate and deploy alternative gas mixtures with a lower impact.(GWP)
CMS + PH-DT specialist group
Dry Gas system: (detector environment control)
-LS1: “membrane plant” (dry air/dry N2 locally extracted) + revised dry air flushing
-pre-LS3: 2nd membrane plant of very similar capacity for HGC calorimeter
and/or colder ECAL/HCAL (could EN adopt this technology now?)
Rationalisation of “pneumatic control” compressed air supply (EN)
HVAC: Major revision of control systems planned by EN-CV in LS2
To check - exclusion periods for USC/UXC work. Vent mode with min ΔP to detectors.
any end of lifetime replacements necessary for HL-LHC era?
18

## Slide 20: Test/surface facilities at CERN
Test beam zones, irradiation facilities, cosmic stands, temp cycling etc at CERN:
LS1 : main work was refurbishment of M1 magnet (3 T in H2 line)
For HL-LHC: revision/replacement of ageing infrastructure
904 electronics integration centre:
For HL-LHC: rack/lab-unit cooling system ageing - to be replaced by new system (GS)
provision for detector assembly (eg HGCAL, new muon stations)
may require revision of building infrastructure (GS)
904/892/27/186 detector development labs and assembly facilities
refurbishment of common, CMS-owned or CMS-specific infrastructure
904 R-002 occupancy will become critical (+ needs refurbishment)
Pt5 workshops : mechanical & electro mechanical
machine & work-area replacement, & refurbishment – advice needed (EN)
19

## Slide 21: Surface facilities: P5 site
SX5 extension: (needed by end LS2)
-re-house workshop areas,
-re-house storage for critical tooling
-self-supporting self-propelled gantry
Assuming turnkey solution (GS)
as for building 3594 completed in LS1.
(CMS takes care of fitting-out/oversight)
SX5: extend detector labs & RP zone
into space freed up by re-housing.
-provide space for EB revision & Phase 2
endcap calorimeter final assembly.
-mech. work & services installation (EN)
Caveats: -UPS barn may be reqd
-synergy with HL-LHC Civ. Eng. plan?
-no provision yet for storing ME1/2, ME1/3
removed while endcap cal replaced in LS3
SX5 + extension (& UPS barn?)
20

## Slide 22: SX5 extension: detail
Workshop machine list established : replacement/ refurbishment of existing machines
to conform with latest regulations under study.
Storage volume estimated in detail : about 500m2 needed
21

## Slide 23: HL-LHC works at P5: LS2
Underground works done during LS2.
CMS has to co-exist with (& try to profit from) large scale Civ Eng
Aim to get new buildings
for CMS on this plan before
discussions with French authorities
Very recently revised, all CMS barracks to go
-on-site storage space?.
Also bear in mind endcap replacement in LS3
-passage needed for crawler crane
22

## Slide 24: Conclusion
Period up to end of LS2 is vital for the HL-LHC upgrade of CMS in LS3,
as well as for finishing the upgrades needed to get to LS3
YETS 2015-16 -focused on magnet cryo consolidation
EYETS 2016-17 -focused on pixel tracker replacement, will be very intense and
requires careful choreography
-strong support required from the Technical Departments
YETS 2017-18 -use for infrastructure preparations to be optimised
LS2 -will require support similar to LS1 from Technical Departments
on-detector planning OK for known baseline, but could fill-up easily
Infrastructure/ facilities consolidation/upgrades
at P5 need to be in place by the end of LS2
to be ready for detector elements arriving at P5 before LS3
elsewhere even earlier (eg assembly/testing labs)
23

## Slide 25: Main themes arising from LS1 experience
LS1 was (mostly) a spectacular success…but
project engineering/supervision resources insufficient
from tech spec to completion on site
shortfall in provision for “host lab” responsibilities
expt collaborations will not be able to compensate indefinitely
re-examine definitions/provision with Tech sector & PH
“limits of supply” of tech depts to be re-discussed/re-optimised.
(tech sector + expts + PH)
cost control scrutiny and oversight could/must be improved
experiment budgets (from outside CERN) keenly scrutinised
frame contract system & internal (monopoly) market flawed
(award of contracts still under-weights technical advice)
perhaps the moment to recall from “LS1 days”……
24