## Slide 1: Detection thresholds in LARP Magnets
Quench detection thresholds adjusted according to VSDS data
Automated voltage spike data analysis provides basic spike parameters, amplitude and width as a function of current
What may impact voltage spikes ?
  Filament size in Nb3Sn conductor (largest effect at 4.5K)
  Overall conductor (magnet) length
  Temperature (smaller spikes at 1.9 K) and ramp rate
Voltage spikes are smaller in a “mirror” structure due to single coil
Current dependent detection thresholds at VMTF : large thresholds used at currents 1.0-5.5 kA, and lower thresholds above 6 kA (~0.4-0.5 V) for a voltage difference between the half-coil signals

## Slide 2: RRP 54/61 vs 108/127
Short TQ coils in a “mirror” structure

## Slide 3: RRP 54/61 vs 108/127
Long LQS magnets

## Slide 4: Conductor length
Long LQS coil: “mirror” structure vs quadrupole
Long and short coils in a “mirror” structure

## Slide 5: Voltage Spikes in MQXFS coil
MQXFSM1 and LQM01 coils in a “mirror” structure, HQ02 quadrupole

## Slide 6: Voltage Spikes in MQXFS coil
Width of spike signals

## Slide 7: Detection thresholds in LARP Magnets
Quench detection thresholds adjusted according to VSDS data
Automated voltage spike data analysis provides basic spike parameters, amplitude and width as a function of current
What may impact voltage spikes ?
  Filament size in Nb3Sn conductor (largest effect at 4.5K)
  Overall conductor (magnet) length
  Temperature (smaller spikes at 1.9 K) and ramp rate
Voltage spikes are smaller in a “mirror” structure due to single coil
Current dependent detection thresholds at VMTF : large thresholds used at currents 1.0-5.5 kA, and lower thresholds above 6 kA (~0.4-0.5 V) for a voltage difference between the half-coil signals