## Slide 1: Enabling new trends in molecular imaging Maurizio Conti, Siemens Healthcare, Molecular Imaging
MEDAMI workshop, May 2, 2016

## Slide 2
Enabling new trends: beyond “diagnostic-only” imaging (1)
Enabling new trends: beyond “generic” FDG imaging
Enabling new trends: beyond “static” imaging
Outline (1)
Enabling new trends: beyond “diagnostic-only” imaging (2)

## Slide 3
Enabling promising applications
Enabling better workflow and patient comfort
Outline (2)
Enabling research
Enabling better image quality

## Slide 4
Hoffman Estates (Chicago)
SPECT
Detectors
Knoxville
PET
Clinical appl
Detectors
PETNET HQ
Forchheim (Germany)
CT components
MR components
Software development
Bangalore (India)
Software development
Siemens Healthcare, Molecular Imaging
Oxford (UK)
Clinical appl

## Slide 5
Where is Knoxville?

## Slide 6
Siemens Healthcare Molecular Imaging, Knoxville

## Slide 7
Enabling new trends: beyond “diagnostic-only” imaging (1)
Enabling new trends: beyond “generic” FDG imaging
Enabling new trends: beyond “static” imaging
Enabling new trends: beyond “diagnostic-only” imaging (2)

## Slide 8
Challenges of theranostics and personalized therapy monitoring

## Slide 9
Beyond “diagnostic-only” imaging
    Use imaging to plan therapy, monitor therapy, assess response to therapy.
    Personalize the medical treatment.
Theranostics
mAb imaging
Hadron therapy
Theranostics
a targeting agent which specifically binds to a molecular target on the surface of a cell or tumor
an imaging agent and a therapeutic drug/radioisotope pair
a linker which serves to connect the two entities
Theramonitoring
Gene targeted imaging
(photon) RTherapy
Y90 therapy
Small compound imaging

## Slide 10
Theranostics: small compound imaging
    Matched imaging–therapeutic agents:
    68Ga-DOTA-PSMA 177Lu-DOTA-PSMA
177Lu
    Matched imaging–therapeutic radioisotopes: 124I and 131I, 86Y and 90Y

## Slide 11
Theranostics: mAbs and immunoPET
Monoclonal antibodies (mAb) can target a specific tumor cell surface marker.
They can be used to disease-selective imaging and disease-selective therapy.
Intact mAbs have a long biological time in humans (days), and they are the format of choice for therapy (pharmaceutical or radiotherapy).
The corresponding imaging radioisotopes must have a long half-life

## Slide 12
Number of therapeutic monoclonal antibodies entering clinical study during 1980 to 2006
Theranostics: mAbs and immunoPET

## Slide 13
PET image of Zr89-mAb-U36 in a patient with a tumor of the right tonsil (neck squamous cell carcinoma, lymph node metastasis). Images were obtained within 1 hour (A), at 24 hours (B), at 72 hours (C), and at 120 hours (D)
89Zr-labeled mAbs showed high tumor detection and accurate quantification.
Moreover, PET with 89Zr-labeled mAbs appeared to be able to predict the biodistribution of 90Y- or 177Lu-labeled mAbs
Theranostics: mAbs and immunoPET

## Slide 14
Proton and heavy ion therapy is prone to uncertainties (anatomical changes, mis-positioning, …)
Need to verify correct application of the planned treatment:
Dosimetry
Localization (edges)
Prediction
Measurement
    Compare:
    clinical treatment plan
    Monte Carlo simulation
    PET imaging
Monitoring of therapy: hadron/proton therapy
Courtesy of Heidelberger Ionenstrahl-Therapiezentrum (HIT)

## Slide 15
PET scan characteristics:
long (30 minutes)
low number of counts [105-106], while typical FDG PET is [107-108]
with high random fraction (>80%)
    PET imaging of irradiation with proton or 12C induces β+ emitter (15O, 11C, etc)
    Hadron and proton therapy are used for local treatment of tumors: liver, brain, eyes, prostate.
    Very weak PET signal from nuclei activation inside the tissue.
Courtesy of Heidelberger Ionenstrahl-Therapiezentrum (HIT)
    Off-beam PET (about 5-10 minutes post-therapy)
Monitoring of therapy: hadron/proton therapy

## Slide 16
Courtesy of Heidelberger Ionenstrahl-Therapiezentrum (HIT)
OSEM, 21 subsets,
3 iterations
OSEM+PSF+TOF, 14 subsets, 2 iterations
Reconstruction methods need to be tailored/optimized for challenging applications
image size 200x200, 5 mm Gaussian filter
Monitoring of therapy: hadron/proton therapy

## Slide 17
90Y is used as a therapeutic radioisotope in microspheres for targeted radiation therapy of the liver. Very weak PET signal from e+e- pair creation (BR=30x10-6)
PET scan characteristics:
long (30 minutes)
low number of counts [105-106], FDG PET is [107-108]
with high random fraction (>80%)
high flux of Bremsstrahlung singles in 90Y
LSO
Monitoring of therapy: Y-90 therapy

## Slide 18
Courtesy of University Hospital, Nantes, France
Patient after Y-90 microspheres injection in the liver
(50000 counts, 95% random fraction)
non TOF
TOF
Monitoring of therapy: Y-90 therapy

## Slide 19
Enabling new trends: beyond “diagnostic-only” imaging (1)
Enabling new trends: beyond “generic” FDG imaging
Enabling new trends: beyond “static” imaging
Enabling new trends: beyond “diagnostic-only” imaging (2)

## Slide 20
Challenges of theranostics and personalized therapy monitoring

## Slide 21
    Challenges for PET physics and reconstruction:
    non-pure beta decay -> accurate prompt gammas corrections
    Long positron range -> range deblurring
    Fewer counts -> reconstruction optimization
    Unknown kinetics or biological mechanisms
Challenges of theranostics and personalized therapy monitoring

## Slide 22
Positron Emitting Isotopes
Isotope | Half-life | Branching (b+) in % | Emax (MeV) | Emean (MeV) | Rmax (mm) | Rmean (mm)
11C | 20.4 min | 99.8 | 0.960 | 0.386 | 4.2 | 1.2
13N | 10.0 min | 99.8 | 1.199 | 0.492 | 5.5 | 1.8
15O | 2 min | 99.9 | 1.732 | 0.735 | 8.4 | 3.0
18F | 110 min | 96.9 | 0.634 | 0.250 | 2.4 | 0.6
64Cu | 12.7 hours | 17.5 | 0.653 | 0.278 | 2.5 | 0.7
89Zr | 78.4 hours | 22.7 | 0.902 | 0.396 | 3.8 | 1.3
Isotope | Half-life | Branching (b+) in % | b+ Emax (MeV) | Branching (g) in % | g E (PG) (MeV)
68Ga | 67.8 min | 87.7, 1.2 | 1.899, 0.821 | 3.2 | 1.077
76Br | 16.2 hours | 25.8, 6.3 | 3.382, 0.871 | 74.0, 15.9 | 0.559, 0.657
82Rb | 1.3 min | 81.8, 13.1 | 3.378, 2.601 | 15.1 | 0.777
86Y | 14.7 hours | 11.9, 5.6 | 1.221, 1.545 | 82.5, 32.6 | 1.077. 0.627
124I | 100 hours | 11.7, 10.7 | 1.535, 2.138 | 62.9, 11.2 | 0.602, 1.691
Pure beta emitters
Prompt gamma emitters

## Slide 23
Challenges: prompt gamma corrections
F18 decay scheme
“simplified” Y86 decay scheme
Radial profile of phantom: comparison of different radioisotopes
124I
90Y
68Ga
89Zr,11C, 18F, 64Cu

## Slide 24
Scanner characterization for all radioisotopes
Prompt gamma correction prevents image artifacts
124I
PG

## Slide 25
Quantification with OSEM starts failing below 200k trues
PSF and TOF can overcome this limitation (best results PSF+TOF)
OSEM
OSEM+PSF
OSEM+TOF
OSEM+PSFTOF
Quantification and reproducibility at extremely low counts
Monitoring of therapy: hadron/proton therapy

## Slide 26
Enabling new trends: beyond “generic” FDG imaging
Enabling new trends: beyond “static” imaging
Enabling new trends: beyond “diagnostic-only” imaging (1)
Enabling new trends: beyond “diagnostic-only” imaging (2)

## Slide 27
Tracers used in PET
Radiopharmaceutical | # of Times Scanned | % of Total
Fluorodeoxyglucose (F-18) | 346442 | 90.68%
Rubidium chloride (Rb-82) | 10211 | 2.67%
Fluorethylcholine (F-18) | 5087 | 1.33%
Germanium (Ge-68) | 1790 | 0.47%
Gallium Solution (Ga-68) | 1468 | 0.38%
Fluoro-L-dopa (F-18) | 1264 | 0.33%
DOTA-TATE (Ga-68) | 1241 | 0.32%
Ammonia (N-13) | 1233 | 0.32%
Sodium Fluoride (F-18) | 1111 | 0.29%
Oxygen (O-15) | 938 | 0.25%
Solution (F-18) | 823 | 0.22%
Acetate (C-11) | 806 | 0.21%
Choline (C-11) | 787 | 0.21%
DOTA-TOC (Ga-68) | 726 | 0.19%
Fluorethyltyrosin (F-18) | 723 | 0.19%
Methionine (C-11) | 606 | 0.16%
Solid Solution (Ge-68) | 598 | 0.16%
Crossed Rods (Ge-68) | 550 | 0.14%
PSMA (Ga-68) | 442 | 0.12%
18F-FDG
(fluorodeoxyglucose)
Use of PET tracers, 2013
Main application today: oncology
Measuring Glucose Metabolism
But it is changing…

## Slide 28
Compound is inhibitor of a specific site in PSMA (high affinity):
It bounds strongly and only to PSMA
Low molecular weight (fast uptake)
* Banerjee et al.: “Synthesis and Evaluation of Technetium-99m- and Rhenium-Labeled Inhibitors of the Prostate-Specific Membrane Antigen (PSMA)”J Med Chem 51, 4504-4517, 2008
** Courtesy of Martin Pomper
Beyond FDG: PSMA tracers for prostate cancer

## Slide 29
A.Afshar-Oromieh et al.: “Comparison of PET imaging with a 68Ga-labelled PSMA ligand and 18F-choline-based PET/CT for the diagnosis of recurrent prostate cancer”, EJNMMI 41: 11-20 (2014)
R.C.Mease et al.: “N-[N-[(S)-1,3-Dicarboxypropyl]Carbamoyl]-4-[18F]Fluorobenzyl-LCysteine, [18F]DCFBC: A New Imaging Probe for Prostate Cancer”, Clin. Cancer Res. 14: 3036-3043 (2008)
* Courtesy of Martin Pomper
Ga68-PSMA
F18-PSMA
F18-choline
Ga68-PSMA
F18-PSMA
Beyond FDG: PSMA tracers for prostate cancer

## Slide 30
Beyond FDG: Amyloid imaging for Alzheimer’s disease
C11-PiB
Several candidates to image amyloids, associated with Alzheimer’s disease

## Slide 31
Diabody
55 kDa
Peak 1-2 h
Minibody
80 kDa
Peak 4-6 h
Intact antibody
150 kDa
Peak 50-100 h
Engineering Antibodies
Targeted by mAbs:
Immuno diseases
Specific cancer cells
Beyond FDG: mAbs for specific cancer or immuno diseases

## Slide 32
Enabling new trends: beyond “generic” FDG imaging
Enabling new trends: beyond “static” imaging
Enabling new trends: beyond “diagnostic-only” imaging (1)
Enabling new trends: beyond “diagnostic-only” imaging (2)

## Slide 33
Beyond static SUV imaging
Is there more information beyond the simple static tracer uptake?
Tracers kinetics can bring information on:
  Disease characterization in general
  Disease preferred mechanism
  Disease differentiation between aggressive and indolent
  Drug mechanisms
Kinetics parameters of a model can carry that information
Imaging of the kinetic parameters can carry local information

## Slide 34
Results: patient 1 and 2
F18-choline | Prostate tissue | 
 | Healthy | Tumor*
K1 | 0.14 | 0.25
k3 | 0.007 | 0.02
Ki | 0.02 | 0.12
Beyond static SUV imaging: kinetic models and tumor characterization
Parametric imaging (Ki)
static imaging (SUV)
Courtesy of D. Townsend, CIRC, Singapore
Prostate cancer:
18F-choline

## Slide 35
Beyond static SUV imaging: kinetic models and tumor characterization
Reference is muscle tissue
Prostate cancer:
FACBC
and
PSMA

## Slide 36
Enabling promising applications
Enabling better workflow and patient comfort
Enabling research
Enabling better image quality

## Slide 38: Better image quality via PSF reconstruction
3D-OSEM
3D-OSEM+PSF

## Slide 39
Better image quality and higher sensitivity via TOF reconstruction
non-TOF
TOF
Data Courtesy of University of Tennessee, Knoxville, Tennessee, U.S.A.
T1
T2
x
0

## Slide 40
Non-TOF
TOF
Better image quality and higher detectability via PSF and TOF reconstruction
Kadrmas, Casey, Conti, Jakobi, Lois, Townsend; “Impact of Time-of-Flight on PET Tumor Detection” J Nucl Med 2009; 50:1315–1323

## Slide 41
Better image quality via reconstruction and scanner architectures
TOF reconstruction, large bore and high-capacity table enable studies of a broad patient population
Non-Hodgkin's Lymphoma: 13-year old, 155 cm, 217 kg, 90.3 BMI, 15.6 mCi
Images courtesy of the Ohio State University, USA

## Slide 42
Area with least motion found automatically and selected independently from respiratory pattern
Ignores different respiratory cycles and sums up all data
Static
HD•Chest
Respiratory motion can lead to bias in interpretation
respiratory gating technologies are methods to “freeze” motion
Better image quality via motion freeze

## Slide 43
HD•Chest
Non-gated image demonstrates only a single lesion
Gated image shows a second and third lesion
Non-gated
Courtesy of the University of Michigan, Ann Arbor, MI, USA
* Based on volumetric resolution available in competitive literature for systems greater than 70 cm bore size. Data on file.
Better image quality via motion freeze

## Slide 44
Better image quality via motion correction
Cancer
Cardiac
I.Hong, S.Furst, J.Jones, M.Casey, “The strategy of elastic motion corrections”, Proc. IEEE Medical Imaging Conference, Seoul, South Korea, 10/27-11/2, 2013.
Compared to motion freeze (or gating),
motion correction
recovers more data
(reduced noise)
EMC=elastic motion correction
Gated image

## Slide 45
Enabling promising applications
Enabling better workflow and patient comfort
Enabling research
Enabling better image quality

## Slide 46
Promising applications for PET: prostate cancer
New applications for improved detectability: prostate cancer
    Prostate cancer is the leading cancer for men
    Imaging is very poor, “blind” biopsy is the common diagnostic tool
    Objectives:
    Imaging to guide or replace the biopsy
    exploit potential of new specific PET tracers (PSMA based and others)
    Accurate localization of primary tumor and local affected lymph nodes
Detection and localization of extra capsular lesion and lymph node
Simulation with 11C-choline (6mm lesions, SUV 8)
Low resolution OSEM
5.5 mm pixel, 6 mm filter
High resolution OSEM+PSF+TOF
2 mm pixel, 2 mm filter

## Slide 47
0.1 mSv PET
10 mSv PET
New applications for improved detectability: lung cancer
Low-dose CT screening of persons with high risk for lung cancer):
effective, but 95% false positives
Low-dose PET/CT screening of persons with high risk for lung cancer?
More effective, and much fewer false positives!
Low-dose PET/CT ?
Promising applications for PET: lung cancer screening
J. Schaefferkoetter, J. Yan, D. Townsend, and M. Conti, Phys. Med. Biol. 60 (2015) 5543–5556

## Slide 48
18F-FDG uptake in different tumors:
(blue) Ki=0.015; (red) Ki= 0.043.
Influx Ki = K1*k3/(k2+k3)
Tumor aggressiveness and survival in patients with pancreas cancer can be predicted by kinetic 18F-FDG uptake parameters generated by dynamic PET studies.
Specific tracers are being developed.
Promising applications for PET: pancreas cancer
J Nucl Med. 2013;54:12-18.

## Slide 49
Enabling promising applications
Enabling better workflow and patient comfort
Enabling research
Enabling better image quality

## Slide 50
Data Courtesy of University of Tennessee, Knoxville, Tennessee, U.S.A.
non-TOF (120 sec)
TOF (60 sec)
  An immediate direct advantage, for patient comfort, of improving reconstruction technology: shorter scans
Better patient comfort: shorter scans
  TOF reconstruction produces comparable image with ½ of the counts

## Slide 51
IQ and fast
FlowMotion Clinical Result
Clinical Demand
Stop and Go
Increased Contrast
and Fast Speed
Reduced Motion
Increased Resolution
Speed
or Hi-Rez
Hi-Rez
or Gating
Gating
or Speed
Speed
Speed
IQ or fast
Hi-Rez
HD•
Chest
Speed
Continuous
Flowing
Bed Motion
Stop & Go
Stop & Go
Stop & Go
Stop & Go
Stop & Go
* Based on volumetric resolution available in competitive literature for systems greater than 70 cm bore size. Data on file.
Enabling better workflow: continuous bed motion

## Slide 52
Conventional
FlowMotion
90 cm
HD•Chest
0.8 mm/s
0.8 mm/s
0.5 mm/s
2.0 mm/s
Hi-Rez
Higher Speed
Enabling better workflow: continuous bed motion

## Slide 53
Enabling promising applications
Enabling better workflow and patient comfort
Enabling research
Enabling better image quality

## Slide 54
Enabling research
    Provide tools to researchers to access full raw data:
    List mode data*
    Raw position, energy and time data
    Data corrections (normalization, scatter, attenuation)
    Provide platform to researchers to test new reconstruction:
    Flexible reconstruction interface
    Ability to accept research inputs (i.e. alternative scatter correction)
    Ability to provide intermediate and debug outputs
*http://usa.healthcare.siemens.com/molecular-imaging/petlink-documents

## Slide 55
Enabling research
    Siemens present toolkit: JSRecon/e7tools
    Great flexible research tool !

## Slide 56: Contact page
Maurizio Conti Director, PET physics and reconstruction Siemens Healthcare, Molecular Imaging
Research & Clinical Collaborations
810 Innovation Drive Knoxville, TN 37919, USA
Phone: +1 (865) 218-2390
E-mail:
MaurizioConti@siemens.com
Answers for life.

## Slide 57
Knoxville, Tennessee
Thanks!
Grazie !
Merci!