## Slide 1: Beam Loading Dog-leg experiment status
Frank Tecker (BE-OP)
on behalf of the dogleg team
Introduction
2014 Data Analysis
2015 Spring Run
2015 Autumn Run
Conclusion
18/1/2016

## Slide 2
Drive beam 1-3A
130-50 MeV
12 GHz RF from klystron
12 GHz accelerating structure
CLEX
LINAC
DELAY LOOP
COMBINER RING
BL - BDR experiment
CTF3
J.L. Navarro
Dogleg Beam-Loading Experiment
Unloaded
Loaded (CLIC)
Increasing current
Gradient along the structure
Average gradient 100 MV/m
Beam loading changes the field distribution for the same average gradient ⇨ how is the break-down rate affected?
Reactivated an old beam line (dogleg)
~1.2 A DB current (like CLIC Main Beam)
Measure BDR with/without beam for a direct comparison

## Slide 3: Measurement goals
Final goal: measure BDR
  loaded
  unloaded
for the same average gradient
Alternative: measure BDR for constant input power (when power limited) => BDR decreases?
even decelerate the beam => antiloaded
Unloaded
Loaded
Antiloaded
Unloaded
Loaded

## Slide 4: 2014 Measurement Analysis
Constant input power
Constant average gradient
Unloaded
Loaded
Unloaded
Loaded
Chart: 
| 0.0 | 0.0142 | 0.035 | 0.04912 | 0.06735 | 0.09201 | 0.112 | 0.1357 | 0.1588 | 0.1806 | 0.1975 | 0.2076
Unloaded 34 MW | 91.66 | 92.18 | 93 | 93.67 | 94.2 | 94.9 | 95.2 | 95.5 | 95.5 | 95.2 | 94.6 | 94.3
Loaded 34 MW | 91.66 | 90.66 | 89.56 | 87.26 | 85.16 | 82.36 | 80.06 | 77.36 | 73.31 | 69.76 | 64.28 | 58.36 | 51.37
Unloaded 22MW | 72.66 | 73.18 | 74 | 74.67 | 75.2 | 75.9 | 76.2 | 76.5 | 76.5 | 76.2 | 75.6 | 75.3
Gradient, (au)
Length, m
Chart: 
| 0.0 | 0.0142 | 0.035 | 0.04912 | 0.06735 | 0.09201 | 0.112 | 0.1357 | 0.1588 | 0.1806 | 0.1975 | 0.2076
Unloaded 34 MW | 91.66 | 92.18 | 93 | 93.67 | 94.2 | 94.9 | 95.2 | 95.5 | 95.5 | 95.2 | 94.6 | 94.3
Loaded 34 MW | 91.66 | 90.66 | 89.56 | 87.26 | 85.16 | 82.36 | 80.06 | 77.36 | 73.31 | 69.76 | 64.28 | 58.36 | 51.37
Unloaded 22MW | 72.66 | 73.18 | 74 | 74.67 | 75.2 | 75.9 | 76.2 | 76.5 | 76.5 | 76.2 | 75.6 | 75.3
Gradient, (au)
Length, m
Data biased 
Transmitted power channel broke
=> most relevant interlock missing 

## Slide 5: Interlock vs Cell position
Transmitted power channel
Interlock triggers
many Downstream Breakdowns not detected

## Slide 6: Running March-April
Xbox RF acquisition fixed (broken transmitted channel)
Beam quickly set up
BDR measurements
  unloaded (26 MW)
  loaded (26 MW)
  anti-loaded (10 MW)

## Slide 7: A hot spot on the Dogleg
Initially: breakdown distribution inside structure as expected 
Later: breakdowns mostly detected at beginning of structure 
=> hot spot for breakdowns had developed there
structure unusable for beam-loading experiment => changed
Aug 2014
Apr 2015
position along structure

## Slide 8: 2015 History (cont.)
T24 structure was changed in shadow of the access system change in May against a TD26CC that had been previously in CTF2
pulse compressor breaking down, backlash => taken out => limited in RF power
June: Realignment of the 8mm-diameter collimator (2mm off)
several unavailability periods with
  Xbox klystron solenoid power supplies
  drive beam gun
  pulse compressor installation (6-12 August)
  drive beam priorities for other users
restarted drive beam RF for dogleg on 14 Oct., RF setup completed but no beam due to dogleg line magnet problem
finally restarted on 25/26 Nov

## Slide 9: Quad scans
Dogleg optics initially rematched based on quad scans in girder 10
followed by empirical optimisation

## Slide 10: Beam Setup - Transmission
Gradually improving transmission
no big corrector kicks any more
still small losses through the accelerating structure
intensity varying over time
=> need to improve still

## Slide 11: Dispersion
horizontal dispersion not fully closed
vertical dispersion can still be improved

## Slide 12: Empirical optics setup
Empirical optics found for best transmission has
  non-closed dispersion
  small horizontal waist
=> need to improve dispersion correction

## Slide 13: Beam Phase + Energy profile
BPR0475W (bunch length) + energy nicely flat along the pulse
beam phase constant after transient

## Slide 14: Loaded RF pulse
Beam phase and RF phase have to match to assure constant loading along the pulse
Beam and RF setup OK and output power matches expectations

## Slide 15: Vacuum behaviour
Vacuum on dump side increased when going to 25 Hz beam frequency
Stabilized at 10 Hz
Losses limiting operation?
10 Hz

## Slide 16: Full history of the TD26CC-N1 structure
Gradient
#BDs
Pulse length = 100 ns
150 ns
200 ns
180 ns (CLIC pulse shape)
First measurements with beam ~1.5 A
Pulse width
#BD-clusters
BDR
12/08/2015
14/12/2015
Jorge Giner Navarro
See Jorge Giner Navarro’s talk: “High-gradient structure performances” on Monday afternoon

## Slide 17: BD position in TD26CC-N1 structure
Initially BD more at downstream end
Later at the beginning
Now uniformly distributed
See Robin Rajamäki’s talk: “Breakdown positioning” on Tuesday afternoon
In the structure
Phase information + timing
Aug
Sep
Oct
Nov
Dec
Robin Rajamäki

## Slide 18: 2016 initial plans
Structure conditioned to ~95 MV/m @ CLIC pulse 180ns with ~10-5 BD per pulse
pulse compressor presently limits significant increase in power
=> compare unloaded and loaded operation at constant power back off in input power to have same unloaded gradient as loaded
100 BDs with 10-5 BDR => 111h = 4.6 days @ 25 Hz (longer! if we are limited by vacuum to 10 Hz)
=> we need to
  schedule some days to improve beam quality and transmission
  plan routine running week-ends (+ nights)
after some results we can think of pulse compressor replacement

## Slide 19: Summary
2014 relevant data compromised
Good stable running April/March but structure with hot spot
Difficulties: Pulse compressor replacement, klystron solenoid power converters, drive beam unavailability and priorities
Pulse compressor easier to operate but still power limiting => replacement in 2-3 month?
last weeks 2015: finally restart with beam
  almost full transmission achieved
  orbit without strong corrector kicks
  12 GHz RF interlocks set up
2015 run time very limited, need much more in 2016
consolidate beam setup => ready to go for data taking
A BIG Thank You to everyone involved!

## Slide 20: Acknowledgements
R. Corsini
S. Doebert
W. Farabolini
D. Gamba
L. Malina
J.L. Navarro
T. Persson
P. Skowronski
CTF3
M. Filippova
A. Grudiev
D. Gudkov
P. Guyard
S. Lebet
A. Olyunin
A. Samochkine
A. Solodko
P. De Souza
RF/PM
XBOX
T. Argyropoulos
N. Catalan
S. Curt
A. Degiovanni
J. Giner Navarro
G. McMonagle
R. Rajamäki
S. Rey
I. Syratchev
J. Tagg
L. Timeo
B. Woolley
W. Wuensch
BI
M. Kastriotou
E. Nebot
VAC
A. Bruton
H. Rambeau
SU
P. Bestmann
T. Dobers
J.F. Fuchs
M. Tortrat

## Slide 21: Spare slides

## Slide 22: 2015 running overview
Xbox RF acquisition fixed (broken transmitted channel), cleanup
March: machine restart, rapid beam setup and measurements
May: Accelerating structure had to be replaced due to hot cell (during access system change) pulse compressor breaking down, backlash => taken out
June: Realignment of the 8mm-diameter collimator (2mm off) Klystron solenoid power supply problems
since July: DB gun problems
6-12 August: refurbished pulse compressor installed
14 Oct: RF setup, no beam (faulty magnet)
25/26 Nov: finally beam again
4 Dec: beam setup, MKX stopped
11 Dec: beam, vacuum crate stopped

## Slide 23: Interlock history

## Slide 24
First “comparison”
Unloaded
Loaded
Beam zero crossing
Antiloaded
Only March and April data.
Error bars account only for statistical errors.
Structure with hot cell
We did not observe an increase of breakdown rate with the beam but…
… the beam could have develop a hot-cell.
Nothing conclusive yet
!