## Slide 1: Scintilators Efficiency Test
Prepared By : Ashiq, Arshad,
Yasir
EHEP Lab NCP Islamabad

## Slide 2: Summary of all Scintilators with Position
S.No | Scintilator No | Efficiency Trend | Comments
1 | CN0013(T), AE0014(C),AT0015(B) | Decreasing | Not a Good Scintilators
2 | CN0013(T), AT0015(C), AE0014(B) | Decreasing | Not a Good Scintilators
3 | AE0014(T), CN0013(C),AT0015(B) | Decreasing | Not a Good Scintilators
4 | CN0013(T), AE0016(C),AT0015(B) | Decreasing | Not a Good Scintilators
5 | CN0013(T), AT0015(C), AE0016(B) | Decreasing | Not a Good Scintilators
6 | AE0016(T), CN0013(C),AT0015(B) | Decreasing | Not a Good Scintilators
7 | CN0013(T), CE0011(C),AT0015(B) | Decreasing | Not a Good Scintilators
8 | CN0013(T), AT0015(C), CE0011(B) | Decreasing | Not a Good Scintilators
9 | CE0011(T), CN0013(C),AT0015(B) | Decreasing | Not a Good Scintilators
10 | CN0013(T), AE0015(C),AT0015(B) | Decreasing | Not a Good Scintilators
11 | CN0013(T), AT0015(C), AE0015(B) | Decreasing | Not a Good Scintilators
12 | AE0015 (T), CN0013(C),AT0015(B) | Decreasing | Not a Good Scintilators

## Slide 3
Voltage (T,B,C) | CENTRE CURRENT(UA) | COUNT (T&B) | COUNT {(T&B)&C)} | TOP | BOTTOM | CENTRE | EFFICIENCY %
1100 | 170 | 0 | 0 | 24 | 37 | 106 | 0
1150 | 177.8 | 0 | 0 | 39 | 70 | 126 | 0
1200 | 185.4 | 0 | 0 | 54 | 108 | 205 | 0
1250 | 193.2 | 2 | 0 | 125 | 160 | 380 | 0
1300 | 201.0 | 8 | 2 | 187 | 285 | 760 | 25
1350 | 208.6 | 11 | 3 | 290 | 515 | 1275 | 27
1400 | 216.4 | 21 | 3 | 460 | 830 | 1800 | 14
CENTRE=AT0015

## Slide 4
Voltage (T,B,C) | CENTRE CURRENT(UA) | COUNT (T&B) | COUNT {(T&B)&C)} | TOP | BOTTOM | CENTRE | EFFICIENCY %
1100 | 172.8 | 0 | 0 | 43 | 54 | 54 | 0
1150 | 180.6 | 1 | 1 | 63 | 82 | 93 | 100
1200 | 188.4 | 2 | 0 | 88 | 117 | 163 | 0
1250 | 196.2 | 4 | 1 | 157 | 170 | 195 | 25
1300 | 204.2 | 10 | 1 | 265 | 275 | 320 | 10
1350 | 211.8 | 21 | 6 | 450 | 520 | 640 | 28
1400 | 219.6 | 37 | 4 | 870 | 880 | 1220 | 10
TOP=AT0015

## Slide 5
Voltage (T,B,C) | CENTRE CURRENT (UA) | COUNT (T&B) | COUNT {(T&B)&C)} | TOP | BOTTOM | CENTRE | EFFICIENCY %
1100 | 170 | 4 | 2 | 33 | 40 | 120 | 50
1150 | 177.8 | 0 | 0 | 52 | 117 | 160 | 0
1200 | 185.4 | 3 | 1 | 65 | 130 | 206 | 33
1250 | 193.2 | 3 | 0 | 120 | 190 | 360 | 0
1300 | 201.0 | 11 | 3 | 200 | 299 | 750 | 27
1350 | 208.6 | 10 | 1 | 280 | 510 | 1280 | 10
1400 | 216.4 | 25 | 7 | 455 | 880 | 2100 | 28
BOTTOM=AT0015

## Slide 6
Voltage (T,B,C) | CENTRE CURRENT (UA) | COUNT (T&B) | COUNT {(T&B)&C)} | TOP | BOTTOM | CENTRE | EFFICIENCY %
1100 | 173.0 | 0 | 0 | 31 | 60 | 60 | 0
1150 | 180.8 | 3 | 1 | 51 | 98 | 98 | 33
1200 | 188.6 | 3 | 1 | 91 | 120 | 130 | 33
1250 | 196.6 | 7 | 0 | 150 | 160 | 140 | 0
1300 | 204.4 | 1 | 1 | 223 | 321 | 287 | 100
1350 | 212.2 | 23 | 1 | 424 | 500 | 470 | 4
1400 | 219.8 | 46 | 5 | 900 | 1750 | 910 | 10
CENTRE=AE0014

## Slide 7
Voltage (T,B,C) | CENTRE CURRENT (UA) | COUNT (T&B) | COUNT {(T&B)&C)} | TOP | BOTTOM | CENTRE | EFFICIENCY %
1100 | 170 | 1 | 0 | 48 | 33 | 107 | 0
1150 | 177.8 | 2 | 0 | 59 | 47 | 138 | 0
1200 | 185.6 | 10 | 1 | 111 | 80 | 204 | 10
1250 | 193.4 | 9 | 2 | 138 | 80 | 400 | 22
1300 | 201.2 | 7 | 0 | 230 | 106 | 716 | 0
1350 | 208.8 | 25 | 0 | 480 | 188 | 1120 | 0
1400 | 216.4 | 46 | 4 | 807 | 330 | 2000 | 8
TOP=AE0014

## Slide 8
Voltage (T,B,C) | CENTRE CURRENT (UA) | COUNT (T&B) | COUNT {(T&B)&C)} | TOP | BOTTOM | CENTRE | EFFICIENCY %
1100 | 170.0 | 0 | 0 | 43 | 23 | 40 | 0
1150 | 177.6 | 0 | 0 | 66 | 28 | 120 | 0
1200 | 185.4 | 0 | 0 | 99 | 63 | 224 | 0
1250 | 193.2 | 2 | 0 | 155 | 78 | 374 | 0
1300 | 201.0 | 4 | 2 | 240 | 149 | 675 | 50
1350 | 208.6 | 8 | 2 | 440 | 195 | 1150 | 25
1400 | 216.4 | 17 | 1 | 840 | 315 | 1860 | 5
BOTTOM=AE0014

## Slide 9
Voltage (T,B,C) | CENTRE CURRENT (UA) | COUNT (T&B) | COUNT {(T&B)&C)} | TOP | BOTTOM | CENTRE | EFFICIENCY %
1100 | 170 | 2 | 0 | 26 | 48 | 100 | 0
1150 | 177.6 | 3 | 0 | 53 | 76 | 140 | 0
1200 | 185.4 | 5 | 0 | 78 | 110 | 216 | 0
1250 | 193.2 | 12 | 0 | 86 | 166 | 394 | 0
1300 | 201 | 19 | 1 | 118 | 285 | 795 | 5
1350 | 208.6 | 22 | 5 | 195 | 530 | 1300 | 22
1400 | 216.4 | 45 | 4 | 315 | 900 | 1900 | 9
CENTRE=AE0016

## Slide 10
Voltage (T,B,C) | CENTRE CURRENT (UA) | COUNT (T&B) | COUNT {(T&B)&C)} | TOP | BOTTOM | CENTRE | EFFICIENCY %
1100 | 173.2 | 2 | 0 | 61 | 39 | 56 | 0
1150 | 181.0 | 4 | 1 | 95 | 53 | 83 | 25
1200 | 189.0 | 6 | 0 | 130 | 106 | 95 | 0
1250 | 197.0 | 14 | 1 | 168 | 158 | 169 | 7
1300 | 204.8 | 16 | 0 | 278 | 230 | 242 | 0
1350 | 216.6 | 51 | 3 | 520 | 505 | 500 | 5
1400 | 220.6 | 94 | 7 | 860 | 850 | 820 | 7
TOP=AE0016

## Slide 11
Voltage (T,B,C) | CENTRE CURRENT (UA) | COUNT (T&B) | COUNT {(T&B)&C)} | TOP | BOTTOM | CENTRE | EFFICIENCY %
1100 | 170.2 | 1 | 0 | 43 | 40 | 94 | 0
1150 | 178.0 | 3 | 0 | 70 | 58 | 147 | 0
1200 | 185.6 | 3 | 0 | 114 | 75 | 216 | 0
1250 | 193.4 | 5 | 1 | 145 | 96 | 400 | 20
1300 | 201.2 | 5 | 0 | 234 | 120 | 725 | 0
1350 | 209.0 | 20 | 1 | 450 | 201 | 1210 | 5
1400 | 216.6 | 31 | 3 | 830 | 350 | 1910 | 9
BOTTOM=AE0016

## Slide 12
Voltage (T,B,C) | CENTRE CURRENT (UA) | COUNT (T&B) | COUNT {(T&B)&C)} | TOP | BOTTOM | CENTRE | EFFICIENCY %
1100 | 170 | 2 | 0 | 50 | 47 | 94 | 0
1150 | 177.6 | 2 | 0 | 85 | 45 | 132 | 0
1200 | 185.4 | 10 | 1 | 120 | 100 | 236 | 10
1250 | 193.2 | 13 | 5 | 174 | 165 | 440 | 38
1300 | 201.0 | 24 | 2 | 283 | 292 | 725 | 8
1350 | 208.6 | 47 | 5 | 510 | 460 | 1200 | 10
1400 | 216.2 | 91 | 3 | 900 | 850 | 2000 | 3
CENTRE=CE0011

## Slide 13
Voltage (T,B,C) | CENTRE CURRENT (UA) | COUNT (T&B) | COUNT {(T&B)&C)} | TOP | BOTTOM | CENTRE | EFFICIENCY %
1100 | 173 | 3 | 0 | 71 | 39 | 91 | 0
1150 | 180.8 | 14 | 0 | 75 | 74 | 133 | 0
1200 | 188.6 | 4 | 0 | 95 | 90 | 220 | 0
1250 | 196.6 | 18 | 2 | 164 | 144 | 406 | 11
1300 | 204.4 | 24 | 1 | 273 | 260 | 675 | 4
1350 | 212 | 59 | 10 | 545 | 490 | 1300 | 17
1400 | 219.8 | 93 | 13 | 890 | 895 | 2900 | 14
TOP=CE0011

## Slide 14
Voltage (T,B,C) | CENTRE CURRENT (UA) | COUNT (T&B) | COUNT {(T&B)&C)} | TOP | BOTTOM | CENTRE | EFFICIENCY %
1100 | 170 | 0 | 0 | 62 | 52 | 96 | 0
1150 | 177.8 | 2 | 0 | 77 | 65 | 128 | 0
1200 | 185.6 | 10 | 1 | 120 | 112 | 212 | 10
1250 | 193.4 | 9 | 2 | 170 | 161 | 360 | 22
1300 | 201 | 12 | 2 | 280 | 290 | 650 | 16
1350 | 208.8 | 36 | 1 | 480 | 500 | 1180 | 2
1400 | 216.4 | 77 | 12 | 860 | 1020 | 2130 | 15
BOTTOM=CE0011

## Slide 15: Summary and Future work
Testing and efficiency of scintelaltors is presented.
Find efficiency by using cross??
Order new PMTs (exact detail from CERN).

## Slide 16: Thank you

## Slide 17: Summary of Scintilators Efficiency
S.No | Scintilator No | Centre Efficiency | Top Efficiency | Bottom Efficiency | Comments
1 | Scintilator # CN0013 | Reference | Reference | Reference | Appears Ok
2 | Scintilator # AE0014 | Decreasing | Decreasing | Decreasing | Appears ok
3 | Scintilator # AE0016 | Decreasing | Decreasing | Decreasing | Appears ok
4 | Scintilator # CE0011 | Decreasing | Decreasing | Decreasing | Appears Ok
5 | Scintilator # AT0015 | Decreasing | Decreasing | Decreasing | Appears ok
6 | Scintilator # AE0015 | Reference | Reference | Reference | Appears Ok