## Slide 1
Resonance Control
at the Compact ERL in KEK
Takako Miura (KEK)
LLRF Group
Takako Miura, Shinichiro Michizono, Feng Qiu, Toshihiro Matsumoto, Hiroaki Katagiri, Dai Arakawa (KEK),
Liu Na (Sokendai Univ.)
SRF Cavity Group
Eiji Kako, Takaaki Furuya, Hiroshi Sakai, Kensei Umemori, Taro Konomi, Masato Egi, Kazuhiro Enami (KEK),
Masaru Sawamura (QST)
on behalf of LLRF and SRF Cavity Group

## Slide 2
©Rey.Hori/KEK
Injector LINAC
Main LINAC
Circumference ~ 90m
Marger
Beam Dump
Compact ERL (cERL) is a test facility of 3-GeV ERL as a future light source.
9-cell SRF cavity x 2
2-cell SRF cavity x 3
QL = 1 107
Nominal beam energy | 35 MeV
Nominal Injector energy | 5 MeV
Beam current | 10 mA (initial goal)
Normalized emittance | 0.1 ‒ 1 mmmrad
Bunch length | 1-3ps (usual) 100fs (short bunch)
Buncher
Photocathode DC gun
Design parameters of the cERL
RF = 1.3 GHz CW
QL = 5105
Introduction of cERL
Requirements of RF stabilities for 3GeV-ERL
0.01%rms,0.01deg.rms

## Slide 3
 | Buncher | Inj-1 | Inj-2 | Inj-3 | ML-1 | ML-2
Cavity | NC | 2cell-SC | 2cell-SC | 2cell-SC | 9cell-SC | 9cell-SC
Cavity Voltage | 114 kV | 0.7 MV | 0.7 MV | 0.7 MV | 8.6 MV | 8.6 MV
Field Gradient (Desgin) |  | 3 MV/m (7.5MV/m) | 3MV/m (7.5MV/m) | 3MV/m (7.5 MV/m) | 8.6 MV/m (15MV/m) | 8.6 MV/m (15MV/m)
QL | 1.1 105 | 1.2  106 | 5.8  105 | 4.8  105 | 1.3  107 | 1.0  107
Cavity Length | 0.068 m | 0.23 m | 0.23 m | 0.23 m | 1.036 m | 1.036 m
RF Power @Low beam current | 3 kW | 0.53 kW | 2.6 kW |  | 1.6 kW | 2 kW
Vector-sum
BUN
Inj1
Inj2
Inj3
ML2
ML1
Current status of high power RF sources
300kW Klystron
25kW Klystron
8 kW SSA
8 kW SSA
16 kW SSA
The "vector-sum" operation in low energy region, β<1, has some difficulties.
-> transit time change
-> RF phase change

## Slide 4
Kly Pf
Resonance Tuning
Tuner board
KLY/SSA
LO
Cavity pick-up signal
IF (10MHz)
LO (1.3GHz(1+1/128)1310MHz)
IF (10MHz)
Pc
Pf
RF FB board
I
Q
MO(1.3 GHz)
RF SW
1.3 GHz
Digital I/O
DAC
(CW/CCW pulse output)
Motor driver
Piezo driver
Offset
+250V
-250V~250V
LO
Fast Interlock
(Arc, Pref, etc.)
0 ~ 500V
Digital LLRF System at cERL
Requirements of RF stabilities for 3GeV-ERL
0.01%rms,0.01deg.rms
Pieso Tuner : fine tuning
Mechanical tuner : coarse tuning
Baseband

## Slide 5
Tuner System

## Slide 6
He jacket
Tuner system of Injector Linac
performance of slide-jack tuner
performance of piezo tuner
20 kHz/rotation
Piezo: 0-500V => 2 kHz
Piezo tuner
0-500V
stroke= 4 m@2K
E. Kako, IPAC2013
2kHz
600 kHz
Drive Shaft
Taper
Slide-Jack tuner
f0
f
f1/2
2Δf1/2 = f0 / QL
Inj1: QL= 1.2 106, 2Δf1/2 =1.1 kHz
Inj2: QL= 5.8 105, 2Δf1/2 =2.2 kHz
Inj3: QL= 4.8 105, 2Δf1/2 =2.7 kHz
Piezo tuner can cover the band width.
The same as KEK-STF tuner system

## Slide 7
Tuner system of Main Linac
Piezo tuner
0V - 500V (offset=250V)
Stroke= 4 m @ 2K
40 m @ 300K
(1 m: 300Hz)
Cavity flange is fixed here
He jacket
Shaft of slide-jack tuner
Slide-jack tuner
Screw
1kHz
QL=
2Δf1/2 = f0 / QL=130 Hz
H. Sakai, SRF2013 @Paris

## Slide 8
435Hz
INJ3
Mechanical Resonance of Inj Cav
Eacc: 1MV/m
Sinusoidal wave (40Vpp) was fed to piezo tuner.
Mechanical resonance is scanned by sweeping the input frequency.
<Injector>
409Hz
INJ2
Phase detector : 20 mV/deg
INJ1
423Hz
260Hz
68Hz
Large mechanical resonance exists around 400 Hz.

## Slide 9
M. Satoh, IPAC2014
Mechanical Resonance of ML Cav
Impulse hammer response
Large mechanical resonance exist near 50 Hz.

## Slide 10
LLRF System

## Slide 11
Total 11 boards are used for operation.
Digital LLRF Boards
 | BUN | Inj1 | Inj2 | Inj 3 | ML1 | ML2
RF FB board | FB0 | FB1 | FB2 (Vec-sum) |  | FB4 | FB5
Tuner board | TN0 | TN1 | TN2 | TN3 | TN4 | TN5
Embedded Linux is working in the PowerPC on FPGA.
Each board acts as an EPICS IOC.
Data acquisition is performed through GbE bus on the backplane.
AMC(Advanced Mezzanine Card)
16-bit ADC x 4
(LTC2208)
16-bit DAC x 4
(AD9783)
Digital I/O
FPGA
Virtex5-FX
RF FB Control
Tuner Control
MicroTCA
(Mitsubishi Electric TOKKI Systems Co.,Ltd.)

## Slide 12
Field Feedback Control
LPF:100kHz
I
Q
A


## Slide 13
-3dB
2Δf1/2
Cavity Resonance
f1/2 =65 Hz for ML cavities (QL=107)
Narrow bandwidth for f0=1.3 GHz
: The phase difference between the input
RF and the cavity pickup signal
-45deg
45deg
f0
Detuning Δf=fc-f0
Amplitude (dB)
For constant acceleration field,
double input power is necessary at f = f1/2
To keep resonance frequency,
tuner should be controlled to maintain  at zero.

## Slide 14
Block Diagram of Resonance Control
Digital I/O
Amplitude
Level judge
Kp
Ki
FB on/off
Offset
DAC
CW
CCW
ADC
Pf
Cav
LPF
for protection
CLK
80MHz
0 <
0 >
ROT1
ROT2

I
Q
I
Q
ADC
Polarity
change
offset
θoffset
Δθ
+
-
Δθ'
IIR LPF
IIR LPF
Mechanical Tuner
Piezo Tuner
Feedback Control:
Δθ = θf (Pf)  θc (cav) => 0
Digital I/O
DAC
CW
CCW
Stepper motor
FPGA Board
I
Q
A
θ
fc=20Hz

## Slide 15
Cavity field Stability & Microphonics

## Slide 16
Waveform of ML Cavities
ML1
ML2
Amplitude
Phase
Amplitude
Phase
Vc: w/o RF Feedback
A = 0.012% rms
 = 0.014 rms
A = 0.013% rms
 = 0.015 rms
A = 0.035% rms
 = 0.3 rms
A = 0.15% rms
 = 0.6 rms
T. Miura, IPAC2014 @Dresden
50Hz
Microphonics
Vc: w RF Feedback
4deg
1.4deg
Field fluctuation by Michrophonics is stabilized by RF Feedback
Notes: If detuning is seriously large, tuner FB control should be improve.

## Slide 17
Hz
dBc/Hz
Phase noise jitter measurement
using Signal Source Analyzer
Agilent E5052B
Phase noise by Microphonics was suppressed well by RF FB.
Phase noise of Vc with FB was almost the same as that of Master Oscillator.
Microphonics is observed at 10 Hz - 400Hz.
50Hz
T. Miura, IPAC2014 @Dresden
ML2
RP: Rotally pump
M.Egi, PASJ2016 (MOP025)
Vibrational state of "ﬂoor" around Main Linac
Vc Phase Noise with RF FB (10Hz-1MHz)=0.017deg
Vc Phase Noise w/o RF FB (10Hz- 1MHz)=0.73 deg

## Slide 18
Countermeasure against Scroll Pump Vibration
vibration by scroll pumps
After
Before
Field gradient
8.3 MV/m : Operation point
(15 MV/m : Design)
For constant input RF power
50 Hz
9-cell SC cavity: QL=107
The rubber sheet was inserted
under the scroll pump. The 50 Hz vibration is suppressed.

## Slide 19
Momentum Error due to Vector-sum Error
&
Improvement by tuner feedback parameter

## Slide 20
 | Inj1 | Inj2 & Inj3 | ML1 | ML2
Amplitude | 0.006% rms | 0.012% rms | 0.003% rms | 0.003% rms
Phase | 0.009 rms | 0.022 rms | 0.010 rms | 0.009 rms
RF Performance
<Measurement condition>
Beam: 5Hz, 3ps rms, 23 fC, total Energy=19.9 MeV
Momentum drift of ~15 minutes period was observed.
36 min.
Momentum stability = 0.013% rms
Measurement of Beam Momentum Stability
for confirmation of RF stability
RF Stabilities for Short Time
Almost satisfied the requirement of 3-GeV ERL

## Slide 21
Inj1
Inj2
Inj3
Time
Detuning phase
(deg.)
What causes Energy Drift ?
Time interval of detuning is similar to the interval of energy drift.
~15 min.
Large ripple depends on valve control for liquid N2.
Input-couplers of injector are cooled
by liquid N2.
<RF source : cavity =1:1 >
Cavity phase is stabilized by RF FB.
<Vector-sum operation>
Vector-sum is constant,
but each cavity phase fluctuates.
Inj2
Inj3
Inj1
ML1
ML2
10deg
Cavity Phase(deg.)
Cavity Phase
Vector-sum error may cause energy drift.

## Slide 22
Vector-sum calibration error
Amplitude & Phase calibration error
(2) for low beam energy ( < 1),
transit time is affected by cavity field.
Vector-sum Error
In order to minimize the momentum fluctuation due to the vector-sum error,
“detuning” should be stabilized.
Possibility of momentum drift caused by vector-sum error
Injection energy : 1.63MeV@Inj2, 2.36 MeV@Inj3
Cavity fields changes => Beam phase changes.

## Slide 23
Result of Resonance FB Control Improvement
Cav |  deg (rms) | f Hz (rms)
INJ1 | 0.23 | 2.2
INJ2 | 0.10 | 2.0
INJ3 | 0.09 | 2.1
ML1 | 0.09 | 0.08
ML2 | 0.16 | 0.18
5 h
Detuning phase
Δθ (deg)
Cavity phase
θc (deg)
Inj2&Inj3 cavity phase
become stable.
Higher FB gain in resonance control is adopted for small detuning.
Detuning due to liquid N2 flow has been compensated.
Cavity Phase
Detuning Phase

## Slide 24
Stability of Beam Momentum (2)
50 min.
Momentum Jitter= 0.0064% rms
Large momentum drift disappeared.
=> Beam momentum jitter ΔP/P =0.003% is achieved.
P/P = 0.003% rms
30 min.
The latest result
Measurement after modification of tuner feedback gain
FB gain is optimized
for Inj2 & Inj3 only

## Slide 25
Summary
Digital control boards are applied to RF feedback and tuner control.
Owing to the stiff cavity structure, so detuning by Michrophonics does not influence to the operation.
The field fluctuation due to Michrophonics is well suppressed by RF feedback.
Vector-sum operation has some difficulty for low beam energy due to different transit time.
By applying high FB gain for piezo tuning, the detuning fluctuation due to liquid N2 flow rate has been suppressed. => Beam energy drift caused by vector-sum error has become small.
0.003% momentum stability is achieved.