## Slide 1: High Power Model of Isolator
Rutambhara Yogi
www.europeanspallationsource.se
24 June 2016

## Slide 2
Suppliers:
Many thanks to
Colleagues:
ESS: Mats Lindroos, David McGinnis, Anders Sunnesson, Morten Jensen,
A. Johansson, Carlos Martins, Rihua Zeng, Rafael Montano, Chiara Marrelli,
Stevo Calic, Bruno Lagoguez, Staffan Ekström, Daniel Lundgreen, Carl Johan Hardh
FREIA: Rolf Wedberg, Lars Hermansson, Roger Ruber, Magnus Jobs & other colleagues
CERN: Eric Montesinos, Olivier Bruner
SNS: Crofford Mark

## Slide 3: Outline
Brief introduction to ESS RF systems
Circulator
High power model of Isolator
Effect of hot water cooling

## Slide 4
ESS: European Spallation Source
Being constructed in Lund, Southern Sweden
Beam pulse width = 2.86 ms
Pulse repetition rate = 14 Hz
Peak proton beam power to the target = 125MW
Most intense pulsed neutron source in the world : peak beam power larger by factor 5 compared to existing spallation facilities
Europe’s one of the largest infrastructure

## Slide 5: Power profile
84 High Beta cavities
704 MHz
New development: IOT
Backup solution: Klystron
26 Spoke Cavities,
352 MHz
400kWp*/20 kWavg
Tetrode
36 Medium Beta cavities
704 MHz
1.5MWp Klystron
1200 kWp*/
60 kWavg
1200 kWp*/60 kWavg
Warm Linac 352 MHz
-1 RFQ & 5 DTL:
Klystrons
3MWp* / 150 kWavg
*- Amplifier power with overhead
Coax upto circulator,
WR2300(hh )
WR1150
WR1150
- 3 MEBT
SSA
30kWp* / 1.5kWavg
WR2300 FH,
WR2300 HH
1-5/8 inch
coax

## Slide 6
Schematic of RFDS for MB/HB

## Slide 7
Racks
Gallery layout MB cell
Gallery layout for Spokes
WR1150

## Slide 8
RFDS layout for HB / MB cell
Space envelope
For Circulator
Space envelope
for Load
DC-1
DC-2
DC-3
Removable
wg section for PSS
Shutter
switch
RFDS layout
same for
MB/HB upto
this flange
IOT

## Slide 9: Circulator
Ideal Circulator: Like a round about
Can have either 3 ports or 4 ports
Can use coaxial line / waveguide
Direction – clockwise / anti-clockwise
2.856 GHz
P = 6 MWp
Pavg = 25 kW
Indigenous development done at SAMEER (India) in 1999. Installed on medical linac.
352 MHz
P = 400 kWp
Pavg = 20 kW
Technology demonstrator
for ESS spoke linac
Notes: After lot of technical negotiations this unpresuruzed cirulator in 6 inch coax was developed.

## Slide 10: Circulator
Nonreciprocal device
Non-reciprocity is achieved due to ferrite
Generally ferrite discs are used
The discs can be either two discs or the multiples of two. Depends on power and thermal management
Difference in μ+ and μ-, leads to non-reciprocity

## Slide 11
Transmission
from port 1 to port 2
Isolation
From port 1 to port 3
RF signal applied to the ferrite disc will generate two counter rotating waves, velocity dependent on propagation direction.

## Slide 12: Construction of 3-port Circulator

## Slide 13: Ideal & Practical Isolator Isolator = Circulator + load
1
2
3
A
Load
Isolator
Amplifier is completely isolated from
unwanted reflections
Ideal: Infinite Isolation
Practical: Finite Isolation
A
1/1000 of
power
Isolation = - 30 dB
Input power = 1500 kW
Only 1500 W = 1.5 kW
will reach amplifier ?
Load – Dissipate RF power from amplifier

## Slide 14
Load
A
Cavity
1
2
3
Isolation = 30 dB
Return loss = 30 dB
Load return loss = - 20 dB
1.5 kW
15 kW
15 kW
Full reflection
from cavity
1.5 kW
Practical Isolator:
1500 kW
So we need a load with good
Return loss

## Slide 15: High Power Model of Isolator
VT = V1 + V2 +V3
A small program is written for effective return loss calculation
(Ref: ESS-0043091)
Forward power /
Circulation direction
Reflected power

## Slide 16
Assumptions:
Amplifier is perfectly matched.
Higher order harmonics are not considered
Uncertainties in S-parameter measurement not considered
Only first order reflections are considered
Worst case S-parameters and worst phases for S11eff calculations are considered.

## Slide 17
Parameters of circulator :
S11= -30 dB
S12 = -30 dB
S21 = -0.5 dB
NO higher order harmonics
Variation of Effective return loss of Isolator with return loss of load
Amplifier
Specification
We need a GOOD LOAD:
-30

## Slide 18
Worst Case Input Return Loss |S11eff| for |S11| = |S12| = 30dB (typical)

## Slide 19
Variation of Effective return loss of Isolator
with Phase of load return loss
Parameters of Circulator:
S11 = -30 dB
S12 = -26 dB
S21= -0.05 dB
Parameters of Load:
S11 = -30 dB
Value is not important.
Phase of load return loss: stable
Waveguide section can be added to get desired phase shift

## Slide 20: Testing of concept
Courtesy: Magnus Jobs at FREIA
Short
Towards VNA
Load with return loss
variation
Trend of both the curves is same
Good return loss when phase of load return loss at circulator is about 180 o
Further experimentation is planned with high power
Circulator and load at Lund test stand.
Experiment

## Slide 21: Additional boundary conditions in ESS: High temperature water cooled loads
Initially Outlet temp
Controlled to 60oC
Klystron
Load
Klystron
Load
N Number
Inlet temp
50oC
Power dissipated across load
With HV on, RF off:
ΔT ~ 10 oC
With HV on, RF on:
Nominal beam operation,
Klystron testing,
ΔT ~ 5 oC
Temperature rise due to power dissipation across collector:

## Slide 22: Effect of unstable phase of load return loss
At same power level, cooling water temperature will vary from 50o C – 80o C depending upon operational scenario.
If phase of load return loss also varies, S11eff seen by amplifier will vary automatically, which may affect klystron stability.
# Klystron operation will be affected
Hence stable phase of load return loss will be preferred.

## Slide 23
New Load development for ESS
Hot water cooled loads
Frequency = 704 MHz
Power ≥ 1500 kWp
Pulse width = 3.5 ms
Pulse repetition rate = 14 Hz
Inlet water temp ≥ 50o C,
Outlet water temp ≤ 80o C
Temp: 50oC, 60oC, 80oC,
80oC: After 24 hours
S11 > 30 dB
BW > ±10 MHz
Return Loss
Phase of Return Loss
ΔΦS11< 10o
Prototypes by MEGA, Thales will be delivered soon.

## Slide 24
Thank you !

## Slide 25: Functional Requirements of RFDS
Transportation of RF power from amplifier to the cavity coupler with minimum losses
Dissipation of unwanted RF power during different operational scenario - Loads
Protection of amplifier from the unwanted reflections during filling and emptying of cavities or due to sparks – Circulator
To satisfy requirements of personnel safety system (PSS) during different operational scenarios – removable waveguide, PSS flange, switch
Detection of sparks in the RF system and give signal to the local protection system (LPS) – Arc Detectors

## Slide 26
Measurement of forward and reflected power from amplifier and cavity with the required accuracy, granularity - Directional couplers
Take care of thermal expansion-contraction, dimensional changes in building layout such as changes in the tunnel length, differential settlement of tunnel with respect to gallery etc – Bellows
Simplification of cooling system design for energy recovery – High temperature water cooled loads
Depending on requirement, to ensure stable klystron operation by changing distance between klystron and circulator – extra flanges

## Slide 27: Prototype RFDS at ESS
# Hot water cooled loads – New development for ESS
Being developed by three suppliers: AFT, MEGA and Thales
Received AFT load in Lund test stand,
MEGA and Thales load will be shipped by end of July / August
# Circulators being developed by
three suppliers: AFT, MEGA and FMT
FMT circulator will be shipped in
a few days.
DR for MEGA circulator: 27 June
AFT circulator: Expected delivery
by Aug end.
# High directivity directional
couplers (D > 40 dB)- by MEGA
3.8 km Waveguide
E-elbows > 750
H elbow > 550
For MB/HB Linac,
Huge requirement of
waveguides and elbows !

## Slide 28
Flange and waveguide design
Waveguide extends through
the flanges
Fabrication becomes extremely
easy and avoids discontinuities
Flanges will be welded along the
periphery (*)
Surface cut along flange surface to ensure flatness, perpendicularity and parallelity
Flat flanges (without RIM)
Time spent at ESS < 20 hours.
*
*
Received prototypes from
two companies.
Failed in visual inspection
Costs by metal manufacturers expected to be 20-30% cheaper than RF companies ~ MEUR

## Slide 29
# Arc Detectors (AD)
Procured two types of Arc Detectors
- AFT : uses optical fibre, response time < 10 μs
- Microstep/CERN AD – don’t need optical fibre
* Four redundant photodiodes: to detect spurious arcs
* Two blind photodiodes: to detect radiation induced arcs
Tests will be performed in the Lund test stand and then preferred solution will be selected.