Tracker Readout
- Tracker systems have been stable since last cm.
  - No major changes to Hardware or DAQ code since last cm.
- Just prior to shifts starting in February cryostat 3 had a minor shudder and was taken offline,
  - Cryostat 3 warmed up, and was then cooled down.
  - Suspect a small amount of water entered the cryostat and resulted in the disconnection of some VLPC channels (50+),
  - To recover the problem a waveguide was successfully was moved from a bad (50 dead ch) module to a better module (8 bad ch), this was done quickly and efficiently thanks to the whole tracker team!
- Ongoing work in the tracker lab to assemble spare front end boards and buffer boards.
  - 4 validated good boards.
  - Need to calibrate temperature/bias for spare boards.

---
Tracker Controls
- Control system has also been stable since the last cm.
  - No major changes.
  - Now using “Pro” version.
- Pierrick is adding state machine to tracker controls:
  - Will add alarm handler to monitor alarm on hardware problems
  - More sophisticated control of archiver

---
Tracker: Bias Calibration
Optimum Bias is a balance between dark counts
●
(noise) and quantum efficiency.
Top plot gives light yield as a function of noise.
Quantum efficiency flattens off, however noise
grows exponentially.
Ideal region is ~0.02 = 2% probability of a
generating a hit over the first photo-electron
charge within a 180ns gate. 220kHz singles rate.
Optimum bias is performed by scanning a range
of bias voltages. The optimum is then selected
where the dark count ratio is closest to 0.02.
A common bias is shared between groups of 64
channels.
The optimum biases are checked by eye once.
Plots from D. Adey Thesis. For more information
read it!
Calibration & Data Quality
8
Inefficient
Noisy

---
Tracker: ADC Calibration: Data Collection
For the tracker reconstruction to work properly a
calibration file is required which converts from
the ADC values to a meaningful ‘npe’ value.
Data for this is collected by the
“RunPedCalib.py” script, which is run by the
shifters at the start of each shift.
The calibration dataset contains two runs, one
with the internal light injection system enabled,
and one with this disabled.
All tracker calibration data is stored on a RAID1
disk, nightly backed up to micenet backups and
finally uploaded to the grid once a week.
Each calibration data set requires 2.6MB of disk
space.
9

---
Tracker: ADC Calibration: The Fit
 At the heart of the calibration is the fit. A
This strategy was empirically found to produce
combined poisson-gaussian function is used
the best reliability and results, at the expense
which accurately describes the ADC distribution.
of cpu time..
The fit is conducted simultaneously (one
The chi2/ndf values of each fit are checked,
minimisation) on LED and no LED histograms,
and if they exceed a threshold the channel is
sharing common parameters.
flagged for human investigation.
LED
No LED
10

---
Tracker: ADC Calibration: The Workflow
The Calibration process follows the workflow:
The “Manual” process has a human check
stage.
Can be time consuming.
–
Not good for daily updates!
–
11

---
Tracker: ADC Calibration: The Workflow
Addition of an “Automatic” calibration update.
This only provides an incremental update to
the location of the pedestal
The VLPC gain is very stable, which we have
checked in MICE.
Process uses an existing calibration as a
starting point and only updates the pedestal
location.
Checks are finally performed on the light yield
and noise rates.
Plan is to integrate this as part of the data
collection script, to alert the control room
before taking data that the tracker is not
comparable to a “good” reference state.
Online running still under development.
12

---
Tracker: ADC Calibration: Auto Checks
Light Yield Residual (pe)
Dark Rate Residual
Channel
Plot the residual dark rate, and light yield between this and a previous calibration. If these change
then the detector state is not ‘stable’.
Summary of all tracker calibrations since Sept 2015 available at:
http://heplnv152.pp.rl.ac.uk/tracker/summary.xhtml
–
13
Pedestal Change

---
Tracker Calibration
- Improvements made to calibration process to facilitate regular calibration updates:
  - Master calibration performed at beginning of ISIS cycle
  - Daily updates to calibration account for small shifts and are verified against “master” performance.
  - Calibration information can be found at: http://heplnv152.pp.rl.ac.uk/tracker/summary.xhtml.
- Aim to automate the “per shift” calibration updates
  - And launch from a button!