## Slide 1: A Wall Current Transformer for bunch-by-bunch intensity measurements in the LHC
Michal Krupa 10/03/2016

## Slide 2: Outline
Bunch-by-bunch intensity measurements in the LHC
Wall Current Transformer design and installation
Laboratory and beam measurements
Summary

## Slide 3: Bunches in the LHC
Typical LHC values:
Bunch intensity: 5·109 – 1.5·1011 particles
Bunch length: 1 - 2 ns
Bunch spacing: 25 ns

## Slide 4: From FBCT to ICT / WCT
Run 1: Fast Beam Current Transformers (FBCT) used for bunch-by-bunch intensity measurements in the LHC. Sensitivity to the transverse beam position and bunch length.
LS1: Design of new monitors for absolute bunch-by-bunch intensity measurements. Two developments in parallel:
  Integrating Current Transformer (ICT) by Bergoz
  Wall Current Transformer (WCT) by CERN
2015: Commissioning of ICT and WCT
2016: Installation of two WCTs as the operational system

## Slide 5: FBCT principle of operation

## Slide 6: Bergoz ICT
Commercially available toroid compatible with the existing LHC FBCT mechanics and cooling system

## Slide 7: WCT principle of operation

## Slide 8: WCT principle of operation

## Slide 9: WCT design
All parts cut in half – installation and removal does not require breaking the accelerator vacuum

## Slide 10: WCT design
Calibration winding: - single turn - calibration with current - low resistance at low frequencies (minimising power dissipation) - high impedance at high frequencies (decoupling from beam)
Signal winding: - few turns - low impedance - averaging two nearest transformers on the PCB
Signal addition: the four outputs are passively summed (averaged) outside, but close to, the monitor

## Slide 11: Magnetic cores

## Slide 12: LHC Point 4 installation in 2015
System A (operational): no change during LS1
System B (development): two new monitors (WCT + ICT), digital acquisition in parallel to DABs

## Slide 13: LHC Point 4 installation in 2016
System A (operational): two new WCTs installed during YETS 2015/2016 to be commissioned in 2016

## Slide 14: LHC WCT installation
Design launched: 01/10/2013 First installation in the LHC: 02/03/2015

## Slide 15: Bandwidth
Test bench not very reliable above 1 GHz
Lab WCT: - work in progress - RF bypass - 1.1 GHz filter - low cut-off: 500 Hz
LHC WCT: - RF bypass - 400 MHz filter - low cut-off: 500 Hz
LHC ICT: - no filter - low cut-off: 600 Hz

## Slide 16: FBCT / WCT nominal bunch

## Slide 17: Bunch position sensitivity MD
Dedicated MD on 20/07/2015
FBCT ΔI/I ≈ (0.3 – 0.7)% / mm | WCT ΔI/I < 0.001% / mm Not measurable

## Slide 18: Bunch length sensitivity MD
Ad-hoc MD on 22/09/2015
FBCT ΔI/I < 0.2% / ns | WCT ΔI/I < 0.2% / ns

## Slide 19: Radial modulation
Chromaticity measurement by modulation of the RF frequency.
LHC fill: 3670 01/05/2015 Post LS1 commissioning WCT was rescaled to overlap with the FBCT

## Slide 20: Orbit correction
LHC fill: 4640 21/05/2015 Proton physics WCT was rescaled to overlap with the FBCT

## Slide 21: Satellite and ghost bunches
LHC fill: 4569 03/11/2015 Stable beams
LDM: single photon counting over 3 million turns (5 min)
WCT: averaging on a 12-bit oscilloscope over 500 turns (50 ms)

## Slide 22: Summary
During 2015 performance of the first prototype WCT was evaluated with the LHC beams
The WCT was shown to be insensitive to beam position and bunch length variations
The WCT output pulse is shorter than the output pulse of the FBCT – measurements are easier
During YETS 2015/2016 the two operational FBCTs were replaced with the WCTs
The new WCTs are waiting for commissioning with beam

## Slide 23: Thanks for the attention
Questions?
Special thanks to:
M. Gasior and S. Bart Pedersen, D. Belohrad, N. Chritin, F. Guillot-Vignot, J. Kral, T. Lefevre, P. Odier, L. Soby, BI Day organisation

## Slide 24: BSRT calibration

## Slide 25: FBCT phasing

## Slide 26: LHC WCT installation
Active splitting – no cross-channel reflections
Different gain and bandwidth for each channel

## Slide 27: FBCT / ICT / WCT nominal bunch

## Slide 28: Impedance

## Slide 29: Beam position sensitivity MD
FBCT ΔI/I ≈ (0.3 – 0.7)%·mm-1 | ICT ΔI/I ≈ (0.01-0.02)%·mm-1 | WCT ΔI/I < 0.001%·mm-1 Not measurable
Dedicated MD on 20/07/2015

## Slide 30: Bunch length sensitivity MD
FBCT ΔI/I < 0.2%·ns-1 | ICT ΔI/I < 0.2%·ns-1 | WCT ΔI/I ≈ 0.5%·ns-1 Not understood at the time of the MD
Dedicated MD on 20/07/2015

## Slide 31: Baseline droop and restoration
WCT: low cut-off: 500 Hz
Beam: 1632 bunches 2.1e14 total intensity
LHC fill: 4567 02/11/2015 Injection energy

## Slide 32: WCT bunch response

## Slide 33: Doublet bunches
LHC fill: 4064 26/07/2015 Scrubbing run
WCT: averaging on a 12-bit oscilloscope over 50 turns (5 ms)
Old amplifiers

## Slide 34: Satellite and ghost bunches
LHC fill: 4640 21/11/2015 Stable beams
LDM: single photon counting over 3 million turns (5 min)
WCT: averaging on a 12-bit oscilloscope over 500 turns (50 ms)

## Slide 35: Energy ramp
Bunch length reduction, RF change, orbit change, unbunched beam loss
LHC fill: 4643 21/11/2015 Ramp FBCT, ICT and WCT were rescaled to overlap with the DCCT

## Slide 36: Van der Meer scans
Absolute bunch-by-bunch measurements
LHC fill: 4538 24-25/08/2015 VDM scans No rescaling, all plots as logged

## Slide 37: (un)Stable beams
LHC fill: 4538 26/10/2015 Stable beams ICT and WCT were rescaled to overlap with the FBCT

## Slide 38: WCT calibration circuit
Current source:
~10 ppm shot-to-shot stability
~100 ppm accuracy
~5 µs switch-on time
ICAL limited by VCAL (~1.5 A for RCAL ~ 10 Ω)
RREF: 5 Ω ± 0.01% < 0.2 ppm/oC

## Slide 39: WCT calibration circuit
Current measurement:
Half-bridge principle
Variable reference voltage source controlled by potentiometer and DAC (coarse / fine)
RREF: 10 Ω ± 0.01% < 0.2 ppm/oC
RHB: 100 kΩ ± 0.01% < 1 ppm/oC