## Slide 1
HYDROSTATIC LEVELLING SENSORS
BASED ON EXTRINSIC FIBRE
FABRY-PEROT INTERFEROMETER
TECHNOLOGY
14. International Workshop on Accelerator Alignment · 3 - 7 October 2016 · ESRF – Grenoble - France

## Slide 2
PROJECT OVERVIEW
HL-LHC
AIM
upgrade luminosity of LHC
increase collisions by factor 10
lifetime of inner triplets reached by 2023 @ 300 fb-1
Alignment consequences
maintain alignment concept
with hydrostatic levelling for vertical and roll
larger zone to be covered by monitoring systems

## Slide 3
SENSOR
REQUIREMENTS
SENSOR
MAIN HLS
REQUIREMENTS
annual radiation dose
100 kGy
HL-LHC
16 kGy
LHC
remote DAQ
sensor

## Slide 4
INTERFEROMETRIC
HLS SENSOR

## Slide 5
TESTS IN THE
LABORATORY
RANGE AND LINEARITY
range of 10 mm
5 ml
10 ml
20 ml
5 ml = 249 µm
repeat & multiples
within ± 1 µm
INCLINATION TOLERANCE
3 sensor network: 2 horizontal, 1 inclined
tested at four azimuths, max 1.3°
injection and extraction of liquid

## Slide 6
Laboratoire Souterrain
à Bas Bruit
LSBB
interferometric HLS
(iHLS)
HLS-LINES
separated water /air system

## Slide 7
Laboratoire Souterrain
à Bas Bruit
LSBB
STABILITY
both sensors see same variation
differences of ~5 µm – coupling to rock
different measurement principle
none of them can be classified to be better

## Slide 8
capacitive HLS
(cHLS)
interferometric HLS
(iHLS)
main hydraulic network
Transfer Tunnel 1 (TT1)
CERN

## Slide 9
Transfer Tunnel 1 (TT1)
CERN
SENSITIVITY LINEARITY
injection of 20 ml water
both sensor types see 3 µm
shift in wavefront corresponds to position
extraction and injection of water to main network
difference between cHLS and iHLS ~8%
hypothesis: protective oil layer on iHLS causes shift
Δh ~8%

## Slide 10
Transfer Tunnel 1 (TT1)
CERN
STABILITY
instability on central iHLS sensor
stability fixed after cleaning of sensor
  dust on optics ?
  mechanical problem ?
  tilt of support ?
earth tides on all sensors – coherent with theory
linearity difference on measurement data confirmed
EAST support is subject to true movements

## Slide 11
SUMMARY
3 years: idea, design and development
laboratory:
range, linearity calibration
inclination tolerance
two test installations:
stability
comparison to other sensors
linearity scale factor
complies to most specification requests
PERSPECTIVES
Relative use already possible for tilt
monitoring applications in geopyhsics
CERN: HL-LHC
absolute sensor (external, geodetic reference)
suppress protective oil layer, cancel scale factor
configuration with more sensors
radiation tests on system components

## Slide 12
to all persons involved in the project
Thank you for your helping hands, brilliant ideas and your understanding!