## Slide 1: The Electron Beam Dump of the FERMI FEL
Simulation & measurement of electromagnetic and neutron radiation fields
The FERMI free-electron laser
Geometry
Scoring & results
Optional: Magnetic fields
Map:
Wikimedia Commons
Lars Fröhlich (DESY) ● Katia Casarin, Alessandro Vascotto (Elettra—Sincrotrone Trieste)

## Slide 2: The FERMI Free-Electron Laser
FERMI is a seeded free-electron laser driven by a 1.5 GeV electron linac.

## Slide 3: The Geometry – In Broad Strokes

## Slide 4: The Main Beam Dump

## Slide 5: The Main Beam Dump

## Slide 6: Fluka Geometry
Lattices, for no reason at all.

## Slide 7: Measurement Positions
A–D Ionization chamber currents
A–D Neutron dose equivalent H*(10)

## Slide 9: Top View
Black Hole
Parking
Lot
(z ≈ 0 m)
Simulation
Geometry
(z ≈ 280 m)

## Slide 10: Parking Lot (Top View)
z = 0
x = 0

## Slide 11: Transforms

## Slide 12: Bodies
Clone
Prototype, around (100,0,0)
Clone
Prototype, around (0,0,0)
Translation here cannot be transformation!

## Slide 13: Clone Regions & Lattice Cards
No inverse (negative) transformations here!

## Slide 14: Scoring & Results
Why life is easier if you don‘t measure things.

## Slide 15: Scoring
Simulation only
3D USRBIN of ambient dose equivalent H*(10) (nice to look at)
USRBDX for photon spectra
USRBDX for neutron spectrum in position A
Simulation & measurement
Deposited energy in ion chamber gas (to calculate current)
H*(10) from neutrons in position A
500 pC bunches
10 Hz
(5 nA)
y-avg. ± 20 cm

## Slide 16: Detectors
PTW 32002 spherical ionization chamber
Active air volume: 1182.7 cm³
Outer wall: 3 mm polyoxymethylene
Inner sphere: 10 cm  polystyrene
Keithley 6517 electrometer
Framework Scientific ABC1260 neutron detector
Superheated drop detector
Acoustic detection mechanism
Active volume: 10 cm³ SDD-100 vial
Calibrated vs. H*(10) with AmBe

## Slide 17: Ionization Chamber Current
Ion chamber model in simulation
Score deposited energy in “enclosed air” region
Calculate released charge per primary (34 eV for 1 e−-ion pair)
Scale to bunch charge/current
Hard to reach good statistics  LPB everywhere except target region; weak importance biasing
Inner electrode
Outer electrode
Enclosed air

## Slide 18: Neutron Dose Equivalent
Neutron detector in position A
Score DOSE-EQ in detector position (in air, no model present)
Scale to bunch charge
Compare to measurement

## Slide 19: Energy Spectra
2 runs with special transport & production cutoffs
Detector model consists of air
USRBDX at outer detector sphere
Fluence scoring
Isolethargic plots: Area proportional to number of particles
Photons
Neutrons

## Slide 20: Conclusions
Simplified geometry
Some estimated material properties (concrete composition, graphite density)
For practical purposes: Good agreement with measurements
Nice tool to understand radiation fields around the accelerator
Extremely helpful for “what if” scenarios
L. Fröhlich, K. Casarin & A. Vascotto,
Simulation and measurement of the radiation field of the 1.4-GeV electron beam dump of the FERMI free-electron laser,
Radiation Protection Dosimetry (2014) doi:10.1093/rpd/ncu138

## Slide 21: Spare Slides
… or: What did not fit in elsewhere.

## Slide 22: Energy Dependence PTW 32002

## Slide 23: Energy Dependence SDD-100
(kinetic energy)

## Slide 24: Magnetic Fields
One magfld.f to rule them all.

## Slide 25
… more problems
challenges!

## Slide 26: Not just a title…

## Slide 27: Magnet setup by machine and region name

## Slide 28: Not so nice: Lattices

## Slide 29: Field calculation: More fun with lattices