## Slide 1: Medipix3 Irradiation Studies
Irradiation Studies of a 130nm Large Area Pixel Chip
Richard Plackett – CERN Medipix Group
TWEPP 2009, Paris, 22nd September

## Slide 2: Overview
Introducing the Medipix3 Chip
Updates from Medipix2
Charge Summation
Spectroscopic mode
Example Application
Brief Radiation Damage Overview
Specific Design Issues
Irradiation Measurements up to 460Mrad
Threshold, Noise and Gain Results
DAC Stability
Irradiation Measurements up to 3Mrad
Additional Run in the Worst Case
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation
The Medipix3 Design team is Xavier Llopart, Rafael Ballabriga & Winnie Wong
These Measurements taken and analysed by myself, Xavier Llopart & Rafael Ballabriga
Thanks to them and all the CERN group for their support

## Slide 3: Hybrid Pixel Detectors
sensor
Analogue amplification
Ionizing Particle
e-
h+
Digital processing
Chip read-out
Positive or negative sensor bias
Medipix3 adds communication between the pixel analogue electronics
10001110101
A hit in the sensor deposits charge
The charge passes to the analogue amplifiers
The Medipix 3 charge summing circuit operates
The counter is iterated and read out with the shutter
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 4: Imaging by Photon Counting
Application of HEP idea to x-ray detection and medical imaging, the concept behind Medipix chips.
Photon counting devices provide superior contrast and dynamic range to charge integrating devices.
Counting only hits that pass threshold rejects noise in the sensor giving clearer images.
Dynamic range is limited only by front end response and depth of in-pixel counters.
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 5: Medipix 3
Medipix3 builds on the success of Medipix2 as a single photon counting imaging chip
Added Features
  Analogue charge summing to keep all charge information
  Spectroscopic mode with 8 threshold levels
  Continuous readout mode (no dead time)
  Increased counter depth increasing dynamic range
  Increased readout speed
  Increased radiation hardness from 130nm CMOS process
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 6: Status
First engineering run (12 wafers of 100 chips) delivered early this year
Wafer probing complete (from 11 wafers – 437 Class A + 166 ClassB)
Initial readout system working at low speed
Initial characterisation underway
One wafer diced and bonded to PCBs
First bump bonded assemblies with wafers expected soon
Initial radiation hardness measurements underway
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 7: Charge Summing
10001110101
A hit in the sensor deposits charge across four pixels
The analogue comparators assign the charge to the pixel with the most hits
This prevents ‘lost’ charge by partial hits not passing threshold
The threshold is applied to the summed charge and read out when the shutter closes
Medipix3 can be set to sum charge across four pixel clusters to prevent hits being lost due to charge sharing
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 8: Spectroscopic Mode
By connecting 4 pixels into a larger super pixel, eight threshold levels are available to us in the digital part of the pixel…
Counter 1
Counter 2
Counter 3
Counter 4
Counter 5
Counter 6
Counter 7
Counter 8
Amplifier response
Each threshold is adjustable, allowing a wide range of settings.
This gives us enough flexibility to capture reasonable spectra in many different applications
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 9: MARS Computed Tomography System
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 10: Radiation Damage
130nm is expected to be significantly more radiation hard than 250nm CMOS because of the much thinner gate oxides
Radiation damage occurs because the regular crystal structure of a device is disturbed. This causes a number of effects among which are:
  Change of effective doping concentration (esp. in diodes)
  Increase leakage current (esp. in detectors)
  Charge trapping (esp. in detectors)
  Oxide charging (esp. in CMOS)
  Single Event Upset (in logic circuits)
Photons and hadrons cause different types of damage, point and cluster respectively.
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 11: Analogue Switch Problem
One of the DAC outputs is particularly sensitive due to a minor design oversight
The current drawn by some minimum sized NMOS transistors when the matrix is irradiated can overload the DAC output stage stopping the pixel front end from functioning under nominal bias conditions
Here it happened at less than 1MRad so we were able to take no proper front end measurements during the 400MRad irradiation
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 12: Solution
The chip operation can be recovered by modifying the DAC settings.
In particular the pre-amp reference voltage was reduced modifying the front-end operating point.
Next wafer production fix Cas DAC issue by eliminating the (unneeded) leaky NMOS
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 13: 400Mrad Irradiation
Used a calibrated X-ray machine (Seifert RP149)
Beam profile is smaller than the Medipix3 → Two runs:
On the Pixel Matrix 60Mrad
Threshold Variation
Gain Variation
Noise Increase
On the Periphery 400Mrad
Check DACs
E-fuses
Logic functionality
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation
60 MRad
460 MRad
400 MRad

## Slide 14: Performance after 460MRad
Threshold
Noise
Yield artifact
σ=1.72 ke-
µ=9.3 ke-
σ=12.9 e-
µ=71.6 e-
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation
Threshold can be re-tuned using 5 bit equalisation

## Slide 15: Performance after 460MRad
Yield artifact
Qin=2ke-
Row[0:255]
After 460MRad there is essentially no gain variation observed
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 16: DAC Measurement
Band Gap
Accumulated dose of 396 Mrad
Accumulated dose of 396 Mrad
9mV
NMOS DAC (Preamp)
Accumulated dose of 396 Mrad
33mV
PMOS DAC
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 17: 3MRad irradiation
3 MRad
Worst effect at 3MRad, After this the effect of further radiation is much less.
Irradiated a small area of the chip up to 3Mrad to try and keep Cas DAC working by only damaging a limited number of pixels
The main effect observed is the shift in the threshold and amplifier response caused by the DAC moving…
The front end leaves its normal operating point at nominal bias conditions at 1500krad, although it remains possible to read out the chip up to 3Mrad. The front end can be recovered by increasing the value of the IKRUM DAC.
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 18: 3MRad Noise and Gain Result
Yield artifact
Qin=2ke-
Row[0:255]
In this worst case situation the noise is still less than 100e- and the gain variation is still minimal
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 19: Next Steps
Repeat 400MRad measurement with bare chips and full sensors, recovering performance at each point to take measurements.
Begin hadronic measurements rather than x-rays, possibly in collaboration with LHCb upgrade projects.
Measurements with different sensors eg 3D or diamond
Next wafer production to fix Cas DAC issue by eliminating the (unneeded) leaky NMOS
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 20: Conclusions
Medipix3 has been operated successfully after an exposure to a very large x-ray dose.
This is very encouraging for readout chips in SLHC trackers and other high radiation environments.
Confirmation in a full chip that 130nm is intrinsically more radiation tolerant than 250nm.
You still need to bear radiation damage in mind when designing chips to avoid unexpected problems.
Thank you for your attention.
Introduction to Medipix3
Radiation Damage
460MRad Irradiation
Design Issues
3MRad Irradiation

## Slide 21: Backup Slides
This page has been left intentionally blank

## Slide 22: 460MRad Threshold variation
After annealing we see ~20 THL DAC steps variation (<1ke-) between areas non irradiated and areas irradiated at 60, 400 and 460 Mrad.
This variation can be corrected by the threshold equalization procedure (5 bits)
Row[0:255]

## Slide 23: General behavior after 3 Mrad
Threshold
Noise
Yield artifact

## Slide 24: THL Shift at 3MRad
Main contributor is the protection diode. In order to minimize the pixel to pixel threshold variation Ikrum is set to 16nA which indicated that the leakage of this diode at 3Mrad is ~5 to 10 nA.
This is the worst radiation operation point.
Row[0:255]

## Slide 25: Difference in CAS DAC range
Good chip (F7) shows CAS DAC dynamic ranger larger than other chips → Less ids transistor leakage in analog switches?