SALT128 Testing
Carlos Abellan
Milano, May 18th 2016

---
SALT128 Test setups
Two main branches:
  - SALT128 prototype testing
    - PCB similar to the existing one
    - General purpose
    - Compatible with VLDB, COMET, Flex, standalone
    - Immediate production, first priority
  - SALT128 production testing
    - To be used with a probe station in the wafer (before dicing)
    - Automate testing as much as possible
    - Specific purpose
    - Intend to construct it around COMET board
    - Interaction with probe station, analysis software, results database, etc
    - Need to design tests (analogue performance, digital functionality, switching coverage, pll performance, PSRR, ADC performance, etc)

---
SALT128 Prototype TB
Input I need:
  - ASIC: New pinout, exact dimensions, detailed pin functionality
  - Detailed list of tests to be done: sort out the changes if any (I heard common mode?)
  - Feedback: is there anything we would like to change/improve?
  - List of cards to deliver: Who wants them? How many? When?
Production process only when all input is available: Today t=0.
  - 3 or 4 weeks for design (inc. circulation, review, …). Depends on the changes to make
    - 1st or 2nd week of June ?
  - 2 weeks for PCB production (standard procedure) 3rd or 4th week of June ?
    - 1 week express: need to check for availability, very expensive (x2)
    - Component sourcing in parallel. It can be slower than 1 week.
  - 1 week for component soldering (once they start...)
    - Yes, but be aware of availability issues. 1/2 weeks waiting? Outsourcing it?
  - 2 days for wire bonding (once they start...)
    - Outsourcing/Volunteers? Marina volunteered yesterday.
    - Ship to Syracuse by mid-July?

---
Input I need:
  - ASIC: New pinout, exact dimensions, detailed pin functionality (more or less)
OUTDATED VERSION!

---
Input I need:
  - ASIC: New pinout, exact dimensions, detailed pin functionality (more or less)
OUTDATED VERSION!

---
Schematics well advanced:

---
Schematics well advanced. Main changes:
- Less regulators: no one used them anyway...
- Power consumption reading now COTS (MAX4173)
- ASIC now uses 6 e-links instead of 3: where should I map them?
- Profit to test the ADC that will be used for production testing. The e-link configuration is compatible with the ones for SALT.
  - We need to sample 12 channels!!
    - Since it's DC, maybe a relay?

---
ADC Testing:
  - We use a MAX144 2 channel 12 bit @ 100KSPS ADC to sample a 5 bit DAC (could seem like an overkill)

---
ADC Testing:
  - We use a MAX144 2 channel 12 bit @ 100KSPS ADC to sample a 5 bit DAC (could seem like an overkill)

---
ADC Testing:
  - We use a MAX4238 (Zero drift) together with precision resistors.
  - x2 gain and only used from 0.2V to 0.8V gives better linearity. Uses a compensated V source to calibrate it further.
    - This gives a [0.4, 1.6] output range (smaller than ADC dynamics, but we have many bits left).

---
SALT128 Production TB
Production process similar to the proto_TB one:
  - [New design (I don't dare guesstimating) + 2 weeks for PCB production (standard procedure)+1/2 weeks for component soldering (once they start...)] * 1.5 safety factor
    - Prototype Test Board (a big part can be reused) +
      - ADC to sample and calibrate internal DACs (+ OP amps, + ...)
      - Precise power consumption measurement
      - Any other requests? (Depends on what you want to test) Now or never?
  - Probe card: I never designed one. I need to learn the basics, find a company that does it, get time and price estimations, … Anyone with experience can help me with that?
    - I talked to the people who designed the test system for, some conclusions:
      - Their card has 110 needles (to have an order of magnitude): (We want 247!!)
        - 5x Boards = around 12k€
          - We really want to be sure it's perfect before submitting
        - Delivery time estimated in 9 weeks once everything is accepted
          - We have to foresee a pipeline not to be stopped while we wait
        - The manufacturer keeps some manufacturing information private
          - We give instructions and they give back closed products, will require several iterations to make sure there is no communication failure (we only want to do it once!)

---
SALT128 Production TB
Input I need:
  - Results of the prototype testing: will drive the production tests
  - Full-featured COMET board + adequate mini-HDMI daughter boards.
  - Detailed list of tests to be done, to plan for:
    - Analysis software: Chris Betancourt, Iaroslava Bezshyiko (from Zurich). Online data analysis, report generation, result management, ...
    - Hardware design: automate the measurements (Carlos)
      - Probe card + reuse as much as possible the proto_TB design
    - Probe station interface (MariaPilar)
      - We never operated one. Will require some time to learn.
    - System wide design:
      - Chuck? Wafer manipulation? Clean room equipment? ESD cautions?
    - HHRR, logistics, coordination with stave production...
      - Fact: 1 second of test/ ASIC x 6kASICs = 1.67h of testing
        - 10 mins of test / ASIC = 25 working weeks = 6.25 months
          (8h/day, 5 day/week)

---
SALT128 Production TB
What we're aiming for:

---
SALT128 Production TB
What we're aiming for:
Clean Room
Coordinate Tests
Online Data Analysis
Control SALT128
Probe card movement
Handle a database
User Interface
…
Power
Supplies,
etc...
Probing Station
DAQ
Probe Card
Wafer

---
Conclusions
For the prototype testing:
  - Where should I put the extra pairs? (Fit cable mapping, ...)
  - Will upload schematics to Twiki.
  - Need to fix a date for the schematic review.
  - A couple of weeks for routing + PCB review
  - Production delays
- For the production testing:
  - Any help from experienced people will be highly appreciated
  - I already contacted other groups at CERN who did it, looks like it will take some work
    - 9 weeks production delay!!
    - Very expensive, no room for failure... (will be even slower)
  - Full system design: analysis SW, control SW, HW tools, procedures, …
  - Detailed list of tests to be done
  - We're going to be busy quite some time with that.