## Slide 1: CLICdet_2015 Detector Model
CLICdp-Draft-2015-019
Section 6, ECAL optimization
Proposal: reduce the number of layers from 30 to 25
How did we come to this proposal?
What is wrong with it?
Why is CLIC/ILC photon energy resolution poor?
ILC TDR σ(E)/E = 0.01 ⊕ 0.175/√E
ATLAS σ(E)/E = 0.007 ⊕ 0.10 /√E
Can one improve our energy resolution?
Discussion

## Slide 2: CLICdet_2015 Detector Model
Fig. 19, Jets, σ(E)/E; no degradation when ↘ nLayers
from 30 to 25 => reduce the number of layers.

## Slide 3: CLICdet_2015 Detector Model
Fig 17, 10 GeV γ;σ(E)/E ↗ when nLayers ↘
Is it dramatic?
30 layers, SiW, σ(E)/E=5.6%
σ(E)/E≈0.177/√E
25 layers, SiW, σ(E)/E=6.0%
σ(E)/E≈0.19/√E
(modest increase)
15 layers, SiW, σ(E)/E=8.0%
σ(E)/E≈0.25/√E
Expect 0.177x√2=0.25 (ok)
High E photons, “σ(E)/E not driven by sampling term”.
Photons, σ(E)/E is not a major physics driver.
=> Can reduce the number of layers from 30 to 25

## Slide 4: CLICdet_2015 Detector Model
Steven studied σ(E)/E as a function of nLayers for 100 GeV γ and 5x5 mm2 cells in SiW ECAL.
30 layers, σ(E)/E=1.8%
σ(E)/E = 0.18/√E
Same value as for 10 GeV γ
25 layers, σ(E)/E=3.3%
σ(E)/E = 0.33/√E
Significant σ(E)/E degradation
Why?
15 layers, σ(E)/E=5.3%
σ(E)/E = 0.53/√E
Expect 0.18x√2=0.25
Why so large?
Calibration problem?

## Slide 5: Origin of poor σ(E)/E and of degradation when ↘ n layers
Calorimeter talks
http://www.desy.de/~garutti/LECTURES/ParticleDetectorSS12/L10_Calorimetry.pdf
σE/E(sample) ~ 1/√Nch
Nch=number of charged particles crossing the active layer
Nch=E/t; t: absorber layer thickness in Xo
σE/E(sample) ~ √t/√E
=> The resolution is ~ absorber layer thickness.
The smaller the thickness t, the larger the number of times the shower is sampled by the active layers (and the larger the number of detected particles), the better the energy resolution

## Slide 7: Shower Max and Containment
Model CDR:20+10 layers. Arrows indicate the two regions: first thin plates 2.1mm, then 4.1 mm.
10 GeV γ a large fraction of shower sampled by thin plates
100 GeV a larger fraction of shower sampled by thick plates

## Slide 8: Origin of poor σ(E)/E and of degradation when ↘ n layers
σE/E(sample) ~ K/√E
K=0.032 . √(Ec. t/F)
Ec=critical energy, F=Fano parameter, t in units of (Xo)
 | t (mm) | Ec [MeV} | F | t (Xo) | K(%)
CLIC (W) | 2.1 (and 4.1) | 8.1 | 0.115 | 0.64 | 21
ATLAS (Pb) | 2. | 7.3 | 0.115 | 0.36 | 15
 | t (mm) | Ec [MeV} | F | t (Xo) | K(%)
CLIC (W) | 3 | 8.1 | 0.115 | 0.86 | 26
CLIC (W) | 4 | 8.1 | 0.115 | 1.14 | 29
Assuming same Fano value; the formula gives the right order of magnitude
The mix of 2.1 and 4.1 mm plates may explain why CLICdet is worse than ATLAS
Absorber 23 Xo; Reducing the number of layers => t (Xo) ↗ => K ↗
=> Energy resolution degradation when reducing the number of layers

## Slide 9: e⁺e⁻-> h(750) -> γγ [CLIC at 3 TeV]
Left plot σ(mγγ) for e⁺e⁻->h(750)->γγ; ATLAS, CLIC σ(E)/E
At 3 TeV <Eγ> ~ 450 GeV
σ(E)/E is a major driver for di-photon resonance studies
=> Must reduce σ(E)/E for photons

## Slide 10: High Energy Photons
NB. remark by Philipp:
at e.g. E=450 GeV, the constant term of 1% already dominates over the 20%/√E
Answer by JJB:
not quite; at 400 GeV the stochastic and constant term would be equal (1%), which implies that the stochastic term still considerably worsens the energy resolution (twice 1% added in quadrature…)

## Slide 11: Discussion
Which other tests to understand the origin of the energy resolution degradation when reducing n layers?
How to improve the photon energy resolution?
(need to study 200, 500 GeV photons ?)
What else?

## Slide 12
Backup slides

## Slide 13
The plots shows sigma(E)/E vs E; Sigma(E) is Rms90