## Slide 1: BNL MAGNET TEST FACILITY
Joseph F Muratore
Superconducting Magnet Division
Brookhaven National Laboratory
Upton, NY 11973 USA
SC Magnet Test Stands Workshop CERN June 13-14, 2016

## Slide 2: BNL MAGNET TEST FACILITY
Cryogenics Facility
2. Power Supplies and Upgrade for MQXF Tests
3. Redesigned Vertical Test Dewar for 1.8 K Magnet Tests
4. Data Acquisition
5. Quench Protection
6. Magnetic Field Measurements

## Slide 3: BNL MAGNET TEST FACILITY
BNL Facility North (High) Bay Area (30 t crane)
Test Facility in Background

## Slide 4: BNL MAGNET TEST FACILITY
Building 902 Test Facility Floor Plan and Overview
Vertical Test Dewars 2-6
Horizontal
Test
Stands A - E
CONTROL ROOM
MQXF
30 kA Power Supply
2.7 m Deep Test Dewars
6.1 m Magnet Deep Test Dewars
CTI 4000 He Refrigerator
CVI Air Liquide He Refrigerator

## Slide 5: TEST FACILITY UPGRADES FOR MQXF
The former cable test 30 kA power supply (2 x 15 kA) reconfigured to test magnets to 24 kA (2 x 12 kA) with new IGBT switch and energy extraction circuits.
1.8 K vertical magnet test dewar (4.2 m useful depth) with new 24 kA top plate, newly designed LHe vessel, 4.5 K heat shield, lambda plate, heat exchanger, and 24 kA water-cooled cables
  Diagnostic and monitoring instrumentation with PLC for CTI Model 4000 refrigerator (main liquefier) and CVI refrigerator (backup liquefier), along with hardware upgrades.
  New data acquisition, monitoring, and control software (NI LabVIEW)
  Upgraded data acquisition loggers and scanners for speed and precision
  Added new backup compressors
  New dirty gas compressor and purifier
  Other upgrades
SUMMARY

## Slide 6
BNL CRYOGENICS FACILITY
BNL Magnet Division has a large infrastructure of cryogenic test equipment, but:
  The cryogenic test facility has been operational and in use since the 1970’s, supporting SSC, RHIC, LHC, etc.
  Despite repairs as needed, equipment reliability decreases with age.
    BNL management supported funding the purchase of critical spares, preventive maintenance of critical components, and installation of backup systems to improve reliability and reduce risk for the Hi-Lumi production magnet testing program (RISK MITIGATION).
Previous 1.8 K testing was limited to superconducting cable samples (some years ago), not magnets.
CRYO CHALLENGES

## Slide 7
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL VERTICAL CRYOGENICS FACILITY

## Slide 8
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL VERTICAL CRYOGENICS FACILITY
1.8 K Process Flow Diagram

## Slide 9
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL CRYOGENICS FACILITY
CTI MODEL 4000 REFRIGERATOR/LIQUEFIER
Primary source of liquefaction
1500 W at 4.5K
Two reciprocating expansion engines rated at 250 rpm
70 g/s (150g/s with Creare circulator cold pump) forced flow capacity at 5 - 12 atm in HTF-Magcool
He liquefaction: 320 L/hr (with both expansion engines and a Koch Model 1600 Wet Expander running)
Recently upgraded with a new LN2 heat exchanger, doubling the number of inline purifiers to four, rebuilt expansion engines, and new diagnostic and control software (LabVIEW).
This work was funded by BNL.

## Slide 10
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL CRYOGENICS FACILITY
CTI Model 4000 Refrigerator Reciprocating Expansion Engines
250 rpm
Each has two cylinders: 9.1 cm bore; 7.6 cm stroke
Largest in world (we think)
Recently rebuilt.
This work was funded by BNL.
Koch Model 1600 Wet Expander
Has two cylinders
60 rpm
7.6 cm bore; 5.1 cm stroke
Largest in world (we think)
Recently rebuilt

## Slide 11
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL CRYOGENICS FACILITY
CVI – AIR LIQUIDE He LIQUIFIER
Secondary source of liquefaction
1000 W at 4.5K (used at 700 W)
Two turbine expanders rated at 130000 rpm
In-line gas purifier
He liquefaction: 160 L/hr
Total facility LHe rate (both liquefiers): 480 L/hr
Recently upgraded for use as a back-up to mitigate risk in the event of CTI Model 4000 failure:
1) New first stage LN2 – He heat exchanger
2) New chilled water heat exchanger
3) New and upgraded relief and isolation valves
4) Re-plumbing and new piping to increase efficiency
5) New and refurbished vacuum pumps
6) New instrumentation (including temp sensors and PLC controls) for turbine protection
This work was funded by BNL.

## Slide 12
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL CRYOGENICS FACILITY
COMPRESSORS (Pt 1)
Mycomm 800 (2-Stage)
PRIMARY UNIT
Power = 597 kW
Discharge pressure = 18 bar
Delivers 160 g/s
Purchased critical spares (BNL funded):
New 2nd stage compressor head
New 400 hp motor
New oil heat exchanger
New helium heat exchanger
All maintenance & refurbishment items completed
Sullair 500
Power = 373 kW
Discharge pressure = 18 bar
Delivers 80 g/s
Provided and installed by BNL funding.
Sullair 350
Power = 261 kW
Discharge pressure = 18 bar
Delivers 51 g/s
Purchased critical spares (BNL funded):
new compressor screw
new oil pump
Can supply either the CTI Model 4000 refrigerator/liquifier and Magcool, or the CVI liquifier
All-purpose; presently configured to work with Mycomm to provide 210 g/s delivery. Can also be used with Sullair 500 as a backup in the event that Mycomm is down. (Risk Mitigation)
All-purpose; recently installed & can be configured to work with Sullair 350 as a backup in the event Mycomm down.
Mycomm and Sullair 350 compressors
Sullair 500 compressor

## Slide 13
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL CRYOGENICS FACILITY
COMPRESSORS (Pt 2)
Gardner Denver 60
Power = 44.7 kW
Discharge pressure = 17 bar
Delivers 213 m3/hr  9.65 g/s
New Gardner Denver compressor
purchased with BNL funding.
Sullair 100 (3 units: CS4, CS5, CS6)
Power = 74.6 kW
Discharge pressure = 12 bar = 1.2 MPa
Delivers 20 g/s (80 g/s @ 8 bar input)
New Sullair 100 compressor
purchased with BNL funding.
Linde Model 500A He Gas Purifier
with dedicated compressor
All new system and commissioning contract purchased with BNL funding.
Creare circulator pumps
Presently, only CS4 (new unit) being used, for pumping Nash 1.9 K station outlet to high pressure gas storage tanks. All three are used for horizontal testing (Magcool facility).
Presently configured for Magcool operations:
(1) Nitrogen cooldown (Cooldown I)
(2) Warmup
(3) Pump and purge
Pumps ‘dirty’ He gas to dirty gas bag and gas recovery tank.
Cleans and recovers ‘dirty’ gas.

## Slide 14
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL CRYOGENICS FACILITY
Nash-Kinema Vacuum Pump
Power = Two stages at 74.6 kW each
Suction pressure = 0.0837 bar @ 4.5 K
Discharge pressure = 1.2 bar
Delivers 2.7 g/s @ 1.8 K
Cooling capacity = 40 W @ 1.8 K
Pumps on LHe in heat exchanger - down to 16 mbar for 1.8 K operation.
Upgrades included:
1. Reconfigured exhaust. Connected helium exhaust gas line to new Sullair 100 compressor and inline purifier, to reclaim helium gas (was previously routed to dirty gas facility).
2. Replaced legacy soldered copper water lines with welded stainless steel.
  3. Purchased new spare booster and liquid ring vacuum pumps.
4. Refurbished existing / purchased new spare chilled water heat exchangers (BNL funded).
5. Replaced drive belts.
6. Repaired service air supply line.
7. New LabVIEW controls with remote HMI panel.
8. Installed new vacuum jacketed transfer line, valve to
test dewar.

## Slide 15
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL CRYOGENICS FACILITY
LIQUID HELIUM STORAGE DEWARS
  Storage Dewar #1 3785 L
  Storage Dewar #2 3785 L
  Storage Dewar #3 10000 L
      Total liquid helium storage capacity is 17570 L.
Total He gas storage capacity is approximately 8.5 x 106 L.
Liquid nitrogen storage capacity is approximately 40000 L.

## Slide 16
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL CRYOGENICS FACILITY
VERTICAL TEST DEWARS
Present Configurations:
Dewars 2 and 3: for longer magnets (4.2m)
Dewar 2 (610mm): now updated and configured for 1.8 K/1 bar operation with lambda plate, new 24 kA leads and water-cooled cabling, and hanging fixture
Dewar 3: wider (711 mm) for larger magnets 4.5K
Dewar 4: 4.5K cable testing
Dewar 5: 1.8K/1 bar cable testing
Dewar 6: common coil type magnets (Nb3Sn and HTF) and other short magnets (2m); cable testing

## Slide 17
SC Magnet Test Stands Workshop CERN June 13-14, 2016
PRIMARY POWER SUPPLIES
(1) 30kA (2 x 15 kA) @ 14VDC Previously cable testing, now configured 24 kA (2 x 12 kA) MQXF operation at 1.8 K
(2) 10kA @ 20VDC “HTF”, 10VDC, 30VDC, 40VDC, 50VDC)
(3) 8.5kA (9.2kA) @15VDC “VTF”, 30VDC
(4) 8kA @ 12VDC “Short sample test facility magnets”
(5) 10kA @10VDC bipolar
(6) 7.5kA @ 25VDC “Fast Ramp”, 50VDC, 100VDC (GSI ramp rates to 7800 A/s)
(7) Dual Acme 5500A @ 125V “Room temperature field measurements and calibration magnets for measuring coils”
(8) Variety of trim power supplies (bipolar and unipolar)
All power supplies can be configured for a wide range of magnet tests.
Energy extraction is available with

## Slide 18
SC Magnet Test Stands Workshop CERN June 13-14, 2016
30 kA POWER SUPPLY UPGRADE

## Slide 19
SC Magnet Test Stands Workshop CERN June 13-14, 2016
30 kA POWER SUPPLY UPGRADE
15 kA power supply switch cabinet with IGBT and snubber circuits. Each power supply is configured for 12 kA with 6 water-cooled IGBT and snubber circuits, allowing 24 kA magnet operation.
Plumbing and flow gauges with signal acquisition for IGBT water cooling and monitoring.
IGBT data acquisition and instrumentation panel
Existing three water-cooled free-wheeling diodes in parallel.
NEW: Upgrades to previous cable sample 30 kA power supply ( two 15 kA in parallel) to operate magnets to 24 kA.

## Slide 20
SC Magnet Test Stands Workshop CERN June 13-14, 2016
30 kA POWER SUPPLY UPGRADE
Power Supply Energy Extraction Switch
24 kA IGBT-based Energy Extraction Switch:
  Switch consists of 2 modules, each with six high power IGBTs in parallel. (One module of 6 for each 15 kA power supply.)
  IGBT rated to 3600 A per unit & at least 1200 V C-E voltage.
  IGBT current & C-E voltage will be limited 2000 A and 840 V.
  Stray and mutual inductances in the power supply circuit generated much higher switching transients than expected.
  So each IGBT has been equipped with a snubber circuit designed to limit high voltage switching transients to 600 V.
  40 channel IGBT Switch data monitoring and logging system for each module: Lab View program debugged and tested.
24 kA water-cooled cables: Finished DI water cooling circuit, installed drain outlet, and verified flow (>68 L/min) and pressure drop. 15 kA water-cooled cables flow is 30 L/min.
IGBT
Rated collector current = 3600 A
Max operating current = 2000 A
Max C-E voltage per IGBT = 840 V
(These values are with 66 mΩ dump resistor.)
Energy Extraction
66 mΩ center tapped to ground
Stainless Steel
1 MJ energy absorption rating
Existing all stainless steel 66 mΩ dump resistors will be replaced by new high power ceramic 66 mΩ 10 MJ resistors for MQXF tests.

## Slide 21
SC Magnet Test Stands Workshop CERN June 13-14, 2016
1.8 K TEST DEWAR ASSEMBLY
New Top Plate
Lambda Plate
Redesigned Inner He Vessel
Existing Outer Dewar
MQXF Magnet
4.5 K Heat Shield
Top is 38.1 mm thick G10 (lower heat load)
Bottom is 19.05 mm thick 304 stainless steel (for added strength and sealing to He vessel interior flange)
Bonded together with Stycast 2850FT
Total heat load < 1.4 W
New top plate to accommodate lambda plate and fixturing for MQXF 4.2 m quads and 24 kA vapor-cooled leads and heat exchanger for 1.8 K operation.
Rated at 5.86 bar
Components below lambda plate and internal to heat exchanger are designed for 5.86 bar,
above lambda plate for 3.45 bar.

## Slide 22
SC Magnet Test Stands Workshop CERN June 13-14, 2016
1.8 K TEST DEWAR ASSEMBLY
+200 mm extended helium dewar bottom
Heat exchanger in magnet helium passage (left) and outside Mirror O.D. (right)
Upper He vessel
Lambda plate sealing detail
Lower He vessel
Outer 4K heat shield
He vessel flange
G-10/SST Lambda plate
2K-300K warm bore
4K Al heat shield
Overall hanging assembly, or insert (mirror shown)
Upper dewar assembly:
Vent and burst disc in lambda plate (for Safety Committee approval)
Gravity seal at lambda plate
Serviced by 30 t crane with 6.79 m hook clearance height above dewar flange.
Test Dewar Overview
Lambda Plate Spring-Energized Face Seal to He Vessel

## Slide 23
SC Magnet Test Stands Workshop CERN June 13-14, 2016
1.8 K TEST DEWAR ASSEMBLY
Top plate 3” (76.2 mm) thick 304 stainless steel
24 kA vapor-cooled leads
1.8 K Top Plate Details
(2) 24 kA main coil vapor-cooled leads
(2) 1 kA CLIQ vapor-cooled leads; port for 3rd lead provided
(3) instrumentation wiring ports
(2) helium fill lines
Heat exchanger ports, manifolds and solenoid valves under plate
Warm bore port
Helium vent port
Heat exchanger
CLIQ Leads
Connector Trees
3rd CLIQ Lead
He fill lines
(if needed)
Quench Heater Connector Tree

## Slide 24
SC Magnet Test Stands Workshop CERN June 13-14, 2016
1.8 K TEST DEWAR ASSEMBLY
4.5 K Heat Shield
Helium supply from 4.5 K liquid
Helium return in gas above 4.5 K liquid
SST manifolds, pipes
Flexible Cu connections
to Al
“blue/green” sections of helium vessel and lambda plate flange welded as a subassembly to permit precise, flat machined surface of lambda plate flange after welding (also locally thicker sections to accommodate bending stresses)
Heat shield is bolted to outside of lambda plate flange (supports weight, provides conductive cooling)

## Slide 25
SC Magnet Test Stands Workshop CERN June 13-14, 2016
QUENCH PROTECTION
The quench detection system is based on a redundant scheme which relies on a number of voltage signals. When any one of the following signals reaches its specified threshold voltage after a specified validation time, a fast discharge is instigated (current decays through dump resistor). Typical threshold voltages and validation times are listed here.
Half coil voltage Vhalf1 - Vhalf2 50-100 mV 10 ms
Quarter coil voltage V1 – V2 50-100 mV 10 ms
Quarter coil voltage V3 – V4 50-100 mV 10 ms
Total coil voltage Vtotal 50-100 mV 10 ms
SC lead voltages VSCL+ and VSCL- 10-50 mV 5 ms
Splice voltages V12, V34, V24 8-10 mV 5 ms
Current Derivative Vtotal – L dI/dt 0.1-5 V 10 ms
There are also designed into the system a number of interlocks which also rely on various signals which, if they do not meet specified conditions such as voltage threshold or on-off state, will instigate a slow discharge, with the dump resistor out of the circuit. These include:
Vapor-cooled lead voltages 80-100 mV 500-1000 ms
Strip heater capacitor bank not charged
IGBT switch temperature or voltage too high
Others
QUENCH DETECTION
Note: In Nb3Sn magnets, current-dependent thresholds as high as several volts at low currents are necessary and are set in the software. This is to avoid false QD trips due to flux jump spikes.

## Slide 26
SC Magnet Test Stands Workshop CERN June 13-14, 2016
CONTROL, MONITORING, AND DATA ACQUISITION SOFTWARE
RECENTLY UPGRADED WITH ALL NEW SOFTWARE
  POWER SUPPLY CONTROL PROGRAMS (LABVIEW)
  INSTRUMENTATION CONTROL PROGRAMS (LABVIEW)
  DATA ANALYSIS PROGRAMS (LABVIEW / DIADEM) include:
    DATA HANDLING, PLOTTING, AND ANALYSIS
    MULTIPLE SIGNALS DISPLAYED
    SIGNAL MANIPULATION
    WAVEFORM CALCULATOR
    ∫(I2 dt) CALCULATION

## Slide 27
SC Magnet Test Stands Workshop CERN June 13-14, 2016
DATA ACQUISITION SYSTEMS
FAST DATA LOGGER DAQ
National Instruments PXIe-4300 (16 Units)
Number of channels = 8 differential, simultaneous sampling
Analog input = ±10 V, ±5 V, ±2 V, ±1 V
ADC resolution = 16 bit
Sampling rate = 250 kHz max per channel
Time scale resolution = 10 ns
TOTAL CHANNELS = 128 differential, simultaneous sampling
Usage = fixed (main) voltage taps and configurable (auxiliary) voltage taps
Status: Assembled and Tested
SLOW DATA LOGGER DAQ
National Instruments PXI-6289 (4 Units)
Number of channels = 16 differential
Analog input = ±10 V, ±5 V, ±2 V, ±1 V, ±0.5 V, ±0.2 V, ±0.1 V
ADC resolution = 18 bit
Sampling rate = 500 kHz max multichannel (31.25 kHz max)
Time scale resolution = 50 ns
TOTAL CHANNELS = 64 differential
Usage = monitoring of vapor-cooled lead voltages, splice voltages, SC lead voltages, temperatures, LHe levels, and other instrumentation
Status: Assembled and Tested
64 Channel Slow Logger
128 Channel Fast Logger

## Slide 28
SC Magnet Test Stands Workshop CERN June 13-14, 2016
DATA ACQUISITION SYSTEMS
QUENCH DETECTOR DAQ
National Instruments PXIe-6356 (1 Unit)
Number of channels = 8 differential, simultaneous sampling
Analog input = ±10 V, ±5 V, ±2 V, ±1 V
ADC resolution = 16 bit
Sampling rate = 1.25 MHz max per channel
Time scale resolution = 10 ns
TOTAL CHANNELS = 8 differential, simultaneous sampling
Usage = signals used for quench detection, such as half coil and quarter coil voltage differences, SC lead voltages, etc.
Status: Assembled and Tested
STRAIN GAUGE LOGGER DAQ
Agilent 34970A Data Acquisition/Switching Unit (2 Units)
Number of channels = 60 differential, scanning
ADC resolution = 22 bit max
Scan rate = 250 ch/s max
TOTAL CHANNELS = 120 differential, scanning
Usage = strain gauges
8 Channel Quench Detector
Note: For mirror test, a portable Cytec scanner system will be used. All 64 channels have been assembled and tested.

## Slide 29
SC Magnet Test Stands Workshop CERN June 13-14, 2016
MAGNETIC FIELD MEASUREMENT
Existing inventory of over 60 magnetic field measuring coils, with variety of lengths and radii:
External drive coils, various radii and lengths, some with warm bore tubes
Moles, 1 m long and smaller, various radii, with electric (warm) or gas-driven (cold) motors
Example: Coils 84 and 86 shown below are external drive coils suited for 8 cm aperture magnets
2 radial dipole bucking windings
2 radial quadrupole bucking windings
1 tangential winding
digital bucking
Other devices: Hall probes, fluxgate magnetometers, vibrating wire
Magnetic field measuring probe designed for MQXF will not be needed for the upcoming mirror test but will be available for the testing of the prototypes starting in Feb 2017.
COIL #84
COIL #86

## Slide 30
SC Magnet Test Stands Workshop CERN June 13-14, 2016
MAGNETIC FIELD MEASUREMENT
More examples of existing measuring coils
Mole RA-7
D0 Warm Mole 13 cm
External Drive Coil 76 8 cm
Warm DX Mole in AGS Corrector Magnet
Mole RA-3
Coil 96 2.54 cm
Coil 203 9-Windings 30 cm

## Slide 31
SC Magnet Test Stands Workshop CERN June 13-14, 2016
SUMMARY
The cryogenics facility refrigerators, compressors, and other peripheral devices have been refurbished and re-configured to test MQXF quadrupoles at 1.8 K and 1 bar (nom).
New compressors, purifiers, critical spare parts, and other updates were provided with BNL funding. These items have been purchased to provide backup systems and mitigate risks in the event of equipment failure.
The Magnet Division 30 kA power supply (formerly used for cable testing) and peripheral systems such as energy extraction and quench heater firing units are being converted to test magnets to 24 kA and at 1.8 kA.
A completely new and redesigned vertical test facility (in an existing outer vacuum vessel) is being built to accommodate 1.8 K testing of MQXF magnets, with new inner helium vessel, 4 K heat shield, 1.8 K heat exchanger, lambda plate, new top plate with 24 kA Cu leads, and 24 kA water-cooled cables to test dewar.
The old data acquisition systems have been replaced by completely modernized and improved NI hardware and LabVIEW-based software.
The updated and reconfigured 1.8 K vertical test facility commissioning will be done with a short sample of Nb3Sn cable instrumented with taps and a spot heater. The mirror magnet will be present on the hanging fixture but not connected to the power leads until the cable test has been successfully completed.
Magnetic field measuring probe, spike detector system, and quench antenna data acquisition system are yet to be built, but will not be needed for the upcoming mirror magnet test.

## Slide 32
SC Magnet Test Stands Workshop CERN June 13-14, 2016
SUMMARY
Cryogenics refrigerators, compressors, and other peripheral devices can be configured to provide for a variety of testing scenarios, including forced flow supercritical He or liquid He testing, and testing down to 1.8 K.
Magnet Division power supplies cover a wide range of testing applications for loads of various inductances and resistances, current and voltage levels, current ramp rates, and quench protection requirements.
Data acquisition systems are designed for flexibility to accommodate many magnet and instrumentation configurations.
There is a large inventory of magnetic field measuring coils of different sizes and design for a wide variety of applications.
Further upgrades of facility systems are in progress.

## Slide 33: THANK YOU FOR YOUR ATTENTION
Acknowledgements for contributions to this presentation: Mike Anerella, Sebastian Dimaiuta, Piyush Joshi, Paul Kovach, Andy Marone, Bill McKeon, Jesse Schmalzle, Dan Sullivan
SC Magnet Test Stands Workshop CERN June 13-14, 2016

## Slide 34
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BACK UP SLIDES

## Slide 35
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL VERTICAL TEST FACILITY
VERTICAL DEWAR TOP PLATES
  New top plate to accommodate MQXF 4.2 m LHC Inner Triplet quads and 24kA leads and lambda plate/fixture for 1.8K
  Magnet hanging fixtures and parts for a wide range of magnet types
  Variety of feedthroughs and instrumentation connector assemblies
  Gas-cooled and liquid-injection power leads for various power applications
  Associated temperature sensors and level probes with each top plate
  Presently eight top plates of various configurations available

## Slide 36
SC Magnet Test Stands Workshop CERN June 13-14, 2016
1.8 K TEST DEWAR ASSEMBLY
1.8 K Heat Exchanger
Top Plate Interface
Lambda Plate Interface
Components below Lambda plate are designed for 85 psi (5.86 bar) external pressure.
Crescent-shaped bottom cup
Outer vacuum jacket above lambda plate designed for 50 psi (3.45 bar) external pressure.
All internal volume components designed for 85 psi (5.86 bar).

## Slide 37
SC Magnet Test Stands Workshop CERN June 13-14, 2016
1.8 K TEST DEWAR ASSEMBLY
2” Fike burst disc designed for subcritical flow (API 520). Venting 4150 g/s helium at 5.86 bar to 3.45 bar backpressure.
Required area is 14.8 cm2. Provided area is 20.0 cm2.
Certified by Fike: 2.2 bar @
74.8 K (lowest they can certify to). Calculated allowance to LHe temperature.
Service (Quench) Valve
  Designed for 750 g/s with a pressure rise of under 0.34 bar @ operating conditions (approx. 1.2 bar).
  Contains internal 1.8 K make up valve, to eliminate low pressure vapor under lambda plate.
Relief conditions:
  3.45 bar above lambda plate
  5.86 bar below lambda plate
Lambda Plate Relief System

## Slide 38
SC Magnet Test Stands Workshop CERN June 13-14, 2016
1.8 K TEST DEWAR ASSEMBLY
Lambda Plate Spring-Energized Face Seal to He Vessel
“The load required to completely compress the seal is approximately 50-60 lbs per linear inch. That’s about 4200-5100 lbs. …….the difference in height between a completely compressed seal and one in the relaxed state is large enough to handle gaps if the parts are not completely flat or if the seal is not completely compressed.” P. Kovach, designer.
Lambda plate sealing detail

## Slide 39
SC Magnet Test Stands Workshop CERN June 13-14, 2016
1.8 K TEST DEWAR ASSEMBLY
⅛” aluminum Shell
Flexible copper attachments
Vertical tubes (white & red) ⅜” X .028” wall, seamless 304 Stainless Steel.
Header tubes (Blue) ⅝” X .065” wall, seamless 304 stainless steel.
1/16” fillet welds (TYP)
4.5 K Heat Shield

## Slide 40
SC Magnet Test Stands Workshop CERN June 13-14, 2016
QUENCH PROTECTION
Quench Protection Heater Firing Units (HFU)
PRESENT HFU SYSTEM
Number of capacitor banks 4
Number of capacitors per bank 15 units
Capacitor voltage rating 450 V
Capacitance per unit 3100 μF
Total capacitance per bank 46500 μF
Maximum energy density that can be delivered vs capacitance with present system

## Slide 41
SC Magnet Test Stands Workshop CERN June 13-14, 2016
DATA ACQUISITION SYSTEMS
DAQ SYSTEMS STATUS:
128 Channel Fast logger: Assembled and tested.
64 Channel Slow logger: Assembled and tested.
64 Channel Strain Gauge logger: 32 Channels tested, remaining channels to be assembled.
8 Channel Quench Detector: Assembled and tested.
Cable end connector assembly complete on dewar end.
Acquisition software written/debugged/tested.
All isolators / pre-amplifiers assembled, calibrated & tested.
128 Channel Fast Logger
8 Channel Quench Detector
64 Channel Slow Logger
DAQ-isolation amplifier cabinet
STATUS SUMMARY
(FY16) - Spike Detector & Quench Antenna Data Acquisition System to be designed & built.

## Slide 42
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL CRYOGENICS FACILITY
Inline Purifiers
  Replaced Polyflow lines
Valve Boxes
  Replaced Polyflow lines
Quench Tank
  Replaced valve & legacy soldered copper lines with stainless steel welded pipe; rerouted quench recovery line to improve flow
  Replaced legacy soldered copper small diameter return manifold piping with large diameter welded stainless steel
  Purchased & installed new return manifold valves
“before” legacy copper lines & valves
“now”
Note: all this work is BNL funded

## Slide 43
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL CRYOGENICS FACILITY
4160V transformer relocated into place – conduit/wiring installation complete
Sullair skid modified for new location
  Electrical panel removed
  Frame cut/shortened/re-welded
All electrical & plumbing installations are complete
Waiting for turn-on
Cost underrun! Surplus funds approved for use in commissioning helium dirty gas purifier system (Linde contract)
Note: all this work is BNL funded
* backup unit, risk mitigation against Mycom compressor failure
Sullair 500* (500 hp, 94 g/s) Installation

## Slide 44
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL CRYOGENICS FACILITY
(FY17)*
Re-commission Kinney Vacuum Pump (backup to Nash Pump, 2.0 g/s He @ 1.8K)
* Pending final budget authorization

## Slide 45
SC Magnet Test Stands Workshop CERN June 13-14, 2016
BNL CRYOGENICS FACILITY
Nash Vacuum Pump Controls : New lab View based controls tested for remote and local operation
Finished hardware and software debugging of the new Lab View based Inline Helium purifier and its associated valve system.
Helium Recovery compressor: Retrieved vendor programmed ladder logic code for modification and addition of remote control logic. initial testing revealed damaged pressure sensor (as received). Replacement ordered.
New Lab View based controls for Quench recovery valve system being tested.
Monitoring system and controls for various level sensors, temperature sensors, heaters, vacuum sensors etc. are in process of being calibration and testing.
Cryogenic Controls

## Slide 46
SC Magnet Test Stands Workshop CERN June 13-14, 2016
30 kA POWER SUPPLY UPGRADE
24 kA water-cooled cable to dewar installation is complete
(Note also helium vacuum line)

## Slide 47
SC Magnet Test Stands Workshop CERN June 13-14, 2016
1.8 K TEST DEWAR ASSEMBLY
Pressure relief valve with “backflow” poppet (not shown)
Note: exiting magnet leads will be cut to ~ 300 mm past splice box; is this acceptable for production magnets?
Main current lead connections to 24KA vapor cooled leads (not shown)
CLIQ leads (not installed for Mirror)
Lambda Plate Design Details

## Slide 48
SC Magnet Test Stands Workshop CERN June 13-14, 2016
1.8 K TEST DEWAR ASSEMBLY
ASSEMBLY COMPLETE
top hx end
bottom hx end
Test Dewar Status p.12K Heat Exchanger Assembly

## Slide 49
SC Magnet Test Stands Workshop CERN June 13-14, 2016
1.8 K TEST DEWAR ASSEMBLY
Helium Dewar & Outer 4K Heat Shield Status
6/10/16 –
Heat shield welded to helium vessel.
Insulation blanket installed.
Heat shield (at BNL)

## Slide 50
SC Magnet Test Stands Workshop CERN June 13-14, 2016
1.8 K TEST DEWAR ASSEMBLY
Assembly underway
G-10 & stainless steel plates potted with Stycast epoxy
Components being installed
Test Dewar Status p.3 Lambda Plate

## Slide 51
SC Magnet Test Stands Workshop CERN June 13-14, 2016
1.8 K TEST DEWAR ASSEMBLY
top hat detail parts being fabricated:
  24KA vapor cooled leads received
  Top hat received, being inspected
  (3) Instrumentation towers assembled
  Wiring harnesses being assembled
  Most parts promised 5/16
  Delivery issue with Ultem 24KA lead insulators, material sent to 2nd vendor, parts due 5/19

## Slide 52
SC Magnet Test Stands Workshop CERN June 13-14, 2016
MAGNETIC FIELD MEASUREMENT
Inventory of measuring coils