## Slide 1: Test stand @ CERN
Marta Bajko CERN
1st International Workshop on Superconducting Magnet Test Facilities @ CERN 13th-14th of June 2016
https://indico.cern.ch/event/507584/

## Slide 2: Content
Magnet test stand layout @ CERN
Vertical cryostats : on-going upgrades
  HFM
  Cluster D
Horizontal test benches: on-going upgrades
  Cluster A
  Cluster E
Associated facilities upgrades
Summary and conclusions

## Slide 3: Magnet test stands layout @ CERN
CLUSTERS
F E D
CLUSTERS
C B A
STRING and SC LINK
LN2 TEST ZONE
CRYOGENIC ZONE
Superconducting RF CAVITY TEST ZONE
HORIZONTAL BENCHES
THE SUPERCONDUCTING MAGNET TEST STAND AT CERN in SM18
Is located in a building of 7200 m2 over which the test stands for magnets occupies about 50%.
INSTALLATIONS
3 vertical cryostats,
1 feed-box for supercritical He,
1 cryostat for LN2,
10 horizontal benches
OPERATION
I= 0.12 kA-20 kA
T = 1.9 K to 70 K He
withstanding U = 1kV to 3 kV

## Slide 4: Magnet test stands layout @ CERN
INSTALLATIONS
3 vertical cryostats ,
1 feed-box for supercritical He,
1 cryostat for LN2,
10 horizontal benches
OPERATION
I= 0.12 kA-20 kA
T = 1.9 K to 70 K He
withstanding U = 1kV to 3 kV

## Slide 5: Vertical test stand upgrade
The vertical cryostats zone, called Cluster G is about 400 m2 . It is under upgrade with an extra space called Cluster D of 150 m2 . It will accommodate the test of larger diameter magnets @ higher operating current for HL LHC.
Cluster G
Cluster D

## Slide 6: Cluster G upgrade
The circuit is distributing current towards vertical cryostats orto the feed box via a commutator
“LONG” CRYOSTAT :
3.8 m long and
Φ 600 mm @ 1. 9 K
20 kA POWER CONVERTER
with EE and DUMP RESISTOR
“DIODE/LEAD” CRYOSTAT
@ 4.2 K
“SIEGTAL” CRYOSTAT
1.4 m long
Φ 800 mm @ 1.9 K
CRYOGENIC
FEED BOX FOR
He 4.2-80 K
“HFM” CRYOSTAT for MAGNET TESTS WITH
Φ 1500 mm and 2.5 m length
PRE-COOLING WITH He GAS AND QUENCH BUFFER FOR COLD He RECOVERY
Ref. HFM cryostat is financed by EUCARD and the WP19 task of HL-LHC

## Slide 7: Vertical test stand upgrade
Cluster D

## Slide 8: Cluster D test stand
EE and DUMP AT THE 2 m LEVEL
PREPARATION ZONE AT 0 AND -3 m LEVEL
POWER CONVERTERS30 kA AT GROUND LEVEL
TWO ODH ARE IN THE -3 m ZONE ( mostly justified by presence of N2)
NEW OVERHEAD CRANE ALLOWS HADLING 25 t AND UP TO -3 m LEVEL
CONTROL ROOM NEAR THE CLUSTER
COLD He RECOVERY BUFFER - 10 m3 at 5 bar, UNDER THE WORKING AREA IN CONFINED SPACE
ELECRICAL AND CRYO-RACKS NEAR THE CONTROL ROOM
-10 m
-3 m
0 m
2 m
5.5 m
Integration by A.Kosmicki
CRYOSTAT : 4.8 m LONG AND
Φ 800 mm @ 1. 9 K

## Slide 9: CLUSTER D: electrical circuit
EL CIRCUIT PARAMETERS
Max stored energy 10 MJ
Short Sample current 21.5 kA
Magnet inductance 8.27 mH/m
Nominal Ramp Rate 11 – 20 A/s
Max Ramp Rate 200 A/s
Max current 30 kA
Max EE voltage 1 kV
Reaction time : < 10 ms
DESIGN DRIVEN BY: cost reduction and operation optimization implying the use of two existing 15 kA power converters for the new circuit
Ref. H. Thiessen, C. Giloux, G.J. Coehling, P. Viret

## Slide 10: Cluster D : energy extraction
DESIGN DRIVEN BY : Reaction time < 10 ms, which implied the use of IGBTs
THE 30 kA SWITCH IS MADE WITH TWO 15 KA UNITS EACH COMPOSED BY A MODULE OF 7.5 kA
EACH UNIT OF 15 kA IS EQUIPPED WITH A DUMP RESITOR
FOR 5 MJ
DESIGN GOLAS FOR BUS BARS
Equal paths for all currents
Equal impedances
Compact and cold:
(T max 110 °C and 80 °C on junctions)
Ref. G.J. Coehling, A. Dinus, K. Petersen

## Slide 11: Cluster G and D: cryostat and insert
CRYOSTAT PARAMETERS
(Cluster G/D)
Max internal pressure : 5 bara
Magnet temperature : 1.9 K
Useful diameter : 1500 / 800 mm
Useful length in 1.9 K bath : 2.5 / 5.2 m
Max. thermal gradient : 50 K
Magnetic measurements: yes
Estimated lifetime : 20 years
Number of thermal cycles : < 1000
Number of Quenches : < 10000
INSERT DOES NOT HAVE ANY CRYO LINE CONNECTION
DESIGN DRIVEN by optimization in operation
Minimizing time and risk at hydraulic connection
Minimizing the loss of He after quench
USE OF CRYOGENIC BUFFER FOR COLD He
Ref. Cryostat: A. Van De Craen, M. Struik
Cryogenics: V. Beda, A. J. Lee
Insert: M. Strychalski, P. Viret

## Slide 12: Cluster G and D : WCC and Current Leads
DESIGN DRIVEN BY Optimization in operation :
Flexible cables
Current lead operation homogenized with the rest of the test stand and minimizing the flow
CONVENTIONAL GAS COOLED Cu LEADS:
length: 1.7 - 2.1 m
top flange Φ 260 - 280 mm
connection : SC cable
Current Leads
SPECIFIED PARAMETERS
Main circuit: 20 kA / 30 kA
Secondary circuit: 15 kA
Design pressure : 5 bara
Max. Voltage to GRD: 3.5 kV
He vapor flow @ 4.2 K:
20 kA lead: 0.87 g/s
30 kA lead: 1.82 g/s
15 kA lead: 0.88 g/s
Water Cooled Cables
SPECIFIED PARAMETERS
Current: 5 kA DC
Max in. temperature: 28 °C.
Pressure in operation: 16 bar.
Maximum delta T: 10 °C.
Water speed : < 2.5 ms-1
Max. temperature : 50 °C.
Operating Voltage: 1 kV
(if fault to ground)
Ref. WCC: C. Gonzalves Perez
Current Leads: S. Gianelli, P. Moyret
500 mm2
CABLES SECTION

## Slide 13: Cluster DAQ system
DAQ
SPECIFIED PARAMETERS
Input channels: +/- 10 V analog
Nr of HF channels: 200 differential
HF frequency: 200 kHz
HF,MF,LF resolution: 16 bit resolution
HF,MF,LF accuracy: 1 mV
MF frequency: 50 kHz
Nr of LF channels: 144
LF frequency: 1 kHz
Timing: GMT synchronization
Example of HF system
Data Analysis with DIADEM
Ref. DAQ: H. Reymond, M. Charrondiere, M. I. Tapani
Analysis: H. Bajas, F. Gomez de la Cruz

## Slide 14: Cluster D in pictures
Ref.
Construction:
H. Botella, R. Morton, E. Perez Duenas,
Safety: E. Paulat

## Slide 15: Horizontal test bench upgrade
The horizontal test zone Cluster A to F excluding D is composed by 10 benches. They are equipped with 5 independent electrical circuits allowing to test magnets up to 15 kA @ 1.9 K. Two clusters are planned to be upgraded to 20 kA capacity.
Cluster A
Cluster E

## Slide 16: Horizontal test bench upgrade
Cluster A
Magnet weight: 25 tonnes
Operating temperature: 4.2 K and 1. 9 K
Max. Operating current: 15 kA, 2 x 0.6 kA
Max. Pressure in the magnet in case of quench: 20 bars
Max. Test voltage: 3 kV
Magnet weight: 25 tonnes
Operating temperature: 4.2 K and 1. 9 K
Max. Operating current: 20 kA, 2 x 2 kA
Max. Pressure in the magnet in case of quench: 20 bars
Max. Test voltage: 3 kV
Target for prototype test: 90% ISS ≤ 20 kA
Target for series test : 108% Inom = 17.8 kA
The UPGRADE is DRIVEN by HL-LHC needs:
Cluster E
Cluster E will have the same characteristics. Foreseen later.

## Slide 17: Facilities Upgrade
  CRYOGENIC COOLING PRODUCTION: + 35 g/s LHe
  Needs essentially for HL-LHC IT STRING run in parallel with magnet testing
  DEMINERALISED WATER PRODUCTION: + 150 m3/h
  Needs for demineralised water entirely coming from magnet operation
  HANDLING: 25 t and longer rope
  Needs for overhead crane entirely coming from magnet operation
  nCONTROL ROOM
  Needs to extend the small control room of the vertical . Test facility to be used
  also for horizontal benches and Sc link
Remarks
UPGRADE DRIVEN BY The recommendation enabling to carry out the full test programme with no contraints
For 2019
For mid 2016
Operational
Operational

## Slide 18
SM18 2020-2022

## Slide 19: Thank you for your attention

## Slide 20: Spare slides on Facilities upgarde

## Slide 21: Test stands : cryogenics upgrade
An additional 35 g/s liquefier (thus bringing the total production capacity for LHe at 4.2 K to 60 g/s)
The existing total pumping capacity in SM18 at 1.8 K is 12 g/s, combining the capacity of WPU1 and 2.
If need for IR STRING a dedicated 3rd WPU shall be installed.
Foreseen for 2019
Ref. L. Serio

## Slide 22: Test stands : demineralized water production
At present SW18 houses a demineralised water station with a thermal capacity of 800 kW limiting the operation to two test stands
The demands for the magnet test stations, including the IT String, will rise to a peak of 1.7 MW (142 m3/h).
Operational from April 2016
Ref. A. J. Broche

## Slide 23: Test stand : HADLING upgarde
25 T HANDELING CAPACITY WAS KEPT
To enable insertion of test objects into the HFM cryostat the operating height, must be increased to 6 m (presently 5.3 m).
Cluster D, on the other hand, will be deeper than the existing vertical facilities. To handle the inserts the operation depth between the hook and the top of the cryostat must be increased to 9.3 m.
Operational from February 2016
Ref. R. Rinaldesi

## Slide 24: New Control room
Ref. M. Charrondier
The motivation to build a new control room is driven by:
Safety and
Working conditions
CONTROL ROOM AT THE PROXIMITY OF THE CRYOSTATS
CONTROL ROOM WITH LARGER SPACE
Foreesen for Jul 2016