## Slide 1: FREIA Laboratory Facility for Research Instrumentation and Accelerator Development "Gersemi" Vertical Test Stand
Roger Ruber for the FREIA Team
Dept. Physics and Astronomy
Uppsala University
SC Magnet Test Stand Workshop - CERN - 13 June 2016

## Slide 2: Uppsala University
Oldest university in Scandinavia (1477)
Sweden
  9.7 million (pop.), 450'000 km2, 430 GEur (BNP)
Uppsala
  25'000 students, 9'000 staff, 630 MEur annual budget
  7 faculties of theology, law, medicine, pharmacy, arts, social sciences, languages, educational sciences, science and technology
  university library and hospital
Faculty of Science and technology
  10'000 students, 1'800 staff
  historical profiles: Linnaeus, Rudbeck, Celsius, Ångström, Siegbahn, Svedberg
  R&D areas:
    physics, chemistry, biology, earth sciences, engineering, mathematics, IT

## Slide 3: Uppsala Accelerator History
1940's: The(odore) Svedberg proposes to build a cyclotron
Gustaf Werner synchro-cyclotron (1947 - 2016)
  nuclear physics & oncology
CELSIUS ring (1984 - 2005)
  nuclear & particle physics
CTF3/CLIC (since 2005)
FLASH/XFEL (since 2008)
ESS (since 2009)
FREIA laboratory (est. 2011)
Skandion clinic (est. 2015)

## Slide 4: FREIA Laboratory
Facility for Research Instrumentation and Accelerator Development
cryogenics
- liquid helium
- liquid nitrogen
control room
- equipment controls
- data acquisition
radio-frequency (RF) power sources
3 bunkers
with test stands
horizontal cryostat
vertical cryostat
Competent and motivated staff
collaboration of physics (IFA) and engineering (Teknikum).
Funded by KAWS, Government, Uppsala Univ.

## Slide 5: Overview of Activities
Cryo Test Stands
ESS SRF Linac
High Power RF
Controls & Data Acquisition
THz Coherent Light Source
RF = Radio Frequency
SRF = Superconducting RF
Cryo Distribution
ESS Neutrino Super-beam
HiLumi LHC
CLIC / CTF3

## Slide 6: Cryogenics
Helium liquefaction
  150 l/h at 4.5K (LN2 pre-cooling)
  2000 l LHe dewar/buffer, 3+1 outlets
  cryostats connected in closed loop
Gas recovery
  100 m3 gasbag
  3x 25 m3 /h compressor
  10 m3 200 bar storage
2K Pumping
  ~3.2 g/s at 10 mbar
  ~4.3 g/s at 15 mbar
  110(90)W at 2.0(1.8)K
Liquid nitrogen
  20 m3 LN2 tank
Sub-atmospheric pumping station
Helium gas recovery system

## Slide 7: HNOSS Horizontal Cryostat
HNOSS: Horizontal Nugget for Operation of Superconducting Systems
Main Vacuum Vessel
  3240 x ø1200mm inner volume
  “beam” axis at 1600mm
Valve box (on top of main vessel)
  distribute cryogens
  4K and 2K pots, JT-valve, heat exchanger
  5K supercritical helium
LN2 and LHe transfer lines
  interconnection box to distribute cryogens to HNOSS and CM
Cold gas re-heater
Control system

## Slide 8: Gersemi Vertical Cryostat
Under construction
Operation modes
  vacuum
  sub-atmospheric liquid bath
  pressurized liquid bath
Main Vacuum Vessel
  4436 x ø1100mm inner volume
  2869 mm below lambda plate
Valve box
  distribute cryogens
  4K pot, JT-valve, heat exchanger
  5K supercritical helium
LN2 and LHe transfer lines
Cold gas re-heater
Control system

## Slide 9: Cryostat Assembly

## Slide 10: Lambda Insert

## Slide 11: Cryogenic Operation Scheme

## Slide 12: High Power RF Amplifiers
352 MHz, 400 kW, 3.5 ms, 14-28 Hz
Uppsala design
tetrode tube TH595
prototype for ESS SRF spoke linac
industrial manufacturing
  Itelco-Electrosys (Orvieto, IT)
  DB Elettronica (Padua, IT)
352 & 400 MHz, 50 kW, CW
CERN (loan since Feb.2015)
tetrode tube TH571b

## Slide 13: Controls
Controls and interlock systems
EPICS interface, data archiver
connecting different sub-systems
  Linde, Cryo Diffusion, Leybold
different hardware
  Siemens PLC, Nat.Instr. cRIO
Radiation monitoring system
Rotem MediSmart
2 inside, 3 outside bunker
In-house development LLRF
Nat. Instr. PXI and LabVIEW
self-excited loop with digital phase control
extended RF measurements

## Slide 14: Test Stand ID Card
name FREIA Laboratory
1 location Uppsala University (Uppsala, Sweden)
number of people working at the test stand 20 (of which 6 scientists, 4 PhD students)
2 surface of the test stand approx. 1100 m2
3 operating temperature 1.8K to 4.2K (LN2 cooled thermal shield)
cooling techniques pressurized or sub-atmospheric bath, vacuum
4 cooling phases 300 - 4.2 K, 4.2 K - 1.8K
5 cooling and pumping capacity 140 l/h liquid production, 4 g/s pumping
6 shared cryogenics yes, with horizontal cryostat test stand
7 PWC max. current, max voltage ±2 kA, ±10 V (planned, CERN design)
8 HVWL
9 number of cryostats 1 vertical (Gersemi), 1 horizontal (HNOSS)
10 capacity of cryostats (useable, l x d) 2.8 m x 1.1 m [vert.], 3.6 m x 1.1 m [hor.]
11 handling tools 1 x 6.5 t overhead crane
12 interlock safety NI CompactRIO, LabVIEW
13 DAQ cards and used software NI PXI, LabView
14 quality control tools
15 magnetic measurement capability

## Slide 15: This work is supported by Knut and Alice Wallenberg Foundation, Swedish Government, Uppsala University
In collaboration with
Dept. Physics and Astronomy and Dept. Engineering Sciences, Uppsala University
Accelerator and Cryogenic Systems, Orsay, France