The NA61 TPCs (2)
Overview:
principle
Mechanics
readout
Position resolution
Calibration
Particle identification

---
Electronic gain calibration
Important for precise position determination: centroid calculation.
However, due to the fact that the spread of parameters is small within a chip (significantly
larger from chip to chip and batch to batch) and in NA61 neighboring pads are connected
to neighboring channels and there are 16 channels per chip there is a only a small
improvement due to calibration.
Not so important for dE/dx calibration: elementary observable has already large spread
        (due to the Landau distribution of signal)
more important: systematic shifts due to mechanical tolerances leading to gas gain variations

---
Illustration for non working electronic gain calibration
uncalibrated
Extracted gain
factors
“calibrated”
useless
This is due to a deficiency of the PASA chip: with all channels having a signal
Generated by the calibration pulser),
output drops depending on position of channel on the chip

---
Response of preamp pulsing one channel vs pulsing all channels

---
Finally channel-by-channel gain calibration not so important:
gains inside a 16 channel chip are quite homogeneous
mapping of pads to electronics channels is 1:1 --> centroid is still very precise
much more important: overall (large scale) gas gain calibration described
In the next slides

---
Calibration pulser
Even though the electronics gain calibration is not applied there is a calibration pulser
in NA61.
It's purpose is first time calibration and then diagnostics: find dead channels, dead chips
(groups of 16 channels), problems of front-end-cards (32 channels), readout problems....
We even built a new calibration pulser system. The old one was not maintainable any more
and the control did not fit into the present DCS system.

---
Time calibration
Time shifts due to:
- trigger cable length variations (large)
- shaping time variations: variations chip to chip (relatively small)
Groups of 16 channels from one chip

---
Full gain calibration
includes gas gain calibration

---
Instead of calibration pulser used: Krypton calibration
        (the only (gain) calibration in NA61)
advantage: takes also gas gain variations into account
due to mechanical variations across chamber surface

---
Response Simulation
Response Simulation
Slide
From Danilo Vranic
Simulation
keV
Kr-83
Rb-83
86.2 d
571
30%
562
62%
26% 74%
42
1.83 h
147 ns
9.4
 6%
0
keV
Isomeric 41.5 keV level decays entirely to the 9.4 keV level by IC
In addition, Auger electrons and X-rays are produced

---
Setup
Setup
Slide
From Danilo Vranic
From ISOLDE at CERN
gas output
decaying Kr-83 particle
bypass line solid Rb-83
gas input
Kr-83 particles

---
With and without calibration:
Difference between sector types of MTPC

---
Example for gain variation across the chamber surface
(this is from ALICE, I don't have an equivalent plot from NA49/NA61)

---
1 slice = 2 pads
Pad #
This and the following plots from A. Rybicky's thesis:
https://edms.cern.ch/file/816032/1/thesis_int_final.pdf
NA49 Kr calibration results: due to lack of statistics: sum of all pad-rows in a sector

---
NA49 Kr calibration results

---
Pulser vs krypton calibration

---
Old comment:
Important for NA61: repeat the Krypton calibration
all prerequisites in the gas system are existing
we need only 83Rb from ISOLDE, some Mbq (5 Mbq is the limit set by CERN safety)
and: lots of analysis work .....
As of today:
Kr calibration data taken with high statistics.
This allows gain calibration < 1% per pad!

---
Particle identification by dE/dx

---
dE/dx signal amplitude distribution
Landau distribution

---
“Truncated mean” method
After truncated mean method
is applied
Raw data: Landau distribution
Lower cutoff
Measured charge (ADC channels)
Effect of different truncation limits
Upper cutoff
measurements

---
dE/dx resolution for different gases
dN/E = E-α
α smaller for heavier gases but N overall larger for heavier gases --> compensating effect
dE/dx resolution about the same for Ne and Ar

---
From Pb + Pb interactions
For the extraction of particle yields:
Gauβ fits
one needs to know the exact form of the dE/dx
curve (this parameter is kept fixed in minimization)
dE/dx

---
NA49/NA61 specific problems
Good news: electronics calibration not critical for dE/dx resolution:
random (Gaussian) distribution: additional 10% “noise” doesn't matter:
        single measurement has a ~50% width
dE/dx measured in different TPCs with different gas mixtures: relativistic rise different
difficult to combine
Additional problem at high multiplicities, as in Pb-Pb at top SPS energy:
Basline shift
Common mode effect
Fixed!

---
Baseline shift
Relevant for Pb-Pb collisions at the highest energies:
→ high track densities
→ systematic shifts of dE/dx as function of drift length
VTPC
NE/CO2
Simulation: Garfield,
Rob Veenhof
NA49 laser signal

---
Baseline shift
MTPC
(Ar/CO2/CH4)
Simulation: Garfield,
Rob Veenhof
NA49 laser signal
Drift velocity of ions
not correct in
simulation

---
Baseline shift
Cluster ions in gas-based detectors
Y. Kalkan,a,b,1 M. Arslandok,c A.F.V. Cortez,d Y. Kaya,e ˙I. Tapana and R. Veenhof f ,a
2015 JINST 10 P07004
Alice: no field wires
Exact shape of baseline
Movement depends on
Distance of pad to cluster
Centroid!!
From Mesut Arslandok, Alice

---
Baseline shift
Alice: no field wires
Alice laser signal
From Mesut Arslandok, Alice

---
Common mode effect
at the time I called it “lateral cross talk”

---
Common mode effect
at the time I called it “lateral cross talk”

---
Correction of baseline shift and Common mode effect in Pb-Pb
Common mode effect
Drift length
Thesis Christof Roland (now CMS)
Baseline shift

---
Possible solution of baseline shift problem with Alice electronics
This what should happen in the Alice readout:
Pedestal subtraction
Moving average filter
This last step never worked in Alice up to now.
May be now experts found a trick in the firmware.

---
the end