The NA61 TPCs (1)
Overview:
principle
Mechanics
readout
Position resolution
Particle identification
Calibration

---
Structure of a TPC
Working principle of a TPC
Driftgas
High voltage electrode (negative)
Multiwire proportional chamber (0V)
Electric field (elctron drift)
Voltage divider
Field cage
R
R
R
R
R
R
R
R
R
Magnetic Field (measure particle momentum)

---
Signal generation in a multiwire proportional chamber with pad readout

---
Signals of one pad after digitization and zero suppression
(NA49 online monitor)
Drift time of electron cloud
Amplitude (ADC counts)

---
The gating grid
prevents positive ions from drifting back into drift volume: space charge distortions
prevents electrons from entering the amplification region: ageing
Garfield calculation of
field lines

---
The gating grid
Different way of presentation:
drift lines of electrons/ions
Garfield calculation for OROC
chambers of ALICE (no field
wires → larger signals on pads)

---
Gating grid pulser
Principle of operation:
Badly tuned pulser causes noise in the first
couple of time bins
ONE THING TO BE CHECKED BY
SHIFT CREWS!

---
Ageing effects

---
Construction of the VTPCs
This design allows very light construction:
multiple scattering is minimized

---
VTPC:
looking along the beam
direction
field cage is split in two halfs,
the beam area is outside the field
cage

---
MTPC field cage

---
Inside the MTPC

---
chamber diagonal
MTPC chambers are first inserted through the opening,then pulled back
sealing with O-ring

---
VTPC chambers are inserted from top
and sealed with RTV (silicone)
Chamber body
Support frame
(end plate)

---
TPC readout

---
Frontend
Na49 DAQ (schematically)
CT board
70 ms for 512 tb
35 ms for 256 tb
2.2s to fill
242 Optical links
Receiver
32 (64) event buffer
mr
ml
V2
V1
  VSB bus
6x2.5MB/s
Uneven distribution over 6 links:
different noise levels
Eventbuilder
Spill structure
7s
4.5s
FDDI
3.2s
2.4s
PC
Tape
CERN backbone
new:38 s
14.4 (19) s
CDR
16 MB/s

---
Digitized signal of one pad (with zero suppression)

---
Resolution issues: position resolution
Important for momentum (vector) determination: curvature in magnetic field
        emission angle

---
Position resolution (pad direction) as fu of drift length (MTPC)
Diff & gas gain fluct &
pad size &
electr. Noise & threshold
measurement
Diffusion & gas gain fluct
2 x Diffusion
Diffusion only

---
Optimization of pad angles:
track angle relative to beam direction at center of VTPC2
π- from Lambda
P from Lambda
Angle (rad)
All neg tracks
All pos tracks
Final pad angle
RMS
X (pad direction [cm]

---
Effect of track angle
L
projection
δx ~ L*σ *tanα/√Neff
Neff ~ 12/cm (not 40)
centroid

---
Position resolution as function of track angle
Sectors 4 - 6
Sectors 1 - 3

---
Effect of Lorentz angle at wire on cluster width and resolution
asymmetric relative to beam direction, i.e. Positive and negatice particles

---
to be continued