## Slide 1: SM18 update and summary of Superconducting Magnet Test Facility Workshop
M. Bajko CERN
TCC 7th if July 2016

## Slide 2: Content
Test stands @ CERN
Cluster G upgrade
Cluster D upgrade
Service upgrade
Status
SMTS Workshop

## Slide 3: Magnet test stands layout @ CERN
CLUSTERS
F E D
CLUSTERS
C B A
STRING and SC LINK
LN2 TEST ZONE
CRYOGENIC ZONE
Superconducting RF CAVITY TEST ZONE
HORIZONTAL BENCHES
THE SUPERCONDUCTING MAGNET TEST STAND AT CERN in SM18
Is in a building of 7200 m2 over which the test stands for magnets occupies about 50%.
INSTALLATIONS
3 vertical cryostats ,
1 feed box for supercritical He,
1 cryostat for LN2,
10 horizontal benches
OPERATION
I= 0.12 kA-20 kA
T = 1.9 K to 70 K He
withstanding U = 1V to 3 kV

## Slide 4: Vertical test stand upgrade
The vertical cryostats zone , called Cluster G of about 400 m2 is under upgrade with an extra space called Cluster D of 150 m2 to accommodate the test of larger diameter magnets @ higher operating current for HL LHC.
Cluster G
Cluster D

## Slide 5: Cluster G upgrade:the FReSca2 cryostat(HFM)
The circuit is distributing current towards vertical cryostats or the feed box via a commutator
“LONG” CRYOSTAT :
3.8 m long and 600 mm diameter @ 1. 9 K
20 KA POWER CONVERTER
with EE and DUMP RESISTOR
“DIODE/LEAD” CRYOSTAT :
@ 4.2 K
“SIEGTAL” CRYOSTAT :
1.4 m long and 800 mm diameter @ 1. 9 K
CRYOGENIC
FEED BOX FOR
He 4.2-80 K
The new cryostat in this zone is the:
“HFM” CRYOSTAT ALLOWING MAGNETS TO BE TESTED WITH
1500 mm diameter and 2. 5 m length
PRE COOLING WITH He GAS AND QUENCH BUFFER FOR COLD He RECOVERY
Ref. HFM cryostats is financed by EUCARD and the WP19 task of HL LHC

## Slide 6: Cluster G upgrade :pictures

## Slide 7: Vertical test stand upgrade
Cluster D

## Slide 8: Cluster D: working area
EE and DUMP on the 2 m LEVEL
PREPARATION ZONE AT 0 AND -3 m LEVEL
POWER CONVERTERS AT GROUND LEVEL
TWO ODH ARE IN THE -3 m ZONE ( mostly justified by presence of N2)
NEW OVERHEAD CRANE ALLOWS HADLING 25 t AND AT -3 m LEVEL
CONTROL ROOM NEAR THE CLUSTER
COLD He RECOVERY BUFFER - 8 m3 à 5 bar, UNDER THE WORKING AREA IN CONFINED SPACE
ELECRICAL AND CRYO RACKS NEAR THE CONTROL ROOM
-10 m
-3 m
0 m
2 m
5.5 m
Integration by A.Kosmicki

## Slide 9: Services upgrade in the test facility
  CRYOGENIC COOLING PRODUCTION: + 35 g/s LHe
  NEEDS essentially for the running of the HL LHC IR STRING in parallel with magnet testing
  DEMINERALISED WATER PRODUCTION: + 150 m3/h
  NEEDS FOR DEMINERALISED WATER ENTIRELLY COMING FROM MAGNET operation
  HANDLING: 25 T and longer rope
  NEEDS FOR OVERHEAD CRANE CHANGE ENTIRELLY COMING FROM MAGNET OPERATION POWERING FROM THE NETWORK: 2 MVA
  NEEDS FOR NEW OR MODIFIED PC COMING FORM MAGNETS and IR STRING
  nCONTROL ROOM
  NEEDS TO EXTEND THE TOO SMALL CONTROL ROOM OF THE VERTICAL TEST FACILITY TO BE USED ALSDO FOR HORIZONTAL BENCHES AND Sc link
Remarks
UPGRADE DRIVEN BY The recommendation enabling to carry out the full test programme with no contraints

## Slide 10: New Control room
Ref. M. Charrondier
CONTROL ROOM WITH LARGER SPACE than today
Foreesen for Jul 2016

## Slide 11: Cluster D in pictures
Ref.
Construction:
H. Botella, R. Morton, E. Perez Duenas,
Safety: E. Paulat

## Slide 12: Planning
Cluster G
Cluster D
Installation finished
Commissioning in July-August
Operational in September????
Civil engineering finished
Installation started
Commissioning the earliest in September
Operational in December?

## Slide 13
Present plan of commisioning
Magnet delivery

## Slide 14
Alternative plan1 of commissioning
Magnet delivery

## Slide 15
Alternative plan2 of commissioning
Magnet delivery

## Slide 17: HL-LHC as a trigger for the workshop- present
The HL-LHC project with its 100 superconducting magnets to be built within a very large and international collaboration, triggered the idea of this workshop.
See presentation of L. Rossi
Courtesy of E. Todesco HL-LHC WP3
The goal is to HARMONIZE the tests performed in different places and to BUILD A GOOD COMMUNICATION CHANNEL between us: working in the same field at the four corners of the world.

## Slide 18
2016
2017
2018
2019
2020
2025
2030
2035
2040
2015
We are looking into the future
FCC production and test ?!?
The SMTS Workshop would be held typically ONCE A YEAR and at the different test stations. It can be linked to events organized by HL-LHC project or conferences where the participants are usually present.
See presentation of L. Bottura

## Slide 19: Welcome everybody!
Chart: The 65 registered PARTICIPANTS
| USA | J | CH | PL | FR | IT | S | DE
 | 8 | 3 | 37 | 3 | 4 | 2 | 4 | 2
Most of these participants has a valuable long experience and expertise
in superconducting magnet testing