## Slide 1
Institute of Nuclear and Particle Physics
Tests of a MM octant prototype towards
a Micromegas TPC Polarimeter
for srEDM
George Fanourakis1, Spyros Tzamarias2
1N.C.S.R. “Demokritos”
2Aristotle University of Thessaloniki

## Slide 2
Designing the srEDM MM polarimeter
Units in mm - Not to scale
For 5o-20o scattering angle

## Slide 3
A prototype octant design
The r and phi strips are read by 5 Panasonic connectors (128 pins, not all the pins are used). A total of 442 channels.
Initially configured as normal tracking chambers (not as TPC) for testing purposes with a
5mm drift/conversion space)
The readout is protected for sparks by
a resistive layer

## Slide 4
110 inner r-strips with 1 mm pitch (50mm-160mm)
67 next r-strips with 1.5 mm pitch(160mm-260mm)
40 next r-strips with 2 mm pitch (260mm-340mm)
44 outer r-strips with 2.5 mm pitch (340mm-450mm)
261 R-strips,
4 sections
r-strips

## Slide 5
23 inner phi-strips with 1.96o pitch (0.2mm-2mm,)
+23 next phi-strips with 0.98o pitch
+45 next phi-strips with 0.49o pitch
+90 outer phi-strips with 0.25o pitch
181 phi-strips,
4 sections
phi-strips

## Slide 6
Testing the prototype octant
Two r-phi prototype octants have been ordered and constructed in the electronics lab of CERN (Rui)
Connectors
Drift layer HV connection
Resistive layer HV connection
Mesh HV connection
G10-Rohacell casing
Read out strips layer
Drift electrode layer, below Rohacell plane
Gas inlet
Gas inlet

## Slide 7
ΦA
ΦB
R4
R1,2
R3
The one r-phi octant was read via the SRS readout system using the APV chips, developed by the RD51 collaboration. This system reads both the charge and the time of the pulse from each strip. The initial tests were done at RD51 lab with Fe55 and Sr90 sources. We used the mesh signal to trigger and the MMDAQ readout software developed by ATLAS. The gas used was Ar/CO2 (93%/7%)
At RD51 lab (with Eraldo)

## Slide 8
ΦΒ
R1,2
r4
r3
ΦA
ΦA
ΦΒ
R1,2
r4
r3
Left phi
Right phi
1-2 th Outer R
Inner R
3rd Outer R
At RD51 lab (with Eraldo)
MMDAQ screen dump – with Sr90 source

## Slide 9
The one r-phi octant was read via the SRS readout system using the APV chips. We used scintillator counters (setup by Amos group) to trigger and the MMDAQ readout software on Eraldo’s laptop, controlled via TeamViewer. We took lots of data (HV scans at various spots)
At RD51 Test beam (H4)

## Slide 11
Preliminary analysis results
Beam trigger area
Expect hits at R4, some R3 and A-Phi, B-Phi strips
Established the mapping of strip-to-APV-channel (almost).
No clustering as yet.
Efficiency ~99% (after using gravity to obtain uniform gas flow)

## Slide 18
Plans for September tests
Spyros and I plan to arrive on Monday the 3rd of October, prepare for installation and install on Wednesday morning.
Use two R-Phi octants and measure efficiency relative to each other
The second octant has been modified (by Rui) to guarantee uniform gas flow
Take data with a beam telescope (if possible)
(If ready) try another type of large area (50x50 cm) MM detector with 128 pads (1.5x12cm)
Request beam time for the 2nd week of beam tests
Many thanks to Eraldo and Yorgos for the support and to the Weizmann group colleagues for the cooperation and help.