## Slide 1: The New Beam Loss Monitoring System for CERN's Injector Complex
Topical Workshop on Beam Loss Monitors (15-16/09/2016)

## Slide 2: Introduction
This project has undertaken the task to develop up-to-date Beam Loss Monitoring System for the Injectors.
Mainly,
Build a generic, highly configurable and high-performing system
Acquisition part to accept several detector types
Use reprogrammable parts to target all injectors’ requirements
Team members (past and present):
BI-BL: M. Alsdorf, C. Arcola, B. Dehning, C. Hajdu, M. Kwiatkowski, O. Malnasi-Csizmadia, W. Vigano, C. Zamantzas
and support from: E. Effinger, J. Emery, G. Venturini, E. Nebot Del Busto
BI-SW: E. Angelogiannopoulos, S. Jackson, L. Jensen, B. Kolad

## Slide 3: Outline
Specifications
  Dynamic range, Measurements, Supervision, Interlocking
System overview
  Architecture, Cabling
Electronics
Examples of the installations in LINAC4, PSB and PS
Summary

## Slide 4: System Architecture

## Slide 5: Specifications

## Slide 6: Acquisition & Processing
Is timing/synchronisation at the acquisition and/or processing level necessary?
Synchronisation is required with the start of the cycle to
  Perform calculation of integration periods and
  Schedule comparisons with their corresponding threshold values
  Record high frequency observation data
  Schedule the data readout and publish by the CPU
Synchronisation is achieved by
  Use the ‘Start of Cycle’, ‘Beam In’ and ‘Beam Out’ events received through the timing system.
  Dedicated timing card with broadcast in the backplane.
  Sync is done at the processing level (i.e. 2 samples jitter between cards).

## Slide 7: Integration Periods
Continuously the processing electronics calculate 4 integration period values for each channel:
2 s, 400 s, 1 ms and 1.2 s (machine basic period)
  Implemented as moving sum windows in the hardware
  Calculation is refreshed at acquisition frequency
Compare with predefined thresholds
  Machine protection with hardware implementation comparisons on each refresh
  Limit radiation levels with software implementation comparisons at end of cycle
  See also next slide.
Calculate for each channel the maximum values recorded on each integration period during the cycle
  Published for the online displays and
  Send to the long-term logging

## Slide 8: Threshold Comparisons
Hardware implementation part:
All calculated integration period values, i.e from 2 s to 1.2 s, will be constantly checked against their threshold values:
  4 threshold values, one for each of the integration periods.
  Comparisons happen at the refresh period – that is, every 2 s
  In the case the measured values exceed those the beam permit signal will be removed for all users
  The blocked beam permit signal will be latched until an operator acknowledges.
The threshold values will be need to be set unique per channel:
  Each card will process 8 channels
Software implementation part:
All maximum integration period values recorded on the cycle will be checked against a second set of threshold values. The outputs will be used for repeated over threshold function
  Additional threshold values for the same integration periods will also be required.
  In the case found to be over threshold repeatedly n times it will be required to block that user’s injections.
  The blocked beam permit signal will be latched until an operator acknowledges.
  The repeat value n will be settable per monitor in the range of 1 to 16.
The threshold values will need to be unique per user and per channel:
  Each CPU will process 8 cards x 8 channels
  The information of the current user has to be obtained from the telegram per cycle -> dedicated timing card
  Memory for 32 users will be reserved.

## Slide 9: Beam Permit Logic
System [HW and/or SW] will block injections
  i.e. “remove permit” if losses over threshold
System [SW] will remember if the user is allowed to have beam
  i.e. “give permit” if previous cycle for the user was ok (or previous interlocks were cleared)
The Beam Interlock Controller will be configured in the “Non-latch” mode.
  i.e. the system will need to follow timing and notify in advance.
Aiming to keep system latency (from measurement to output) small
  HW: The target for the fast integration periods is < 10 μs
  SW: Block on next cycle
Only data from the current cycle need to be considered.
  Timing in the electronics essential (i.e. possible failure mode)

## Slide 10: Ambient Radiation Measurement
Calculate and log the ambient radiation measured at the end of each cycle
Processing electronics will provide two values:
  total accumulated in the cycle (already described) and
  total accumulated with beam present
Subtraction of the two values in CPU
Additional timing events to be used for the recording
Values will come together with number of samples used in the recording to allow accurate conversion to user-friendly units, i.e. Gy, Gy/s, …
Publish values for the online displays and the long-term logging

## Slide 11: Evolution Over Time buffer
Record detailed observation data
  Publish on the online displays
  Log on demand
LINAC4:
  2 s samples for 1 ms during beam presence
PSB:
  1 ms samples for 600 ms during beam presence
  2 s samples for 1 ms with adjustable start (aka Capture)
Readout and transmission window by CPU during beam-out period

## Slide 12: Evolution & Capture functions
274
805
1200
0
Time [ms]
SoC
SoC
Cycle n
275
LINAC4
2 μs integral x 500 samples
274
Cycle n+1
transmission window
recording window
PSB
1 ms integral x 530 samples
2μs integral x 500 samples [adjustable 1 ms window]

## Slide 13: FDFC: Acquisition Principle
A status signal selects in which branch of a fully deferential stage the input current is integrated.
Two comparators check the deferential output voltage against a threshold, whenever is exceeded, the status signal changes to the complementary value (0 ! 1 or 1 ! 0) and the input current is integrated in the other branch.

## Slide 14: FDFC: Data Pre-Processing
0
Time
counts
2 μs
2 μs
2 μs
ADC values
thhigh
thLow
The number of accumulated counts are combined with the ΔADC values to calculate the integrated loss over a 2 μs period.
Most of the operations are handled by the FPGA. That is,
  Defines start and stop of the acquisition period
  Keeps a count of the number of pulses occurred in the acquisition period,
  Clocks the ADC circuitries and makes differences of the recorded ADC values
  Finally, processes the data and provides the 2 μs integral per channel

## Slide 15: FDFC: Pre-Processing Algorithm
Case | Calculation
1.1 | An+1-An
1.2 | An-An+1
2.1 | (max-An) + (max-An+1)
2.2 | (An-min) + (An+1-min)
3.1 | (max-An) + (max-An+1) + (M-1) * (max-min)
3.2 | (An-min) + (An+1-min) + (M-1) * (max-min)
4.1 | (An-min) + (max-An+1) + (M-1) * (max-min)
4.2 | (max-An) + (An+1-min) + (M-1) * (max-min)
In order to evaluate the integral between two samples the algorithm needs to cover eight cases.
By tracking the slope and using the absolute value of the differences the realisation in the FPGA is significantly optimised.
The above is for an ideal circuit. In reality an additional stage is introduced to compensate for component tolerances, switching delays, and other effects.

## Slide 16: Acquisition channel Block Diagram

## Slide 17: Acquisition principle (FDFC & DADC)
The input channel circuit is able to measure current input from 10pA to 200mA.
The measurement of the input current is performed by two techniques:
Dual Polarity Current to Frequency Converter (FDFC) operates in the range 10pA to 30mA
Direct ADC acquisition (DADC) usable in the range 20.3µA to 200mA
No gain change required:
The switch between the 2 ranges is managed by the FPGA.
If the maximum FDFC counts is reached, the FPGA switches the circuit to the DADC mode.
When the value of the DADC falls below a threshold, the FPGA switches the circuit to the FDFC mode.
The sum of all parts is calculated in the FPGA and transmitted as a 2 us sample.
DPFC
DPFC

## Slide 18: System overview

## Slide 19: System Architecture

## Slide 20: Cabling Overview
Ionization Chamber
Single HV cable to each detector using a fanout module.
Screen of detector power cable is open on the IC side to assure there is no ground loop.
Internal screen to shield low frequency noise (GND only on electronics side, IC is floating).
External screen to shield high frequency noise (and atm not grounded at tunel side).
enclosed in cable tray
tube
~ 200 m
~ 5 m
~ 1 m
BNC
SMA
TRIAX
Connector Types:
CKC50
External Screen 2
Internal Screen 1
Signal
Acquisition
Card
(BLEDP)
HV Power
Supply
(~1500 V)
CBH50
Fanout 1-to-64

## Slide 21: Cabling Overview
HV cable: CBH50
Signal cable: custom-made coaxial double shielded CKC50
Connectors/Plugs: custom-made triaxial

## Slide 22: Cable Management
Signal Cable panel &
HV distribution
HV Cables
Signal Cables

## Slide 23: Electronics

## Slide 24: Acquisition Crate
Acquisition module (BLEDP)
Up to 8 modules with 8 channel each
Main panel
Ref. current Input
LEDs
Power switch
Control Unit
Later version w/ advanced remote functions
Custom Backplane
Support 64 connectors and relays for the input channels and distribute signals
Production completed

## Slide 25: Acquisition module (BLEDP)
Backplane connection
Analogue inputs, power and control
SFP connectors
Gigabit optical and/or Ethernet links
FPGA
Altera Cyclone IV
JTAG connection
Local programming and diagnostics
Acquisition
digitisation of 8 channels
Currently in production

## Slide 26: BLEDP Optimisation
Improve the insulation between input lines and other signals e.g. power supplies.
This has been done by:
  creating special VIAS,
  optimising routing paths and
  creating holes below the integration capacitor.
  several ground areas have been created.
Special insulated VIAS
Holes below integrator capacitors
Areas dedicated to the monitors
Standard Vias and Through hole footprint
Modified Vias and Through hole footprint
Hole
High Insulation Capacitor

## Slide 27: Processing Electronics
Processing Mezzanine
  New design to match the acquisition modules
  2 SFP modules (gigabit Optical links and Ethernet)
  Cyclone IV (150K logic elements) FPGA
Carrier VME64x modules
  Generic BI card, i.e. DAB64x
VME crates with BI custom backplane
  Broadcast of timing events to the modules
  Daisy-chain of beam permit signals
  Same as LHC system

## Slide 28: Processing & Threshold Comparator
2 SFP modules
Gigabit Optical links and Ethernet
FPGA
Cyclone IV 150K logic elements
Connectors
to the mainboard
JTAG connection
Local programming and diagnostics

## Slide 29: VFC-HD module
Beam Synchronous Timing
Ethernet
4 SFP for fast connections to the Front-ends
HPC-FMC slot
Rear Transition Module connection
DDR3
ARRIA V FPGA
VME64x interface
P0 User defined connection
CERN Power Supplies
Next generation of generic CERN BI processing modules in production

## Slide 30: System for Verification
Added a gigabit Ethernet module and connect to the network
Developed TCP/IP and UDP stacks in the FPGA firmware -> NIOSII softcore
Developed a server to be included in the FPGA -> compiled C++
  accept commands, prepare data requested and transmit
Developed client applications in Java, Python and C
  send commands, store and display data
Need to record all raw data
Analyse data offline
Debug & optimise pre-processing algorithm
Characterise noise & analogue circuit.

## Slide 31: Stand-alone/Ethernet version
Ethernet-based version of the system
Very powerful for
  verification,
  commissioning and
  fine observations
500,000 acquisition points
Cycle selection
Front view
Rear view

## Slide 32: Installation overview

## Slide 33: Pictures of the installation
Rack includes:
Acquisition crate
Processing (VME) crate
HV Power supply for the detectors
User Interface to BIS (i.e. CIBU)
Separated enclosed cable tray
Metallic cable tube for the last few meters
Enclosed and extended front for cables and fibres
Fans on the rack top
Standard and Uninterrupted Power Supply connections

## Slide 34: Rack Management

## Slide 35: LINAC4: Info
11 ICs in the machine line and
14 ICs in the transfer line

## Slide 36: LINAC4: Supports
Machine line
Dump and Transfer line

## Slide 37: PSB: Info
Machine/Area |  | Channels | Cables | Detector Type
PSB | Ring (L2 sections) | 32 | 64 | LHC-IC
 | Ring (L3 sections) | 32 | 64 | FIC
 | H- and BI lines | 18 | 36 | LHC-IC
 | H- | 8 | 40 | Diamonds
 | Extraction lines | 28 | 56 | LHC-IC
Detectors:
  8 Diamonds with amplification and splitter stages
  32 Flat Ionisation Champers (FIC)
  72 LHC-type Ionisation Champers
Cabling
  Ring L2 sections complete
  Remaining installation is planned for 2017
Supports:
  Common design for the H- detectors (ICs and Diamonds)
  Transfer lines have not been considered yet

## Slide 38: PSB Injection: IC & Diamond Support
The new H− charge-exchange injection system

## Slide 39: PSB Ring: IC Detector Support
Very tight integration.
Had to verify all locations; no interference / obstruction to other systems.

## Slide 40: PSB Ring/Ejection: IC Custom support
No space to add the detector on the inside of the ring
ACEMs mounted on the same plate

## Slide 41: PSB Ring: FIC Detectors
Rectangular detectors (FIC) placed between beam pipes
Support mounts on the magnet
Isolation tests performed

## Slide 42: PS Info
Detectors:
  17 Diamonds with amplification and splitter stages
  151 ICs are in production (delivery expected end 2016)
Cabling installation is planned for 2017
Supports:
  Common design for the ring detectors (ICs and Diamonds)
  Transfer lines have not been considered yet
Machine/Area |  | Channels | Cables | Detector Type
PS | Ring | 100 | 200 | LHC-IC
 | Ring (observation) | 17 | 85 | Diamond
 | Transfer Lines | 51 | 102 | LHC-IC

## Slide 43: PS Ring Support
Common support for IC and Diamond

## Slide 44: Summary

## Slide 45: Summarising
Reliable tunnel installation
  Monitors will be mostly Ionisation Chamber type
  High immunity to EMI with special cables, connectors and enclosed cabletrays
  Special attention given to grounding and isolation
High performance acquisition electronics
  Automatic range selection
  Dynamic range = 1011
  In-system calibration mode (will have remote execution in the future)
Modular processing electronics
  High availability of resources (~350’000 Logic Elements)
  Embedded bidirectional gigabit links
  Additional gigabit Ethernet link (for dedicated measurements)
  Future plan for upgrade using the new BI processing module
Production and installation is expected to be completed mid. 2017
Test-benches for verification, software development and measurements
Beam measurements in October

## Slide 46: Thank you

## Slide 47: Reserved slides

## Slide 48: Status of Development
System Overview

## Slide 49: Status of Development
Acquisition & Control
on-going
Gigabit Ethernet Link
complete
Stand-alone Application
complete
FPGA Development
Data Pre-Processing
complete

## Slide 50: Status of Development
Data Storage (SRAM)
complete
VME Comm.
complete
FESA Class API
complete
Gigabit Optical Link
complete
FPGA Development

## Slide 51: Status of Development
FPGA Development
Timing
Complete
Interlocks
On-going
CIBU Connection
Complete
Data Processing
Complete

## Slide 52: Status of Development
FPGA Development
Thresholds Management
Not Started
FESA Server
complete
Logging Database
complete
System Supervision
On-going

## Slide 55: Diagnostic Data Example