## Slide 1: CERN developments of an optical readout for MPGDs
Filippo Resnati (CERN) on behalf of the GDD team
MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 2: Where to get the light from
Primary
Excitation and ionisation produced by the particle interacting with the gas.
Some dependance
with the electric field.
Electroluminescence
Scintillation without charge amplification.
Increase linearly with the field.
Easy in pure noble gases.
Avalanche
Exponential scaling with the field.
Proportional to the charge gain.
The extreme is the visible spark.
Almost the same spectra
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 3: CF4
Admixtures which provide useful scintillation
(enough and near the visible range) are rare.
CF4 is not transparent to the CF4 de-excitation to ground.
What scintillates must be something else: (CF3*), CF3+* and CF4+*
A. Morozov et al., NIM A 628 (2011) 360
Pure CF4 primary scintillation
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 4: Ar/CF4 80/20
About 0.4 ph/e- in the visible
Plexiglass cuts below 400 nm
Nicely matching the CCD efficiency curve
Preliminary
CF4
ArI 4p 4s lines
High charge gain -> high light gain
Light yield comparison
NaI: ~40 ph/keV GEM at a gain of 105: ~106 ph/keV
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 5: Optical readout
Record the light emitted during the Townsend avalanche with a camera: use the detector as a scintillating plate.
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 6: Only techniques are new
G. Charpak et al., NIM A258 (1987) 177
Parallel mesh chamber filled with Ar/CH4/TEA 80%/8%/2%
seen by an image intensifier and a camera
Muons and delta rays
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 7: Advantages
Simplicity: like taking a picture
Robustness: as a device off-the-shelf
Versatility: several uses and environments
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 8: The setup
10 cm
M = sensor size / image size
M ~ 0.1, Ω ~ 5x10-4
This implies:
- large sensor - low noise
- fast lens
- a lot of light
Geometrical acceptance
10 mm
2 mm
2 mm
2 mm
Ar/CF4 80/20
Plexiglass
+3 dioptre lens
25 mm f/0.95 lens
CCD camera
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 9: Camera and lens
QImaging Retiga R6
CCD: 2688x2200 4.54x4.54 um2 pixels
ADC: 14 bit
rate: 6.9 fps (20fps with binning)
read noise: 5.7 e- RMS
dark current: 0.0002 e-/p/s @ -20o C
trigger: external bulb + others
Navitar
focal length: 25 mm
aperture: f/0.95
Mount: C-Mount
Sensor type: 1’’ format
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 10
4.4 cm
55Fe X-ray imaging
One of the first images
20 min exposure
70 um hole, 140 um pitch

## Slide 11: X-ray images
Raw data: fast (<1 s) acquisition and no processing time
Charge acquisition (26/10/1998)
~7 cm
Light acquisition
X-ray tube with W target at 20 kV - 40 kV at few mA
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 12: X-ray images
X-ray tube with W target
11 kV
40 kV
8 cm
Straight out from the camera, not even the flat field correction
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 13: X-ray images
Increasing energy:
- X-ray more penetrating
- Worse position resolution (larger charge cloud)
X-ray interacting with the CCD
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 14: Fluoroscopy
50 ms exposure
10 Hz acquisition
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 15: CT and 3D imaging
Image -> Sinograms -> Filtered Back Projection -> 3D image
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 16: Fluorescence
X-ray tube
W target
40 kV, 1 mA
Contemporary industrial painting
Pin-hole
GEMs
Shielding
Camera and lens
17 cm
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 17: What should be expected
W target:
bremsstrahlung +
characteristic lines
Lα Lβ Lγ
Limiting factors (because not optimised):
- low energy -> window and cathode
- high energy -> gas
feasible:
Cr, Fe, Ni, Cu
difficult:
Au, Pb
more difficult:
Al, Ag
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 18: The painting (visible)
Fe
SS
Au
Al
?
Cu
Cu
Ag
17 cm
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 19: The painting (X-rays)
O(1 kHz) in 500 s exposure
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 20: Event by event
α from 220Rn & 216Po
X-rays from 55Fe
muon and δ-ray
shower from a charged pion
For hours in sealed mode at test beam
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 21: Event by event: X-rays
X-rays from 55Fe source
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 22: Analysis of single clusters
Spatial resolution in the limit of long exposure:
280 um for 1 cm drift and 5.9 keV X-rays.
Dominated by e- diffusion and cluster size.
Average profile of 5.9 keV events
Spatial resolution improves using the barycentre of the cluster
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 23: The spectrum
from CCD images
FWHM at 5.9 keV over 8x8 cm2
~36% before gain uniformity correction
24.7% after gain uniformity correction
Full energy
Pileup
Escape
55Fe source
Gain uniformity, vignetting plays a role too
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 24: Fluorescence event by event
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 25: Fluorescence event by event
some color coding
Red
Green
Blue
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 26: Fluorescence event by event
some color coding
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 27: Fluorescence event by event
It’s Zinc
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 28: Possible other application
X-ray crystallography
UV imaging
Neutron imaging
Gamma imaging
…
Filippo Resnati - MPGD Applications Beyond Fundamental Science - 15 September 2016 - Aveiro

## Slide 29
Freeze-frame of an X-ray movie of a flying drone
Radiography of a bat and closeup of the GEM holes
Radiography of a crushed cup with pens and its 3D tomographic reconstruction
Visible picture of a painting and its X-ray fluorescence image. Different colours refer to different materials (energy resolved)
Cu
Au
Ag
Zn
SS
Fe
Cu
1€
Al
Single X-rays from 55Fe and the energy spectrum extracted from the images
Alphas from 220Rn decay and its daughter 216Po
O (100 keV) electrons ~MIPs