## Slide 1: GBT-SCA Updates
Alessandro Caratelli, Rui Francisco, Kostas Kloukinas, Alessandro Marchioro,
(alphabetic order)
LHCb Upgrade Electronics meeting
Thursday, 09 June 2016

## Slide 2: GBT-SCA
I2C master
X 16
I2C master
X 16
I2C master
X 16
Network controller
Arbiter
SEU counter
Clock ctrl
Serial interface
sync
TX
SERDES
RX
Config
CLK
RX
TX
SERDES
Config
CLK
AUX_E-PORT
Aux serial port
E-PORT
TX FIFO
RX FIFO
TX FIFO
RX FIFO
General purpose I/O
Interrupt inputs
SPI master
JTAG master
CLK
RX
TX
SLVS
Cm: +0.2V
Dif: ±0.2V
SDA
SCL
CLK
RX
TX
SLVS
Cm: +0.2V
Dif: ±0.2V
SDA x16
SCL x16
I2C master x16
SCLK
MISO
MOSI
SS x8
TCK
TDI
TDO
TMS
RES
GPIO
x32
Temperature sens.
E-FUSEs bank
ADC
Interface
Calibration
Analog IN
x31
Analog OUT x4
DAC
DAC
DAC
DAC x4
Interface

## Slide 3: Changes of the GBT-SCA V2
A new version of the SCA was prototyped in October 2015 with minor changes
Improved ADC performances
  The ADC of the SCA was modified to improve the TID tolerance
  Better linearity, better noise performances,
Allowed higher voltages supply for the periphery VDD
Minor changes on the I2C masters
  The SDA line can now be tristate

## Slide 4: E-port communication
Status:
Start-up and communication ← Tested
Connect and reset commands ← Tested
Switching between primary and auxiliary port ← Tested
Core Logic
Status:
Digital functionalities ← Tested
Receive/Reply and communication related ← Tested
Enable/Disable (clk gating) of all ← Tested communication interfaces

## Slide 5: I2C ports
Status:
Read operations in 7bit addressing mode ← Tested
Read operations in 10bit addressing mode ← Tested
Write operations in 7bit addressing mode ← Tested
Write operations in 10bit addressing mode ← Tested
SPI port
Status:
Read and write operations with different ← Tested settings and frequencies
Read and write operations on ← Tested commercial hardware (memory)

## Slide 6: JTAG port
Status:
Operations with different settings and frequencies ← Tested
GPIO port
Parallel port as output ← Tested
Parallel port as input/tristate ← Tested

## Slide 7: ADC
Conversion operation and reply on all the 31 inputs ← Tested
Crosstalk between the 31 Inputs ← Tested
Gain Calibration logic ← Tested
Internal temperature sensor ← Tested
100 μA Current sources on the ADC pads ← Tested
Offset ← Tested

## Slide 8: ADC Differential Nonlinearity
Figure 3: ADC DNL extrapolated with Instagram method with linear ramp in input, 1600 average samples per bin, source voltage from 16 bit DAC.

## Slide 9: ADC Integral Nonlinearity
Figure 4: ADC INL extrapolated with Instagram method with linear ramp in input, 1600 average samples per bin, source voltage from 16 bit DAC.

## Slide 10: ADC noise performances
Figure 5: Conversion result distribution with fixed input voltage at 0.1V, 0.5V and 0.9V. One been represents 1 count of the ADC
= 0.98
= 0.92
= 0.68

## Slide 11: ADC Offset
Figure 6: ADC Samples close to 0.0 V. From the fitting of the samples (uncalibrated data without offset and gain correction) it is possible to extrapolate the offset due to the comparator. This offset is correctly evaluated and subtracted by the ADC internal state machine, giving the expected correct result.

## Slide 12: ADC TID tests 1/2
1024
2048
3072
4096
5210
0 1 2 3 4 5 6
Mrad
Conversion Result
Figure 7: ADC conversion result with input voltage from 0.0V to 1.0V with steps of 0.1V. A set of measurements is taken every 200 Krad. Dose rate of 800 Krad/h.

## Slide 13: ADC TID tests 2/2
1024
2048
3072
4096
5210
0 10 20 30 40 50 60
Mrad
Conversion Result
Figure 7: ADC conversion result with input voltage from 0.0V to 1.0V with steps of 0.1V. Dose rate of 800 Krad/h up to 6 Mrad and 8 Mrad/h up to 70 Mrad

## Slide 14: ADC Calibration Procedure
Figure 9: calibration procedure implemented on the chip. The blue line represents the linear interpolation of several samples on the ADC before to apply the calibration procedure, while the green line represent the slope after calibration.

## Slide 15: Production plans
Engineering run Scheduled for the 1st July 2016
Volume of 20 000 samples
Unpackaged cheeps will come back in October 2016
Packaged chips will be ready for early 2017