## Slide 1: An ultra-light drift chamber with particle identification capabilities
F. Grancagnolo
INFN – Lecce, ITALY
WG11 Detector Design Meeting
CERN
17 October 2016

## Slide 2: Road to proposal
Ancestor chamber: KLOE at INFN LNF Daϕne φ factory (commissioned in 1998 and currently operating)
CluCou Chamber proposed for the 4th-Concept at ILC (2009)
I-tracker chamber proposed for the Mu2e experiment at Fermilab (2012)
DCH for the MEG2 upgrade at PSI (under construction at INFN and to be commissioned during spring 2017)

## Slide 3: "Traditional" Drift Chamber
A cylindrically symmetric gas volume with (para-)axial wires defining a strong electric field, strung under mechanical tension for electrostatic stability and fixed at their extremities to the end walls by means of feed-through.
CONSTRAINTS:
The end walls, holding the feed-through (which limit the chamber granularity), the FE electronics and the relative cabling, must be rigid enough to transfer the load due to the wire tension (of the order of several Tons) to the outer cylindrical wall, without deforming.
The inner cylindrical wall, usually, does not bear any load, to minimize the multiple scattering of incoming particles.
The gas tightness relies on the hermetic properties of all surfaces and of all their relative joints.

## Slide 4: The KLOE Drift Chamber

## Slide 5: The KLOE Drift Chamber

## Slide 6: Drift Chamber "Innovations"
Separating gas containment from wire support functions
Using a larger number of thinner (and lighter wires)
No feed-through wiring
Using cluster timing for improved spatial resolution
Using cluster counting for particle identification

## Slide 7
Gas containment and Wire support
Gas envelope
Wire Cage
F.E. electronics
Gas containment
Wire support
Gas envelope can freely deform without affecting
the internal wire position and tension.
Wire cage structure not subject to differential
pressure can be light and feed-through-less.

## Slide 8
Gas Envelope
A structural multivariate analysis software (ModeFrontier®) has enabled to find the optimal shape for the profile of the end plates by minimizing maximum stress and stress on inner cylinder
  ANSYS ACP® has chosen the proper unidirectional prepreg to form ply,
  draping of the laminates and flat-wrap of the optimized model
  Solve buckling problems
  of inner cylinder by
  increasing the moment of
  inertia with use of proper
  light core composite
  sandwich
Inner cylinder:
2 C-fiber skins, 2-ply,
C-foam core, 5 mm
0.036 g/cm2
9×10-4 X0
End plates:
4-ply 38μm/ply
orthotropic (0/90/90/0)
0.021 g/cm2
5×10-4 X0
Example: The Mu2e I-Tracker proposal

## Slide 9
Wire cage (conceptual)
struts
strut
“L” shaped
end ring
spider web:
spokes and stays
stiffening
polygon
tie-rod
screws
Example: The Mu2e I-Tracker proposal

## Slide 10: Verification of validity of principle
scale 1:1
model
to verify
validity
of principle
ANSYS
FEM
analysis
Force on the spoke
245.25 N

## Slide 11
This scheme does not require wire feed-through thus allowing for denser wire spacing, i.e. smaller cells (finer chamber granularity) and for larger field to sense wires ratios.
Larger field to sense wires ratio and, therefore, thinner field wires, help reducing multiple scattering contribution and total wire tension on support structure.
Large number of wires and small cells, however, require complex and cumbersome assembly procedures, which call for a novel approach to the wiring problem.
Wire Cage

## Slide 12: DC stringing: the old way
The Old Way
The Three Μοῖραι (Fates)
Atropoς
Laχesiς
Klωϑϖ
Bernardo Strozzi – Le tre Parche – Venezia, circa 1620
The KLOE Drift Chamber
45 m3 > 52,000 wires He/iC4H10

## Slide 13: DC stringing: the novel way
LACHESIS
Winding the thread
cylinder, wire PCB
KLOTHO
Spinning the thread
coil, clutch, wire spool
LABIRINTH
The Extraction System
ATROPOS
Cutting the thread
laser solder system
THESEUS
The wire handling system

## Slide 14: MEG2 DC EndPlates
peek spacer
wire PC board
end-plates
structural
removable
shaft
wire tension
compensating
wheels
sense wires
PC boards
(16 cells)
NC machined
peek spacer
(½ cell thick
1 sector wide)
fiducial
reference
edge
spokes
spoke
3D
measuring table
400 μm
PCB

## Slide 15: MEG2 DC Wiring
Aligning the wire
PC board on the
spinning drum
Winding the
wire on the
spinning drum
Wire tension set up with a real time strain gauge
feedback system on the spooling el. mag. clutch
Pad image
recognition
and contactless
IR laser
soldering
solder feeder
dual safe
soldering
0.1 g
0.1 g
16 sense wires
16 field wires
acceleration phase
first 2 wires
(discarded)
deceleration phase
last 2 wires
(discarded)
constant winding speed
32 wires
wire
strain gauge
fixed pulleys
cathode
anode

## Slide 16: MEG2 DC Wiring
Lifting the wire
PC boards with
vacuum suction cups
Unwinding the
loose multi-wire layer
over the storage frame
Storing the multi-wire layers
for visual inspection and
wire tension measurement
f2
L
Δf
[mHz]
f [Hz]
Chamber accuracy
stereo angle < 35 μrad
wire position on PCB pad < 25 μm
cell width (wire pitch) < 1 μm
cell height (spacer) < 50 μm
wire tension < 0.1 g
PCB offset vs spoke < 50 μm
chamber length < 200 μm

## Slide 17: MEG2 DC Assembly
wire PC Boards
are lifted up by
adjustable arm ...
... and presented to
the end plates moved
closer by a few mm
A pressure
sensitive tape
holds them in
the correct
position above
the peek spacer.
At completion of each layer (12 sectors), the end plates are
moved away to the nominal length. The layer radial
coordinates and the tension of all wires are then measured.
After last layer has been mounted, the outer structural
carbon fiber cylindrical shell is placed and the inner shaft
is removed. The end plates are sealed, the inner mylar
cylinder is mounted, together with the extensions for the FEE

## Slide 18: MEG2 DC Spatial resolution
Single-hit resolution measured with three different prototypes.
Results are all in agreement, yielding a resolution of about 110 mm averaged throughout the cell.
Further improvements expected thanks to the implementation of a wide bandwidth front end electronics allowing for the exploitation of the cluster timing technique.
σ = 106 μm
(He/iC4H10 = 85/15)
arXiv:1605.07970
Notes: Several prototypes were made in order to make a reliable project for the new drift chamber. Results are in agreement, yielding a resolution of about 110 mm averaged throughout the cell.
CYLDCH will run with a low-Z counting gas mixture of Helium-Isobutane (He:iC4H10) in the fraction 85:15, which is the best compromise between track resolution and minimization
of MCS.

## Slide 19: MEG2 DC Expected Perf.
layer
occupancy
hits in 250 ns window
both views
segment fit + turn merging
discard short segments and isolated hits
full track fit
zoom
3D
track finding
and fit
signal track
michel tracks

## Slide 20: MEG2 DC Expected Perf.
Preliminary
track finding + fit
σϕ = 6.2 mrad
σϑ = 6.5 mrad
σp = 93.4 KeV/c
ε ≈ 90%

## Slide 21: Cluster Timing
acquired
signal
reconstructed
signal
From the ordered sequence of the electrons arrival times, considering the average time separation between clusters and their time spread due to diffusion, reconstruct the most probable sequence of clusters drift times:
￼
For any given first cluster (FC) drift time, the cluster timing technique exploits the drift time distribution of all successive clusters to determine the most probable impact parameter, thus reducing the bias and the average drift distance resolution with respect to those obtained from with the FC method alone.
Maximum
Possible
Spacing
algorithm
First
Cluster
Bias

## Slide 22: Cluster Counting
Experimental distribution
σ = 2.5%
≈ 3.2σ separation
Poisson distribution
σ = 1.7%
≈ 5σ separation
dNcl/dx
20% truncated mean
σ = 4.5%
≈ 1.4σ separation
100 samples 3.7 cm
σ = 3.7%
≈ 2.0σ separation
dE/dx
Thanks to the Poisson nature of the ionization process, by counting the total number of ionization clusters Ncl along the trajectory of a charged track, for all the hit cells, one can reach a relative resolution of Ncl-1/2.
(σ[%]=40.7 n-0.43 L[m]-0.32)
The data taken with a beam of μ and π at 200 MeV/c momentum at PSI, refer to a gas mixture He/iC4H10=95/5, Ncl = 9/cm, 100 samples, 2.6cm each at 45° (to avoid space charge effects), for a total track length of 3.7 m.
A 25 μm sense wire (gas gain 2x105), readout through a high bandwidth (1.7 GHz, gain 10) preamplifier, is digitized with a 2 GSa/s 1.1 GHz, 8 bits digital scope.
  (NIM A386 (1997) 458-469 and references therein)
μ−π 200 MeV/c
μ−π 200 MeV/c

## Slide 23: "Innovative DC" advantages
Gas containment – wire support separation and feed-through-less wiring
  allow to reduce material to ≈ 10-3 X0 for the inner cylinder and to a few x 10-2 X0 for the end-plates, including FEE, HV supply and signal cables (Mu2e proposal design: 1.5x10-3 X0 and 8x10-3 X0, respectively).
Feed-through-less wiring
  allows to increase chamber granularity and field/sense wire ratio to reduce multiple scattering and total tension on end plates due to wires
Cluster timing
  allows to reach spatial resolution < 100 μm for 8 mm drift cells in He based gas mixtures (such a technique is going to be implemented in the MEG2 drift chamber under construction)
Cluster counting
  allows to reach dNcl/dx resolution < 3% for particle identification (a factor 2 better than dE/dx as measured in a beam test)
Recipe for cluster timing/counting in He based gas mixtures:
FEE: 1 GHz BW, x10 gain (S/N ratio ≈ 8) - digitizer: 2 GSa/s sampling rate, >8 bits

## Slide 24: MEG2 DC Front End El.
Op-amp ADA4927 first gain stage: low noise, ultralow distortion, high speed, current feedback differential amplifier achieving wide bandwidth, low distortion, and low noise
(1.3 nV/√Hz) and low power consumption.
THS4509 second gain stage and output driver: wideband, fully differential op- amp, very low noise (1.9 nV/√Hz), extremely low distortion, ideal for pulsed applications.
Current
consumption
per channel:
50 mA at ±2.5V
(500 W/end plate)
pre-emphasis
noise immunity

## Slide 25: MEG2 DC Front End El.

## Slide 26: MEG2 DC Front End El.

## Slide 27: MEG2 DC Read Out
latest DRS (PSI) version
More generally, in a large DCH (60,000 channels, 30% occupancy,
1 μs drift time, 2 GSa/s - 12 bits digitization), at a 5 kHz trigger rate,
expect: > 100 GB/s!

## Slide 28: Clu-Tim/Cou Read Out
AD9625-2.0EBZ
UG534 ML605

## Slide 29: How long a drift chamber can be?
Gravitational sag
wire L along x-axis, f force per unit length along y, T tension of wire
f independent of x, parabolic solution
sag
wire radius r, volume density ρ, gravitational sag
Example:
L=2m; sense: 20μm W(Au), T=27g; field: 40μm Al(Ag), T=17g; guard: 50μm Al(Ag), T=26.5g

## Slide 30: How long a drift chamber can be?
Electrostatic stability
Simplified model of a wire of radius R at voltage V0, placed between two parallel grounded
planes at distance W/2 from the wire
linear charge density
cap. per unit length
perfectly symmetric geometry
assume a wire displacement Δ towards one plane
stability
condition

## Slide 31: How long a drift chamber can be?
Electrostatic stability: a numerical example MEG2
20μm W sense wire, V0= 1500V, W = 7mm, L = 2m, T = 0.25N
linear charge density
cap. per unit length
assume a wire displacement Δ = 50 μm
stability
condition
or
increase sense wire diameter to 50μm, T = 1.56N, for the same gain (same λ) need V0=1750V
or
drawback:
multiple scattering and force on end plates increase by a factor 2.52
(Ti(Sn) instead of W(Au)? to regain a factor 4 in mass and 10 in X0)

## Slide 32: A proposal for FCCee
  112 para-axial layers at alternating sign stereo angles, arranged in 16 equal azimuthal sectors;
  32 square, single sense wire, drift cells per sector (512 total per layer) increasing linearly as a function of layer radius;
  Cell sizes ranging from 6.3 mm to 25 mm from inner to outer radius;
  Alternating sign stereo angles in consecutive layers ranging from 40 to 160 mrad (constant azimuthal angular displacement)
  Length: 5000 mm; fully efficient up to cosϑ = 0.97 (≥16 hit)
  Inner Wall: made of 25 μm of Kapton plus 0.1μm of Au (1.2×10-4 X0) at Radius = 500 mm;
  Outer Wall: Sandwich of 8-ply C-fiber (0° and 90°, total of 250 μm) - 2.5 mm Rohacell30 - 8-ply C-fiber (8.0×10-3 X0) at Radius = 2060 mm (must support 20 Tons - check for buckling over 5 m);
  End plates:
    Wire cage (in analogy to Mu2e I-Tracker): 0.9 g/cm2 - 3×10-2 X0 (incl. power distr., decoupling C's, term. resistors and signal and HV cables).
    Gas envelope made of 8 ply (quasi-isotropic, 10×38 μm = 380 μm) C-fiber plus 0.3 μm Au, for a total of 0.090 g/cm2 – 3.0×10-3 X0;
  Gas: 90% He - 10% iC4H10 (δ = 4×10-4 g/cm3, X0 = 1410 m), - 12.5 p.i./cm, gas gain: 4×105 at V ≈ 1700 V on 50 μm wire, vdrift ~ 2.5 cm/μs – 0.47×10-3 X0/1m track
  Wires: - 57,344 sense (50μm Sn coated Ti); 290,816 field and guard (100μm Sn coated C); for a total equivalent thickness of 1.34×10-3 X0/1m track

## Slide 33: Expected spatial resolution
Expected Performance: Track parameters resolutions
n = 112, B = 2.0 T, Rout = 2.05 m, L = 3.0 m or 5.43x10-3 X0, σxy = 100 μm, σz = 1.0 mm
measurement
multiple scattering (gas + wires)

## Slide 34: Expected spatial resolution

## Slide 35: Expected spatial resolution

## Slide 36: Expected p. id. capabilities
σdE/dx/dE/dx =
= 5.4 L[m]-0.37 %
(Lehraus parametrization)
3.6% for L=3m
thresh.
cluster counting efficiency
ε = 80%
σ /dNcl/dx =
= ε×L×12.5/cm =
1.8% for L=3m
dNcl/dx
Particle separation power as a function of cluster counting efficiency for 2m tracks.
Cluster counting outperforms dE/dx for counting efficiencies as low as 20%.

## Slide 37: Conclusions
We propose for FCCee an innovative tracking system based on a
"ultra-light drift chamber with peculiar particle identification capabilities" using cluster timing/counting techniques.
It consists of a full stereo, single sense wire, square cells:
  Rin = 50 cm; Rout = 205 cm; L = 500 cm; 112 layers × (6.3 to 25.0 mm); 57,344 cells; >16 hits down to cosϑ = 0.97; stereo angles ranging from 40 mrad to 160 mrad;
  Inner cylindrical wall: 1.2×10-4 X0
  Outer cylindrical wall: 8.0×10-3 X0
  End plates (fully instrumented): 3.3×10-2 X0
  Gas + wires: .47×10-3 X0/1m track + 1.34×10-3 X0/1m track
Expected spatial resolutions: σrϕ < 100 μm, σz < 1 mm
Expected momentum resolution: Δp/p = 4.9×10-4, Δϑ = 0.9×10-4, Δϕ = 0.5×10-4 for p = 10 GeV/c and ϑ = 45°
Expected p. id.: π/κ separation > 3σ for p < 850 MeV/c and p > 1070 MeV/c
At current status of the art no need for major R&D