## Slide 1
1
Update on ECAL photon
optimisation
Matthias Weber
CERN
Notes: 1
1

## Slide 2: Reminder: Possible Models
2
 | # Layers | Absorber Thickness [mm] | Si [mm] | Other (gaps, etc) [mm] | Absorber Thickness [X0] | Si+Other Thickness [mm] | Si+Other Thickness [X0] | Total Thickness [mm] | Total Thickness [X0] | Calorimeter Thickness [mm] | Calorimeter Depth [X0]
CLIC_o2_v04 (current) | 17 | 2.4 | 0.5 | 1.5 | 0.65 | 2 | 0.06 | 74.8 | 12.01 | 129.2 | 22.85
 | 8 | 4.8 | 0.5 | 1.5 | 1.30 | 2 | 0.06 | 54.4 | 10.84 |  | 
CLICdet_20_10 (original CLIC_ILD) | 20 | 2 | 0.5 | 1.5 | 0.54 | 2 | 0.06 | 80 | 11.96 | 140 | 23.35
 | 10 | 4 | 0.5 | 1.5 | 1.08 | 2 | 0.06 | 60 | 11.39 |  | 
CLICdet_30_10 (possible high granularity option) | 30 | 1.5 | 0.5 | 1.5 | 0.41 | 2 | 0.06 | 105 | 13.89 | 155 | 22.58
 | 10 | 3 | 0.5 | 1.5 | 0.81 | 2 | 0.06 | 50 | 8.68 |  | 
CLICdet_30 | 30 | 2.65 | 0.5 | 1.5 | 0.72 | 2 | 0.06 | 140 | 23.22 | 139.5 | 23.22
 | 0 | 0 | 0 | 0 | 0.00 | 0 | 0.00 | 0 | 0.00 |  | 
CLICdet_40 (possible high granularity option) | 40 | 1.9 | 0.5 | 1.5 | 0.51 | 2 | 0.06 | 156 | 22.85 | 156 | 22.85
 | 0 | 0 | 0 | 0 | 0.00 | 0 | 0.00 | 0 | 0.00 |  | 
CLICdet_40x195 | 40 | 1.95 | 0.5 | 1.5 | 0.53 | 2 | 0.06 | 158 | 23.39 | 158 | 23.39
 | 0 | 0 | 0 | 0 | 0.00 | 0 | 0.00 | 0 | 0.000 |  | 
Values calculated from geometry xml file

## Slide 3: Energy deposited per unit length CLIC_o2_v04 (CLIC_18_7)
Raw energy deposited in sensitive volume as reported by Geant4, divided by sensitive volume thickness. Normalization starts to fail for high energies due to loss of energy in dead material between ECal and HCal
Raw energy deposited in sensitive volume as reported by Geant4, divided by sensitive volume thickness. Normalization to 1 for high energies affected by loss of energy in ECAL-HCAL gap.
4.8 mm
Already at 100 GeV sizable
amount sampled with the larger (“bad”) layers
3
2.4 mm/layer
22.85 X0

## Slide 4: Energy deposited per unit length CLIC_40
Raw energy deposited in sensitive volume as reported by Geant4, divided by sensitive volume thickness. Normalization starts to fail for high energies due to loss of energy in dead material between ECal and HCal
22.85
1.9 mm/layer
constant sampling performs good for vast energy range
4
Raw energy deposited in sensitive volume as reported by Geant4, divided by sensitive volume thickness. Normalization to 1 for high energies affected by loss of energy in ECAL-HCAL gap.
22.58 X0

## Slide 5: Resolution of total ECAL and HCAL Energy
Fits agree more or less, but constant term <1%, fit agrees with the trend of the points.
Thicker back layers deteriorate
resolution of high energy γ’s
Uniform det30 option better than CDR at high energy, comparable at low energies
The uniformly thin option (det40) is best behaving at high energy since it has the thinnest layers at the back (1.9 mm and 1.95 mm comparable) )
Single photons in the very central part of the detector (θ ≈ 90°) and ϕ ≈ 0°
5

## Slide 6: New Model CLIC_o3_v5
Uniform segmentation of ECAL
40 layers, absorber 1.9 mm, active material 0.50 mm
Model CLIC_o3_v05: other 2.65 mm: total thickness 20.2 cm
Model CLIC_o2_v05: other 1.5 mm: total thickness 15.6 cm cm
6

## Slide 7: Model CLIC_o3_v05: Barrel
Compare old and new model: use point phi=0, theta=90 degrees
Determine relative resolution via distribution of Erel = (E_ECAL+E_HCAL)/E_truth
measure of σ(Erel): RMS and gausian fit
Old model determined through iterative gaussian fits around mean within 3 sigmas
Energy (GeV) | 10 | 50 | 200 | 500
σ(Erel) | 4.72 | 2.11 | 1.12 | 0.777
RMS(Erel) | 4.82 | 2.13 | 1.17 | 0.903
CLIC_o2_v05 | 4.58 | 2.07 | 1.05 | 0.673
30 layers | 5.43 | 2.44 | 1.23 | 0.790

## Slide 8: Model CLIC_o3_v5: Calibration
Use particle gun samples for calibration, 10000 evts: 50 GeV K0L, 10 GeV μ
Photons at 10, 50, 100, 200 GeV, test very high E: 500, 1000, 1500 GeV
Produced with uniform distribution in phi and theta
Two campaigns, one restricted to barrel |cos θ| <0.7, second no restriction in θ
Simulated with DD4hep, use iLCSoft package at
/cvmfs/clicdp.cern.ch/iLCSoft/builds/2016-07-04/x86_64-slc6-gcc48-dbg
Slight modification of ECAL Digitization extraction code

## Slide 9: Model CLIC_o3_v5: Calibration Results
CALIBRECAL="35.652
10 GeV
50 GeV
CALIBRECAL="35.776"
CALIBRECAL="35.652”
Calibration pretty stable at different energies

## Slide 10: Model CLIC_o3_v5: Calibration Barrel
ECAL digitization uses photons contained in ECAL (99% of total deposited Energy)
Barrel
At 200 GeV
4271 evts
100 GeV
5737 evts
As expected requirement of 99 % is more stringent for higher energy samples
CALIBRECAL="35.619”
CALIBRECAL="35.617”
10

## Slide 11: Model CLIC_o3_v5: Calibration at high Energy
Expect slight modification of EM scale in HCAL based on previous study (based on one point in phi=0):
Fit a line in 2D plot of ECAL vs HCAL Energy:
Moderate modification:
ECAL: 0.991
HCAL: 0.950

## Slide 12: Model CLIC_o3_v5: Calibration: whole detector
ECAL digitization uses photons contained in ECAL (99% of total deposited Energy)
Barrel
At 200 GeV
4271 evts
Barrel + Endcap
Use values from samples up to 200 GeV, for Barrel and EndCap mix two peaks start to appear, seems due to different losses in cracks etc at higher energies
12

## Slide 13: Model CLIC_o3_v5: Barrel Resolution
Determine relative resolution via distribution of Erel = (E_ECAL+E_HCAL)/E_truth
measure of σ(Erel): RMS and gausian fit
Energy (GeV) | 1 | 5 | 10 | 15 | 30 | 50
σ(Erel) | 15.14 | 6.95 | 4.873 | 4.00 | 2.829 | 2.281
RMS(Erel) | 22.94 | 7.175 | 4.953 | 4.056 | 2.914 | 2.355
Mean (Erel) | 0.957 | 0.996 | 0.998 | 0.999 | 0.999 | 0.997
Energy (GeV) | 100 | 200 | 500 | 1000 | 1500
σ(Erel) | 1.69 | 1.27 | 0.955 | 0.810 | 0.775
RMS(Erel) | 1.773 | 1.366 | 1.133 | 1.065 | 0.953
Mean (Erel) | 0.997 | 0.997 | 0.998 | 0.999 | 1.001
13

## Slide 14: Model CLIC_o3_v5: Barrel Resolution
Determine relative resolution via distribution of Erel = (E_ECAL+E_HCAL)/E_truth
measure of σ(E): RMS and gausian fit
1500 GeV: not large non gaussian tails
14

## Slide 15: Model CLIC_o3_v5: Barrel Resolution
Determine relative resolution via distribution of Erel = (E_ECAL+E_HCAL)/E_truth
measure of σ(E): RMS and gausian fit, now moderate energy 50 GeV, non gaussian tails smaller
15

## Slide 16: Model CLIC_o3_v5: Barrel Resolution
Determine relative resolution via distribution of Erel = (E_ECAL+E_HCAL)/E_truth
measure of σ(E): RMS and gausian fit, very low energy of 1 GeV start to see inefficincies in reconstruction (not the case for to 5 GeV photons)
16

## Slide 17: Model CLIC_o3_v5: Barrel Overview
Summary plot of resolution in ECAL barrel
17

## Slide 18: Model CLIC_o3_v5: Barrel leakage
Energy (GeV) | 1 | 5 | 10 | 15 | 30 | 50
leakage | 0.131 | 0.343 | 0.490 | 0.576 | 0.7731 | 0.979
Mean of Distribution E_HCAL/(E_ECAL+E_HCAL) in %
Energy (GeV) | 100 | 200 | 500 | 1000 | 1500
leakage | 1.30 | 1.76 | 2.58 | 3.4 | 4.19
E-HCAL, 1500 GeV
E-HCAL, 200 GeV
18

## Slide 19: Old study: Leakage
More X0’s perform generally better (as expected)
Leakage improved for high segmentation with thicker layers (1.95 vs 1.9 mm absorber thickness)
Mean
Example at 1500 GeV
19

## Slide 20: Summary
Study of new model CLIC_o3_v05: no drastic change observed in resolution of photons compared to 40 layer CLIC_o2_v05
Calibration constants stable across different energies
Resolution values drop to around 1 % at very high energies
Resolution distributions
At higher energies non gaussian tails to the lower side
at high energies non negligible contribution from HCAL
20

## Slide 21: Values for X0 calculation
Tungsten 24: 0.37 cm
Silicon: 9.36 cm
G10 (density 1.7 g/cm3): 19.4 cm
Air: 30420 cm
GroundOrHVMix: 86 % Copper
siPCBMix: 82% Copper
21