## Slide 1: Dose reduction in PET scans
Paul Lecoq
CERN, Geneva
This work is performed
in the frame of the ERC Advanced Grant Agreement N°338953–TICAL
and the COST Action TD1401

## Slide 2
PET/CT Combine anatomic and functional/MI information
Geneva University
Invention of the PET/CT
CT 10 to 20 mSv
PET about 10 mSv
Standard radiography
0.1 mSv

## Slide 3
PET/MR Combine anatomic and functional/MI information
CT in PET-CT
Anatomic
Low soft tissue contrast
Fast
Attenuation correction
10 to 20mSv
MR in PET-MR
Anatomic + some functionality
High soft tissue contrast
Slow (depending on MR sequences)
No attenuation correction
NO DOSE

## Slide 4
Commercial PET/MR scanners

## Slide 5
Detector B
Detector A
e-
e+
Patient
tA
tB
Dt = tA – tB = [(d+d1) – (d-d1)]/c
d1 = c Dt/2
d
d1
Time of Flight PET
dt (ps)
10
100
300
500
dx (cm)
SNR*
0.15
1.5
4.5
7.5
16
5.2
3.0
2.3
* SNR gain for 40 cm phantom
SNRTOF = √(2D/cDt) · SNRconv
10ps TOF resolution
dx = 1.5mm
SNRTOF/SNRCONV= 16!

## Slide 6
This is a clear-cut case to shed light on TOF-PET with CTR < 10 ps FWHM and raise a challenge on
≤ 1mSv PET exposure = 1 flight Paris-SF
positron tomography
March 1999: Breitling Orbiter III circumnavigate the globe in 19 days 1 hour 49 minutes and won the Budweiser Cup
1992: FAI raised a challenge for the first balloon circumnavigation
The 10ps challenge
Courtesy of C. Morel, CPPM
July 2016: Solar Impulse closed the loop of
the Round-the-World without fuel attempt
Notes: FAI: Fédération Aéronautique Internationale,
International Aeronautic Federation

## Slide 7: Best timing results at CERN
LSO:Ce, 0.4%Ca, melmount coupled to 3x3mm2 NUV SiPM from FBK, 55%PDE
S. Gundacker et al., CERN
2x2mm2 section
511 KeV
150GeV muons
5MeV deposited

## Slide 8: The detection chain
q2
SiPM
Crystal
electronics
g
Dt
tkth pe =
Dt
Conversion depth
+ tk’ ph
Scintillation
process
+ ttransit
Transit time
jitter
+ tSPTR
Single photon
time spread
+ tTDC
TDC
conversion time
Random deletion 1
Absorption
Self-absorption
Random deletion 2
SiPM PDE
Unwanted pulses 2
DCR
Unwanted pulses 1
DCR, cross talk
Afterpulses
Light generation
Light transport
Photoconversion

## Slide 9: Factors influencing scintillator time resolution
Besides all factors related to photodetection and readout electronics the scintillator contributes to the time resolution through:
  The light production mechanism
    Light yield,
    Rise time,
    Decay time
P. Lecoq et al, IEEE Trans. Nucl. Sci. 57 (2010) 2411-2416
  The light transport in the crystal
    Time spread related to different light propagation modes
  The light extraction efficiency (LYLO)
    Impact on photostatistics
    Weights the distribution of light propagation modes

## Slide 10: SPTR dependence on SiPM area
SiPM bias overvoltage (V)
SiPM bias overvoltage (V)
M. Nemallapudi et al, Single Photon Time Resolution of State of the Art SiPMs, Submitted to JINST, May 1016

## Slide 11: Avalanche development in a 50mm diameter SPAD
Ivan Rech, Politecnico di Milano
Notes: Two factors determine the timing resolution of a SiPM.
the time jitter of the avalnche development within a SPAD
The time transit spread of the signals from different SPADs to the collecting anode.
On the first point the group of Cova and Rech in Milano are studying the influence of the SPAD geometry on the avanche development.
This is a measurement, not a simulation! This is slow motion move showing in 30sec the 1.2ns development of the avalanche in a 50microns circular SPAD
***
photon-triggered secondary avalanches start to develop
***
(The image is the level of the current, but the emitted light is proportional to the current so to end is the same- the emission is centered around 1um and reach 4-500nm in the blue and 1.3 1.4 in the IR-Yes the point is the assisted mechanism- the rim is more intense because it is less quenched by the space charge resistor.)

## Slide 12
Ultrafast
Nano-crystals
ZnO:Ga
  No TOF
  700 ps
  500 ps
  300 ps
P Lecoq et al. Nucl. Instrum. Meth. A 718 (2013) 569
Fast sintillation timing: the way to reconstructionless TOFPET
Digital 3D SiPM
Photonic crystals
Notes: ZnO:Ga 20 fold increase in photon density as compared to LSO:Ce, Ca. 500ps compared to 30ns, light yiled 1/3 of LSO:Ce, Ca

## Slide 13
EXPLORER.ucdavis.edu
EXPLORER: A Total-Body PET Scanner

## Slide 14
Ramsey Badawi
Simon Cherry
EXPLORER: A Total-Body PET Scanner
Notes: to consolidate the full portfolio of the applications case upon which the securing of funding rests
bottom line on this one is that although the project initially sounds pretty technologically daunting, we actually have the technology

## Slide 15
 | Siemens mCT | 
Axial FOV (cm) | 21.8 | 200
Ring diameter (cm) | 84.9 | ~80
Scintillator | LSO | LYSO
Spatial resolution (mm) | 4.1 | <4
TOF resolution (ps) | 530 | <400
Integrated CT unit? | Yes | Yes
DOI Measurement? | No | Yes
No
EXPLORER: Design parameters
Notes: Why…does not surround the patient with detectors, much of the radiation is missed.

## Slide 16
Up to 40-fold reduction in dose
  Whole-body PET at ~ 0.15 mSv
  Annual natural background is ~ 2.4 mSv
  Return flight (SFO-LHR) is ~ 0.11 mSv
  PET can be used with minimal risk – new populations
EXPLORER or 10ps TOFPET
Conventional PET
EXPLORER or 10ps TOFPET Image Gently (Low Dose)

## Slide 17
Clinical Need
  15M babies born pre-term (< 37 weeks) / year
  1.1M die because of pre-term complications / year
  Many surviving pre-term babies have problems as adults (disabled, diabetes, hypertension, heart disease, obesity)
Image Gently (Low Dose)
Maternal-Fetal interaction
The potential for low dose functional studies in maternal-fetal medicine using combined PET and MRI
Terry Jones and Thomas F. Budinger. Journal of Nuclear Medicine 2013, 54: 2017-2018
Identify Problems In Utero
  Trace nutrient transport from mother to fetus (oxygen, glucose, amino acids, fatty acids)
  Image at < 0.05 mSv

## Slide 18
Thank you
Notes: PET and CT are very powerfull imaging techniques but theyr are using ionizing radiation.
Then it depends what you want to see. Imaging the surface of the patient is not damaging but does not bring much information, at least from the medical point of view.
Ionizinng radiation must of course be used with soeme care. But technological advances open the way to examinations at the level or even less than the natural background