## Slide 1: Photon Spectrometry with Dosepix
Winnie Wong, University of Houston and Medipix2 Collaboration
On behalf of the Medipix Design Team at CERN
Eurados Workshop
17 October 2016

## Slide 2: Outline
Hybrid pixel detectors
Single quantum processing
Medipix
Dosepix
    Count rate
    X-ray spectroscopy
            Primary beam
            Scatter radiation

## Slide 3: Hybrid Pixel Detectors
Application-Specific Integrated Circuit (“ASIC”, aka chip)
Bump bonds (electrical connections)
Semiconductor sensor

## Slide 4: Hybrid Pixel Detectors
Appropriate sensor material for application: e.g. Si, GaAs, CdTe
Optimised electronics for speed, power consumption, functionality, etc.
100% fill factor and 100% use of pixel area for circuit implementation
  Permits single quantum processing
Cross-section of the Dosepix detector
Image courtesy of Fraunhofer IZM
Silicon sensor
Copper pillar (bump bond)
Transistors and metal connections on the ASIC

## Slide 5: Single Quantum Processing with Hybrid Pixel Detectors
Analogue
threshold
ToT = 29 x 100 ns = 2.9 μs = 20 keV
10 MHz Clock
V
t
ToT = 16 x 100 ns = 1.6 μs = 13 keV
100 ns
E.g. Medipix2,
Medipix3
Photon counting
E.g. Timepix,
(Timepix2),
Timepix3,
Dosepix
Time over threshold (ToT)

## Slide 6: Photon Interaction with Sensor
n-
Semiconductor Sensor
Bump bond (electrical contact)
Chip (aka ASIC)
+ -
+ -
+ -
+ -
+ -
+ -
- +
- +
- +
- +
- +
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
+ -
+ -
+ -
+ -
+ -
+ -
- +
- +
- +
- +
- +
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
Charge sharing
+ -
+ -
+ -
+ -
+ -
+ -
- +
- +
- +
- +
- +
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
Compton scattering
Photoelectric effect
Notes: Photoelectric effect, signal is seen by one pixel
Photoelectric effect, charge sharing due to diffusion
Compton scattering

## Slide 7: Dosepix
Main application: dosimetry
Developed by: CERN, U. Erlangen, IBA Dosimetry
ASIC:
  16x16 pixels=256 parallel spectrometers:
    1 global voltage threshold (programmable)
    16 energy bin thresholds per pixel (programmable)
    12-bit ToT: energy
    Automatic binning in pixel
Sensor:
  300 μm Si, segmented into
    64 pixels of 55x55 μm2
    192 pixels of 220x220 μm2
3.5 mm
3.5 mm
Si sensor
Dosepix ASIC
Photo source: T. Gabor
Notes: 3048 (15x256) is the theoretical number of bins that can be programmed on the chip, but in practice you would want to assign the same bin programming to multiple pixels to increase the active area per bin to gather more statistics. I typically program 60 unique bins across the chip, with the energy range dependent on the measurement I want to take.

## Slide 8
220 μm
220 μm
220 μm
Dosepix ASIC
(photo)
Silicon sensor
(drawing)
3.52 mm
3.52 mm
3.52 mm
3.52 mm
55x55 μm2 active area with 220 μm pitch

## Slide 9: Pixel Block Diagram
W. Wong, PhD Thesis, 2012

## Slide 10: Spectral Response
55 μm
Pixels
Am241 γ
(26.4 keV, 59.5 keV)
Cu
fluorescence
(8 keV)
Spectrum of gamma emissions from an Am241 source, filtered to block alphas
220 μm
Pixels
60 bins programmed from 5 – 64 keV
(1 keV per bin)
Notes: Assembly29, 120 files, 60 seconds each
Tradeoff between the two pixel sizes: large pixels have less charge sharing, small pixels see lower fluxes

## Slide 11: Pileup
Reset current = 7.4 nA
Reset current = 2.5 nA
Pileup
Notes: 15.38 15.38
1.21 -0.27
11.72 13.3
1.27 4.05
13.32 4.05

## Slide 12: Characterised Response
E. Frojdh et al., NIMA 2015
 |  | Energy Resolution, FWHM (220 μm pixels) |  |  | Maximum Linear Count Rate (220 μm pixels) | 
Reset Current | Minimum Threshold | @ 8.5 keV | @ 20 keV | @ 40 keV | @ 30 keV | @ 40 keV
2.6 nA | 3.54 keV | 1.43 keV (16.8%) | 2.81 keV (14%) | 3.18 keV (8%) | 1.69 Mcounts/s/mm2 | 1.67* Mcounts/s/mm2
7.4 nA | 4.27 keV | 2.17 keV (25.5%) | 2.91 keV (14.5%) | 3.79 keV (9.5%) | 4.11 Mcounts/s/mm2 | 3.57 Mcounts/s/mm2
*Note: Pileup starts well below the maximum linear count rate (at approx. 105 counts/s/mm2) and will distort spectrum measurements.
Power Consumption = 15 mW to process 105 counts/s/mm2
under nominal chip settings

## Slide 13: RQR Measurement Setup
Dosepix location, perpendicular to and centred on the beam
5 m
Toshiba KXO-80G
X-ray generator with a tungsten anode
RX Toshiba KX0-80G at Institute for Radiation Physics, University Hospital of Vaud

## Slide 14: Background Measurement
No beam, no hits (“noise free”)
60 bins programmed from
5 – 150 keV

## Slide 15: Primary Beam Spectrum RQR-5*, 10 mA @ 2s
1 “frame” of data in the 220 μm pixels
kVp
*Reference Radiation Qualities, IEC 61267
60 bins programmed from
5 – 150 keV

## Slide 16: Primary Beam Spectra 10 mA @ 2s
1 “frame” of data in the 220 μm pixels
RQR-5
RQR-7
RQR-9
Characteristic X-rays?
Emean=25.0 keV
Emean=26.2 keV
Emean=27.4 keV
Notes: Mean and Max energy is already useful information for radiation protection officers designing appropriate shielding. This information would be unavailable in passive dosimeters or even active dosimeters if what they measure is integrated energy deposition rather than fluence per energy

## Slide 17: Primary Beam Spectra 10 mA @ 2s
RQR-5
RQR-7
RQR-9
Log
220 μm pixels
55 μm pixels
Emean=25.0 keV
Emean=26.2 keV
Emean=27.4 keV
Emean=20.9 keV
Emean=21.8 keV
Emean=22.6 keV
Emax =127 keV
Emax =150 keV
Emax =150 keV
Emax =68 keV
Emax =101 keV
Emax =127 keV
21.2 kHits/s/mm2
31.0 kHits/s/mm2
67.5 kHits/s/mm2
20.7 kHits/s/mm2
30.7 kHits/s/mm2
67.4 kHits/s/mm2
Pileup of 28 counts (0.066%)
Pileup of 27 counts (0.040%)
Pileup of 21 counts (0.016%)
Pileup of 0 counts
(0%)
Pileup of 3 counts (0.006%)
Pileup of 4 counts
(0.003%)

## Slide 18: Primary Beam Spectra for Increasing mAs
RQR-5
RQR-7
RQR-9
Log
220 μm pixels
55 μm pixels

## Slide 19: Primary Beam Fluence for Increasing mAs
RQR-5
RQR-7
RQR-9
220 μm pixels
55 μm pixels

## Slide 20: Scatter Radiation Measurements
RQR-5
220 μm
Pixels
55 μm
Pixels
Log

## Slide 21: Scatter Radiation Measurements
RQR-5
220 μm
Pixels
55 μm
Pixels
On-going work:
Correct for Si absorption efficiency
Develop spectrum reconstruction
  Study impact of charge sharing, Compton Scattering, etc.
Calculate air kerma
Characterise scatter radiation in CT scan
Scatter Radiation
in the 220 μm pixels
20 mAs
RQR-9 Primary Beam
in the 220 μm pixels
20 mAs
Linear
Characteristic X-rays?
Log
Emean, raw = 27.4 keV
Emean, corrected= 45.1 keV

## Slide 22: Scatter Radiation in CT Scan
Patient imaging facility, University Hospital of Vaud

## Slide 23
Scatter radiation spectra during scan of a QRM anthropomorphic thorax phantom:120 kVp, 200 mA

## Slide 24: Summary
Hybrid pixel detectors permit choice of appropriate sensor material
Absorbed energy can be measured by:
  Pulse Height
  Time over Threshold (ToT)
Our detectors measure fluence at a given energy
  Dosimetric endpoints can be calculated offline with energy dependent conversion factors
  Raw fluence data remains available for future calculations or definitions

## Slide 25: Extra slides

## Slide 26
Scatter radiation spectra during scan of a QRM anthropomorphic thorax phantom:120 kVp, 200 mA
Eurados WG9 Meeting, CERN, 17.10.2016
winnie.wong ‘at’ cern.ch

## Slide 28: Energy Calibration
Characterisation with monochromatic synchrotron radiation at ANKA Synchrotron
E. Frojdh et al., NIMA 2015
ToT relationship with photon energy for two programmable reset currents
Ikrum: preamplifier reset current
IKRUM 15 = 2.5 nA
IKRUM 60 = 7.4 nA
Notes: What causes the two slopes

## Slide 29: Count Rate
E. Frojdh et al., NIMA 2015
Measurements with 40 keV monochromatic synchrotron radiation at ANKA Synchrotron
Measured
with 2.6 nA reset current (IKRUM=15),
in 220 μm pixels

## Slide 30: Medipix3
Main application: X-ray imaging
Developed by: Medipix3 Collaboration
ASIC:
  1.4 cm x 1.4 cm active area (can be tiled for more area)
  Configurable pixel size
      55x55 μm2 (65k pixels)
      110x110 μm2 (16k pixels)
  2-8 thresholds depending on mode
  (programmable)
  6, 12, or 24-bit photon counting: intensity
  Charge summing mode
Sensor:
  300 μm Si, 2 mm CdTe, etc.
1.4 cm
2 x 2 tiles of Medipix3 detectors
Photo source: diamond.ac.uk
1.4 cm

## Slide 31: Photon Spectroscopy with the Photon Counting Medipix3
Single Pixel Mode (SPM)
Notes: Charge sharing due to diffusion

## Slide 32: Photon Spectroscopy with the Photon Counting Medipix3
Single Pixel Mode (SPM)

## Slide 33: Photon Spectroscopy with the Photon Counting Medipix3
Charge Summing Mode (CSM)

## Slide 34: Photon Spectroscopy with the Photon Counting Medipix3
L. Tlustos, PhD Thesis, 2005
Simulation of spectroscopy with monochromatic
10 keV beam

## Slide 35: Photon Spectroscopy with Medipix3 & Si
SPM: Single Pixel Mode
CSM: Charge summing mode
ASIC: Medipix3RX
Sensor: 300 μm Si
E. Frojdh et al. JINST 2014
Threshold scans

## Slide 36: Timepix
Photo source: J. Jacubek
Main application: particle tracking
Developed by: Medipix2 Collaboration
ASIC:
  65k square pixels, 55 μm pitch
    1 threshold (programmable)
    14-bit Time over Threshold (ToT): energy
    14-bit Time of Arrival (ToA): time
Sensor:
  300 μm Si, 2 mm CdTe, etc.
  Micro-channel plate,
  gain gas grid, etc.

## Slide 37: Mixed Field Dosimetry with Timepix
241Am (α,γ)
55Fe (γ)
90Sr (β)
α
γ
5.8 keV gammas
5.5 MeV alphas,
59.5 keV gammas
0.5 MeV betas
Notes: Energy of Fe55 gamma=5.8 keV

## Slide 38: Mixed Field Dosimetry with Timepix
241Am (α,γ)
55Fe (γ)
90Sr (β)

## Slide 39: Mixed Field Dosimetry with Timepix
241Am (α,γ)
90Sr (β)
109Cd (γ)
Notes: Can also get directional information from time of arrival

## Slide 40: Heavy Ion Dosimetry with Timepix
180 MeV/A Ne
600 MeV/A Si
δ
60o
90o
11B
11B
L. Pinsky et al., Proc IEEE Aerospace Conf, 2007
0o

## Slide 41: Space Dosimetry with Timepix
Photo source: NASA

## Slide 42: Space Dosimetry with Timepix
Pinsky et al., Presentation at VCI 2016, Vienna
Stoffle et al., NIMA, 2015
Kroupa et al., Life Sciences, 2015
(Timepix) REM on ISS
Intra-vehicular
TPEC
ISS TPEC
For each track:
Estimate species by footprint
Calculate dE/dx, angle of incidence, path length, etc.
Calculate dose in silicon
Convert to dose in water
Notes: Exposure times are dynamic, aiming for 4% occupancy. This example SAA frame had an exposure time of 0.9s