## Slide 1
A neutron detector for pulsed mixed fields: Speaker: Dott. Marco Caresana Dipartimento di Energia – Cesnef Politecnico di Milano

## Slide 2
The problem is known since the forties. In 1948 C. H. Westcott [2] reported “A feature common to all machines giving particles of the highest energies is that their output is not continuous, but occurs in bursts or 'pulses' separated by relatively long intervals during which the machine gives no output…The pulsed nature of the source is of no consequence when it is used to manufacture radioactive isotopes, ….There remains, however, a large field of investigation where quantitative results are of importance and where methods involving electrical counting of the instantaneous products of the reaction are pre-eminently

## Slide 3
There are plenty of practical situations with particle accelerators used for both scientific and medical applications [3], where the time structure of the secondary stray radiation limits the use of active monitors. Usually the time duration of a single burst can range from few µs to about 1 ms with a typical repetition rate in the range 10 Hz – 100 Hz [4,5,6]
[3] S. Agosteo, Radiation Measurements 45 (2010) 1171-1177
[4] Klett, A., Leuschner, A., IEEE 2007 Nuclear Science Symposium & Medical Imaging Conference, Conference Records, Oct 27-Nov 3 2007, Honolulu, Hawaii, USA.
[5] CERN-DGS-2012-036-RP-TN, Environmental measurements and instrument intercomparison around the PS accelerator complex.
[6] W.A. Barletta, Nuclear Instruments and Methods in Physics Research A (2010) 618 69–96

## Slide 4: Pulsed neutron fields (PNF)
It is well-known that neutron detectors, based on neutrons counting, generally suffer from dead-time effects and have strong limitations when measuring in pulsed radiation fields.
This is a major issue at particle accelerators, where pulsed neutron and gamma fields are present because of beam losses at e.g. targets, collimators and beam dumps.

## Slide 5: Pulsed neutron fields (PNF)
An ideal neutron survey meter for PNF should meet the following requirements:
1. capability to withstand very high instantaneous neutron fluxes with little or no saturation;
2. sensitivity comparable to that of commercially available rem counters;
3. capability to measure correctly the intensity of a single neutron burst;
capability to reject the photon contribution that accompanies the neutron field.
At present there is no instrument capable to fulfill all the requirements

## Slide 6: The LUPIN
LUPIN (Long interval Ultra wide dinamic Pile-up free Neutron detector)
The signal LogOut is acquired with an ADC (10 MHz) and processed via software (current integration over a user settable timebase). The system measures the charge generated by a neutron interaction
The BF3 detector is placed inside a cylindrical polyethylene moderator

## Slide 7: The LUPIN output pulse
Chart: time µs
Chart: time µs
Chart: time µs
Example of a neutron current pulse

## Slide 8: Charge calibration
The charge calibration is the measurement of the mean collected charge (MCC) expected from a single neutron interaction, expressed in fC/n
BF£ counter MCC = 570 fC with HV=1180V
The total charge, measured during the interaction of a radiation burst involving several neutron interaction, divided by MCC is the number of neutron interactions

## Slide 9: Signal analysis
The current profile produced by a neutron burst interaction is shown on the right. The number of interacting neutrons N can be calculated using the following equation:
In the example N=107, integrating over 4ms.
About 50 neutron interactions occur in about 300 μs, corresponding to an interaction rate around 170 kHz.
The sensitivity of the LUPIN in terms of ambient dose equivalent H*(10) is 2.17 0.22 count per nSv.
The acquisition shown in figure corresponds a burst intensity of 49 nSv

## Slide 10: Measurement at S. Raffaele hospital
Electron radiotherapy LINAC: Varian Clinac® DHX - Dual energy
beam directed on the treatment couch, irradiation field of 5 x 5 cm2.
The detector was placed on the therapy couch at 100 cm from the isocenter.
LINAC was operating 6 MV or 18 MV.
6 MV
18 MV

## Slide 11: Conclusions
A number of developments are planned (and most of them implemented) to improve the performance of the LUPIN.
a reading in terms of ambient dose equivalent would be desirable, to employ the instrument as a rem-counter. This implies the use of a specific moderator that is still being designed.
the behavior of the detector needs to be analyzed over a wide range of burst intensities in order to evaluate the linearity of its response.
improvements on the detector sensitivity, on the isotropy of its response and on the signal analysis process can be achieved by employing a 3He proportional counter and by optimizing the electronics.

## Slide 12: LUPIN He-3

## Slide 13: LUPIN BF3
Chart: Time [μs]

## Slide 14: Test at HZB 68 MeV protons on W
The proton beam is delivered in burst with duration of 1 µs, 10 µs, 40 µs

## Slide 15: LUPIN output signal resulting from a neutron burst 1µs long

## Slide 16: Test at HZB 68 MeV protons on W
Setting number | Ion current [pA] | Burst current [nA] | Burst length [μs] | Burst charge Qi [fC] | Reference burst yield [nSv per burst] | Average dose rate [μSv·h-1] | Burst dose rate [Sv·h-1]
1 | 0.5 | 5 | 1 | 5 | 0.077 | 27.72 | 0.28
2 | 1.5 | 15 | 1 | 15 | 0.231 | 83.16 | 0.83
3 | 3 | 30 | 1 | 30 | 0.462 | 166.32 | 1.66
4 | 5 | 50 | 1 | 50 | 0.770 | 277.2 | 2.77
5 | 10 | 100 | 1 | 100 | 1.540 | 554.4 | 5.54
6 | 25 | 250 | 1 | 250 | 3.850 | 1386 | 13.86
7 | 50 | 500 | 1 | 500 | 7.700 | 2772 | 27.72
8 | 75 | 750 | 1 | 750 | 11.550 | 4158 | 41.58
9 | 100 | 1000 | 1 | 1000 | 15.400 | 5544 | 55.44
10 | 250 | 250 | 10 | 2500 | 38.500 | 13860 | 13.86
11 | 500 | 500 | 10 | 5000 | 77.000 | 27720 | 27.72
12 | 1000 | 1000 | 10 | 10000 | 154.000 | 55440 | 55.44
13 | 3000 | 800 | 40 | 32000 | 492.800 | 177408 | 44.35

## Slide 17
Chart: LUPIN 3He
| 0.72996 | 1.57542 | 3.920840000000001 | 7.735419999999999 | 11.44528 | 15.5001 | 39.34238000000001 | 75.537 | 156.26688
Raw data | 3.9984 | 7.1599 | 10.6582 | 13.6512 | 30.9718 | 51.2413 | 84.0374
corrected data | 4.0881 | 7.4525 | 11.3199 | 14.756 | 37.3092 | 71.2704 | 155.8841

## Slide 18
Chart: LUPIN BF3
| 0.45892 | 0.7623 | 1.55848 | 3.89004 | 7.84322 | 11.73634 | 16.19618000000001 | 39.05902 | 74.8132 | 153.75976 | 471.26772
Raw data | 1.5886 | 4.042 | 8.178 | 12.22 | 16.92 | 42.3 | 77.55 | 141 | 393.86
BI corrected data | 1.59 | 4.0511 | 8.2152 | 12.3032 | 17.0799 | 43.3135 | 81.0258 | 152.9277 | 503.5725

## Slide 19: Instruments comparison

## Slide 20: Instruments comparison

## Slide 21: Space charge effect
The lack of linearity for the LUPIN is due to the space charge effect that shields electrostatically the inner anode producing a decrease in the multiplication factor.
The different behavior of the two LUPIN version is in the different geometry.
The small detector has an higher charge density

## Slide 22: Neutron photon discrimination
There are two possible scenarios
A mix and steady photons + neutrons field
A mix and pulsed photons + neutrons field
The techniques to cope with the two situations are different

## Slide 23: mix and steady photons + neutrons field
Chart: time µs
Neutron pulse without photon field
Neutron pulse photon field
The technique is to acquire the baseline variation using a pre-trigger

## Slide 24: mix and steady photons + neutrons field

## Slide 25: mix and steady photons + neutrons field
Chart: Time µs
Chart: Time µs

## Slide 26: mix and steady photons + neutrons field

## Slide 27: Instruments comparison at CERN PS

## Slide 28: Instruments comparison at CERN PS

## Slide 29: Instruments comparison at CERN PS

## Slide 30: Instruments comparison at CERN IRADMAT

## Slide 31: Instruments comparison at CERN IRADMAT
Two meters upstream the target
Instruments position

## Slide 32: Instruments comparison at CERN IRADMAT

## Slide 33
Chart: Expected H*(10) [nSv/burst]
| 3.919500000000001 | 4.442100000000001 | 4.964700000000001 | 6.532500000000001 | 6.968 | 7.926100000000002 | 10.1036 | 10.2778 | 16.11350000000001 | 16.20059999999996 | 16.7232 | 17.5942 | 37.19170000000001 | 37.88850000000001 | 61.841 | 61.841 | 82.74500000000003 | 84.48700000000002 | 166.1868 | 193.8846 | 365.82 | 373.6590000000005 | 871.0000000000001 | 871.0000000000001 | 1750.71 | 1768.13 | 3579.810000000002 | 3658.2 | 6863.48 | 6959.290000000001
LUPIN | 3.36 | 3.36 | 6.24 | 8.64 | 7.68 | 6.24 | 12.48 | 9.6 | 12.96 | 20.64 | 13.92 | 15.36 | 32.16 | 33.12 | 60.96 | 41.76 | 59.04 | 64.32 | 152.64 | 148.8 | 278.88 | 288 | 587.04 | 606.72 | 1012.32 | 1009.65 | 1778.4 | 1752.48 | 2719.2 | 2792.16
Fit | 3.2227 | 3.6582 | 3.8324 | 3.9195 | 4.4421 | 5.1389 | 5.226 | 5.4002 | 5.6615 | 6.1841 | 6.6196 | 8.5358 | 9.581 | 9.581 | 9.581 | 10.452 | 11.7585 | 12.3682 | 14.8941 | 15.2425 | 15.678 | 18.291 | 19.162 | 40.937 | 40.937 | 76.7351 | 77.519 | 87.1 | 95.81 | 162.006 | 162.006 | 191.62 | 200.33 | 383.24 | 409.37 | 905.84 | 905.84 | 1742 | 1768.13 | 3579.81 | 3579.81 | 6924.45 | 6968
Chart: Expected H*(10) [nSv/burst]
| 5.226000000000001 | 5.226000000000001 | 5.661499999999997 | 6.706700000000001 | 16.7232 | 18.90069999999996 | 179.426 | 191.62 | 330.9799999999996 | 357.1100000000001 | 905.8399999999995 | 905.8399999999995 | 1785.55 | 1820.39 | 3675.62 | 3710.46 | 6933.160000000004 | 6950.580000000001 | 6985.420000000001
BIOREM | 5.4694 | 4.375 | 5.4694 | 3.8278 | 13.1278 | 12.5806 | 77.8333 | 70.1667 | 114.7222 | 120.7778 | 220.0833 | 219.25 | 320.8333 | 341.9444 | 506.9444 | 499.7222 | 691.6667 | 677.7778 | 691.6667
Fit | 3.2227 | 3.6582 | 3.8324 | 3.9195 | 4.4421 | 5.1389 | 5.226 | 5.4002 | 5.6615 | 6.1841 | 6.6196 | 8.5358 | 9.581 | 9.581 | 9.581 | 10.452 | 11.7585 | 12.3682 | 14.8941 | 15.2425 | 15.678 | 18.291 | 19.162 | 40.937 | 40.937 | 76.7351 | 77.519 | 87.1 | 95.81 | 162.006 | 162.006 | 191.62 | 200.33 | 383.24 | 409.37 | 905.84 | 905.84 | 1742 | 1768.13 | 3579.81 | 3579.81 | 6924.45 | 6968
LUPIN
BIOREM
Chart: Expected H*(10) [nSv/burst]
| 4.006600000000002 | 4.354999999999991 | 7.839000000000001 | 15.0683 | 16.98449999999997 | 345.787 | 345.787 | 871.0000000000001 | 905.8399999999995 | 3510.130000000001 | 3623.36 | 6793.800000000001 | 6837.35
Ar | 5.2522 | 5.5277 | 9.7477 | 14.6403 | 17.4828 | 335.0589 | 267.6178 | 845.4289 | 862.0656 | 3275.4556 | 3429.3 | 5232.5 | 6359.5
Fit | 3.2227 | 3.6582 | 3.8324 | 3.9195 | 4.4421 | 5.1389 | 5.226 | 5.4002 | 5.6615 | 6.1841 | 6.6196 | 8.5358 | 9.581 | 9.581 | 9.581 | 10.452 | 11.7585 | 12.3682 | 14.8941 | 15.2425 | 15.678 | 18.291 | 19.162 | 40.937 | 40.937 | 76.7351 | 77.519 | 87.1 | 95.81 | 162.006 | 162.006 | 191.62 | 200.33 | 383.24 | 409.37 | 905.84 | 905.84 | 1742 | 1768.13 | 3579.81 | 3579.81 | 6924.45 | 6968

## Slide 34: ASTRID and ASTRID2 accelerators at ISA
ASTRID2
The measurement campaign was performed at ISA, Aarhus, Denmark, around the ASTRID and ASTRID2 storage rings.
ASTRID is a 40 m long electron storage ring that has operated since 1990. It operates as a 580 MeV electron storage ring, injected from a 100 MeV racetrack microtron.
ASTRID2 is a 45.7 m circumference storage ring located next to ASTRID that operates at 580 MeV.
ASTRID is used as an injector for ASTRID2 (one injection every 150-200 s, slow extraction mode), which maintains an almost constant circulating current fluctuating around 200 mA.
In ASTRID, a fast injection mode can also be operated, with an injection frequency from the microtron of 2 Hz.
Measurement of neutron H*(10) rate in 7 selected positions around the storage rings.

## Slide 35: Result of the measurements
ASTRID2
Position | Extraction | H*(10) per burst [nSv] |  |  |  | Average H*(10) rate [µSv/h]
1a | Slow | 5.6 |  |  |  | 0.37 ± 0.11
1b | Slow | 31.7 |  |  |  | 1.6 ± 0.2
2 | Slow | 1.9 |  |  |  | 0.26 ± 0.08
3 | Slow | 2.8 |  |  |  | 0.08 ± 0.01
3 | Fast | n.a. |  |  |  | 2.5 ± 0.8
4 | Slow | n.a. |  |  |  | 0.005 ± 0.001
4 | Fast | n.a. |  |  |  | 0.15 ± 0.05
5a | Slow | 27.7 | 23.3 | 23.5 | 43.7 | 2.4 ± 0.3
5a | Fast | n.a. |  |  |  | 405 ± 63
5b | Fast | n.a. |  |  |  | 41.5 ± 10.1
Position 1b Slow Injection
Position 5a Slow extraction
Slow extraction 5 – 10 mHz
Fast extraction 2Hz
The dose per burst is not measured for fast extraction mode (n.a.) because the dose is sampled every second

## Slide 36: Open question
Can the LUPIN be of any interest for measuring stray neutron around medical facilities? i.e. activities related to the WG9?

## Slide 37: Space charge effect
In case of intense burst the charge generated inside the proportional counter can shield the electric field causing a decrease in multiplication factor with a consequent underestimation of the charge

## Slide 38: Space charge effect

## Slide 39: Space charge effect compensation

## Slide 40: Space charge effect compensation