## Slide 1: Synchrotron radiation X-ray experiments for a pulse-counting type SOI pixel for the soft X-ray measurements
KEK IMSS R. Hashimoto, S. Kishimoto, R. Kumai, N. Igarashi
KEK IPNS Y. Arai, I. Kurachi, T. Miyoshi
SOKENDAI R. Nishimura

## Slide 2: Introduction
Development of the SOI pixel detector for the X-ray structural analysis.
・ Small angle X-ray scattering
experiment for functional
thin films.
・ X-ray diffraction experiment
for functional materials such
as ferroelectrics.
etc…
Area detector using SOI
+P
-P
Crystal, Thin film
Precise analysis of position for hydrogen atoms
Time response of the domain structure by pulsed external electric field
X rays
Complementary use of
hard and soft X-ray on
X-ray diffraction.
Current status
・ Conducting the structural analysis experiment by use of a fine
pixelated detector, SOPHIAS.
・ Development of the pulse-counting type pixel detector.

## Slide 3: Introduction
SOI detector
Double-SOI wafer (p-type substrate)
Back plane : negative bias
Another Si layer (middle Si) is in BOX layer
Negative
Ground

## Slide 4: Introduction
X-ray test for Test-Element-Group (TEG) chip including
pulse-counting type pixel TEG.
In this talk, the experimental results of the evaluation for the counting
type pixel TEG will be shown.
Requests for detectors from experiments
・high spatial resolution
→ pixel size is less than 50 μm square
・high sensitivity
→ detecting low energy X-ray ～2 keV
・high frame rate
→ more than 1 kHz

## Slide 5: Evaluation of counting type pixel in KEK
CPIXPTEG1
CPIXTEG3b
(collaborated work with IHEP)
CNPIX1
・ 52 μm pitch hexagonal pixel
・ 256 ×320 array + TEGs
・ 64 μm square pixel
・ 32 × 32 array
consists of 3 types
of pixel + TEGs
・ 50 μm square pixel
・ 64 × 64 array
6 mm
6 mm
18.4 mm
12.2 mm
6 mm
6 mm

## Slide 6: Evaluation of counting type pixel in KEK
CPIXTEG3b
(collaborated work with IHEP)
CNPIX1
・ 52 μm pitch hexagonal pixel
・ 256 ×320 array + TEGs
・ 50 μm square pixel
・ 64 × 64 array
6 mm
18.4 mm
12.2 mm
6 mm
・ 64 μm square pixel
・ 32 × 32 array
consists of 3 types
of pixel + TEGs
6 mm
6 mm
CPIXPTEG1 → CPIXPTEG2
study of low energy
X-ray detection

## Slide 7: Chip overview
・ 64 μm square pixel
・ 32 × 32 array consists of 3 types
of pixel + TEGs
・ 16-bit counter
・ Double-SOI wafer
CPIXPTEG2
6 mm
6 mm

## Slide 8: Chip overview
CPIXPTEG1 → CPIXPTEG2
・ Increasing contact via of the middle-SOI layer
Only 1 Middle-SOI contact
Counter
Register
Preamp.
Shaper amp.
Discrim.
Pixel layout of CPIXPTEG1

## Slide 9: Chip overview
CPIXPTEG1 → CPIXPTEG2
・ Increasing contact via of the middle-SOI layer
Only 1 Middle-SOI contact
Increasing contacts
Counter
Register
Preamp.
Shaper amp.
Discrim.
Pixel layout of CPIXPTEG1
CPIXPTEG2

## Slide 10: Chip overview
CPIXPTEG1 → CPIXPTEG2
・ Increasing contact via of the middle-SOI layer
・ Study of the backside process for the soft X-ray detection
⇒ Wafer thinning process
Total thickness 75 μm
Only 1 Middle-SOI contact
Increasing contacts
Counter
Register
Preamp.
Shaper amp.
Discrim.
Pixel layout of CPIXPTEG1
CPIXPTEG2

## Slide 11: Schematic of the pixel circuit
・ 4-bit DAC for adjusting
threshold level.
・ Hysteresis comparator.
Counter (16-bit)
Test input
Preamplifier
Shaper amplifier
Discriminator
Gain : 45 μV/e-
@designed value.
Gain : twice
@designed value.
Total gain of preamp. and shaper amp. : 90 μV/e-
→ CPIXPTEG1 evaluation result 33.5 μV/e-

## Slide 12: Hysteresis width of discriminator
Vhyst
Vth1
Vth2
Signal from shaper amplifier
Discriminator output
Discriminator behavior
Against malfunction due to the noise signal.
Designed value of Vhyst = 80 mV.
→ Cannot measure the lower level signal than Vhyst.
Shaper signal level

## Slide 13: I-V curve measurement
CPIXPTEG2 (thinning)
CPIXPTEG2 (normal)
CPIXPTEG1 (normal)

## Slide 14: Experimental setup @KEK/PF BL-14A
Slit
Ionization chamber
Filter holder
Laser/pin-hole
exchange stage
CPIXPTEG2
(backside illumination)
X-ray

## Slide 15: Measurement of X-ray profile
Exposing 6 keV X-rays with 0.8 mmφ in diameter with Vdet = -10 V.
Column
Row
Counts
In measurement using X-ray micro
beam, we missed the position of
X-ray.
15 pixel
15 pixel
Obtained shaper output signal
-200 mV

## Slide 16: Study in test bench
Low gain problem
Test input capacity : 4.2 fF
Input signal : 62 mV
Output : 10 mV
→ Corresponding gain is 6 μV/e-
Preamplifier output signal
Shaper amplifier output signal
Gain was half of designed value.
Total gain is 5 times lower
than 33.5 6 μV/e- of CPIXPTEG1.

## Slide 17: Study in test bench
Increasing the gain by adjusting bias current of Preamplifier.
Required setting of bias current for
core amp Icore and feedback amp Ifb.
Icore = 2 μA , Ifb = 2 μA
(Icore, Ifb) = (4, 0.5)
(2, 2)
(2, 0.5)
(4, 0.5)
Required Icore = 2μA
Ifb = 8 μA
*Shaper bias current was fixed.
(2, 2)
(2, 0.5)

## Slide 18: Study in test bench
Increasing the gain by adjusting bias current of a Shaper amplifier.
(Icore, Ifb)
(2, 8) @required values
(2, 4)
(2, 2)
(4, 2)
(4, 1)
(4, 0.5)
Icore and Ifb of preamplifier were 4 μA and 0.5 μA, respectively.

## Slide 19: Study in test bench
Increasing the gain by adjusting bias current of amps.
Initial
Adjusted
Gain became almost 10 times
larger than initial setting.

## Slide 20: Summary
We have started the study of counting type pixel for soft X-ray.
Back side process aiming the soft X-ray counting was performed.
The evaluation experiment by use of 6 keV X-ray was done but we
could not measure.
→ Shortage of the gain was found out so that we recovered that by
adjusting the bias current of amps.
After checking the discriminator and counter action and other tools
used for measurement, we try again the 6 keV X-ray experiment.
・ counting micro spot X-ray beam.
・ 1- and 2-dimentinal pixel scan, etc…
Future plan
For the soft X-ray counting
Using SOI-APD for the counting type pixel array.
Avalanche mechanism is helpful for soft X-ray measurements.

## Slide 21
Thank you for your attention !