## Slide 1: Velopix, a new hybrid pixel readout ASIC for the LHCB VELO LS2 upgrade
CERN, NIKHEF and USC

## Slide 2: Outline
Introduction to the VELO upgrade (Martin)
The VeloPix chip (Martin):
  Velopix FE (Rafa)
  Digital pixel (+eCDRPLL) (Xavi)
  Super Pixel, EOC, periphery and GWT (Tuomas)
Measurements (Jan)
Conclusion (Jan)

## Slide 3: M. Van beuzekom

## Slide 4: Introduction to LHCb and its upgrade
LHCb is a dedicated experiment searching for new physics by studying CP violation and rare decays of b and c quarks
Forward angle spectrometer with excellent vertex resolution and PID
Detector upgrade in 20192020 for 5x higher luminosity and improved efficiency
No more hardware trigger  all data data read out and sent to CPU farm

## Slide 5: Vertex Locator Upgrade
bla
LHCb will be upgraded in 20192020
 very tight planning
Vertex detector surrounding collision region
  In vacuum
  Close to the beam: 5.1 mm
From silicon strips to pixels
New R/O chip VeloPix, derived from Timepix3
In total 624 ASICs, ~41 Mpixels
Trigger-less readout (~2.9 Tbits/s)

## Slide 6: VELO module
One module = 4 sensors with 3 VeloPix ASICs each
256x256 pixels/ASIC, 55 x 55 µm2 pixels
ASICs thinned to 200 µm
All silicon module, minimal CTE mismatch
Si substrate with m-channel evaporative CO2 cooling
  sensors < -25 °C
Sensor signal at end-of-life ~7000 e-

## Slide 7: ASIC challenges: data rate & radiation hardness
Sensor and ASIC exposed to high, non-homogeneous, radiation fluence
  Only part of the pixel matrix gets the full dose of ~ 370 Mrad
  TID at periphery of chip is factor 10 lower
For data rates calc. we assume collisions in every LHCb bunch crossing
  in reality only 2/3 of bunches collide
  would require a lot of memory to level out
   assume peak rates for ASIC design
Data flow simulations using physics
Monte Carlo data
No trigger, all data sent off chip
Hottest ASIC gives ~15 Gbps
ASIC design starting point: Timepix3
Data rates per ASIC (Gbps)

## Slide 8: From Timepix3  Velopix
 | Timepix3 (2013) | VeloPix (2016)
Pixel arrangement | 256 x 256 | 
Pixel size | 55 x 55 µm² | 
Peak hit rate | 80 Mhits/s/ASIC | 800 Mhits/s/ASIC 50 khits/s/pixel
Readout type | Continuous, trigger-less, TOT | Continuous, trigger-less, binary
Timing resolution/range | 1.5625 ns, 18 bits | 25 ns, 9 bits
Total Power consumption | <1.5 W | < 3 W
Radiation hardness |  | 400 Mrad, SEU tolerant
Sensor type | Various, e- and h+ collection | Planar silicon, e- collection
Max. data rate | 5.12 Gbps | 20.48 Gbps
Technology | IBM 130 nm CMOS | TSMC 130 nm CMOS

## Slide 9: VeloPix Project Overview
VeloPix wafers
Design started in June 2013 (after Timepix3 submission)
Change of technology (IBM 130nm → TSMC 130nm)
The chip was submitted as an engineering run on May 2016
First wafers received on 31st August
Fabricated (and diced) chips back at CERN on 7th September
First beam tests done on 6th November
Production testing later this year (624 chips)
Irradiation campaign in the future with sensors bonded

## Slide 10: VeloPix Chip Architecture
Pixel matrix:
  256 x 256 pixels
  128 x 64 super pixels (2x4 pixels each)
  @40MHz
Packet-based architecture:
  8 pixels/packet + 9 bit time stamp  30% reduction in data rate
Data-driven readout:
  20 Mpackets/s/double column
40, 80, 160 and 320 MHz TMR clock domains in the periphery
1 to 4 configurable serializers (GWT)
Similar to the GBT frame

## Slide 11: VeloPix pixel schematic
1 pixel
Global threshold
(LSB= ~15e-)
Front-end (Analog)
Leakage
Current
compensation
Preamp
Input
pad
TpA
TpB
TestBit
4-bit Local
Threshold
4fF
5fF
Common for 8 pixels
Clock
Manager
24-bits
Super pixel (Digital)
Data Node
24-bits
24-bits
10-bits
FIFO (2x18)
8
Addr (6-bit)
TOA (10-bit)
Hitmap (8-bit)
BXID
To the next
SuperPixel
Synchronous handshake
Synchronous handshake
From previous
SuperPixel
6-bits
Synchronizer
&
Clock gating
Front-end (Digital)
Shutter
Clock
(40/80 MHz)
6-bit LFSR:
Counter
Shift Register
TPulse
6-bit TMR
Pixel memory
6-bits
6-bits
Valid
Event
DataIO
MaskBit
Rafa
Xavi
Tuomas

## Slide 12: R.Ballabriga

## Slide 13: Specifications
Parameter | Value
Technology | TSMC 130nm CMOS
Number of pixels | 256 x 256 @ 55 mm pitch
Analog front-end size | 55 x 14.5 mm
Analog supply voltage | 1.2 V
Detector capacitance | 50 fF
Detector polarity | Optimized for electron collection
Maximum of charge distribution | 16ke- (200mm Si @ 80e-h+ pairs per mm)
Leakage current | Up to 12nA/pixel (non-uniform)
Minimum threshold | ~500e-
Pile-up | Discharge of 16ke- charge in 300ns
Time of arrival resolution | 25ns
Operating temperature (in experiment) | -30oC<T<0oC
Analog power consumption | ~6.1mW/pixel 200mW/cm2 (250mW/cm2) (constant)

## Slide 14
Limited NMOS leakage current (in good Fab)
ELTs do not move!
minL
minL, W>1 (or LVT)
L>min
Ion degradation<5%
FAB 6 (NMOS)
Radiation measurements: S. Michelis, F. Faccio
FAB 14 (NMOS)
The annealing at -30oC is much faster in HVT transistors wrt regular transistors
2 Fabs showed different results in terms of leakage for the NMOS transistors
We decided to design for the worst case i.e.:
NMOS transistors either ELT or HVT

## Slide 15: The front-end circuit
Cfb
Vth
Vout_disc
4-bit offset calibration
-A
-HV
FB

## Slide 16
-A
vdda
vdda
vdda
vdda
Cfb=5fF
-HV
Ileak
VOUT
Cleak=700fF
ELT
ELT
ELT
ELT
ELT
Vbias_Ikrum
x2
Vfbk
Bias block shared between 2 adjacent pixels
ELT
vdda
Vfbk
vssa
vssa_ik
vssa_ik
vssa_ik
Vx
The inner terminals of the ELTs are connected to Vx (to minimize Cdb (current source) and Csb (input pair)
Dummy (matching and decoupling)
VBias_Ikhalf
VIN
Segmented in 4 in the direction of the length

## Slide 17
-A
vdda
vdda
vdda
vdda
Cfb=5fF
Ileak
VOUT
Cleak=700fF
HVT_ELT
HVT_ELT
ELT
HVT_ELT
ELT
Vbias_Ikrum
Vfbk
(500mV)
REG_ELT
vdda
Vin_casc
vssa
vssa_ik
vssa_ik
vssa
Vx
VBias_Ikhalf
VIN
VIN
VOUT
vdda
Vbias_preamp
Vbias_preamp_casc
25nA
830nA
2x830nA
x1 x2
x1 x2
2.5mA
1.2V

## Slide 18: The choice on the capacitance for CLEAK
Vx
Vx
(Vx)
CLEAK implemented as a PMOS LVT
For Power Supply Rejection Ratio
For minimization of its gate leakage current!

## Slide 19: The comparator
Designed for minimizing time-walk of the system (target below 25ns)
Designed based on cascading low gain stages was selected
Constant current consumption in analog part

## Slide 20
VIN
VCAS
VCAS
VTHR
VOUT_s1-
VOUT_s1+
IEQ_BAR
IEQ
1.2mA
1.2V
ELT, reg
ELT, reg
VOUT_s2-
VOUT_s2+
1.2mA
1.2V
ELT, reg
VOUT_s1-
VOUT_s1+
hvt
hvt
Gain ~4, BW ~40MHz
Gain ~3.2, BW ~40MHz

## Slide 21
VOUT_s2-
VOUT_s2+
Inverter_B_XL
Analog Digital
ELT, hvt
hvt
Gain ~40
Analog_Disc_Out
Digital_Disc_Out

## Slide 22: Simulation ToA, ToT
Threshold at 750e-, reference pixel (bottom column).
Corners: typ,ff,ss, radiation corners (at 100Mrad and 400Mrad)

## Slide 23: Implementation of the DAC
VbiasDAC_local
VbiasDACcasc_local
55nA
Bn<0>
55nA
Bn<1>
B<1>
x1 x1
…
I_DACOutL
I_DACOutR
x2
x2
110nA
VbiasDACcasc
VbiasDAC
ELT
VbiasDAC and VbiasDACcasc are provided from the periphery. / Less loading to the DAC in the periphery.

## Slide 24
14um
18um
2.5um

## Slide 25
Test capacitor
DACs
Comparators
Test pulse switches
Feedback capacitor
Leakage current compensation capacitor
Input transistors
IKRUM current sources
Preamplifier
Target transistor DAC
Blocks “double pixel”
NMOS transistors on DNW
Ikrum current sources
Shielding
Most contacts are double
Left/right most columns different schematic and layout
Dimensions 29um x 55um

## Slide 26: X. llopart

## Slide 27: eCDRPLL (simplified) Block Diagram
refClk40MHz
analog
data
control
clock
2.56 Gb/s
Phase / Frequency Detector
Loop filter
Divider
(8 / 6 / 4 / 2)
Phase Shifter
Wien-bridge
BIST
VCO
feedbackClocks
vControl
vcoClock
clock240_320MHz
resampledData
extControl[x:0]
State Machine
smCtrl[x:0]
smCtrl[x:x]
smCtrl[x:x]
smCtrl[x:x]
smCtrl[x:x]
smCtrl[x:x]
smCtrl[x:x]
clock120_160MHz
clock60_80MHz
clock40MHz
recoveredData
vcoClockPhases[15:0]
http://cern.ch/proj-gbt/eCDR-PLL
P. Leitao

## Slide 28: ePLLCDR Description
Radiation-hard and phase deterministic PLL & CDR
  TSMC 130nm, 960 x 420 µm
Two Frequency Multiplier modes
  VCO@240MHz - 40, 60, 120 and 240 MHz simultaneous output clocks
    260-ps phase shift resolution
    Targets GBT-FPGA/backend applications
  VCO@320 MHz – 40, 80, 160 and 320 MHz simultaneous output clocks
    195-ps phase shift resolution
    Targets typical TTC/e-link applications
  BIST for jitter characterization (to be run at production time)
Clock and Data Recovery
  VCO@320 MHz only
    40, 80, 160 and 320 MHz simultaneous output clocks
    40, 80, 160 and 320 Mbps selectable output recovered data
    Uses reference-less wien-bridge VCO calibration (no external clock required)
    195-ps phase shift resolution (data pushed-out on clock’s falling edge)
http://cern.ch/proj-gbt/eCDR-PLL
P. Leitao

## Slide 29: ePLLCDR preliminary experimental data
Phase Shifter
  INL and DNL < 0.5 LSB
PLL full characterization on going
CDR and BIST are still being characterized
http://cern.ch/proj-gbt/eCDR-PLL
P. Leitao
Jitter measured using the phase noise (Agilent E5052B)
  Input Clock:
    40 MHz (Agilent 81133A)
  Output Clock:
    320 MHz
    PLL w/ Default parameters

## Slide 30: CERN TSMC 130nm HD Library
Library type | Tap-less CMOS standard cell library
Transistors | Pmos+NmosHVT
Antenna cell | 1
WellSubTap cell | 1
Filler cells | 8
DeCap cells | 5 (thin gate) 4 (thin gate + diode) 4 (thick gate)
Sequential cells | 2
ClockGate cells | 1
ClockBuffer cells | 3
Combinatorial cells | 54
Total cells | 83
Technology | TSMC 130nm 1p7m4x1z1u
Metals used | M1 to M2
Corners [prerad, 100M and 400MRad] | 1.2 25C TT 1.1 80C SS 1.3 -20C FF

## Slide 31: VOTERI_B_XL_TSMC_HVTN
2.46 µm
3.68 µm

## Slide 32: HD Library Characterization
Use Liberate
9 different corners:
  TT 25C 1.2
    preRad
    100MRad
    400MRad
  SS 80C 1.1
    preRad
    100MRad
    400MRad
  FF -20C 1.3
    preRad
    100MRad
    400MRad
Observation:
  400MRad_TT  preRad_SS
Corners used:
  SLOW: SS 80C 1.1 400MRad
  TYP: TT 25C 1.2 PreRad
  FAST: FF -20C 1.3 PreRad
100MRad_tt
preRad_tt
preRad_tt
preRad_tt
preRad_tt
preRad_ff
400MRad_tt
preRad_ss

## Slide 33: Velopix Pixel
@Shutter | SelectOverflow_FullRange_TOToverflow | TOToverflow [1:0] | Select_1hitTOT_iTOT | Select_PC_TOT | @SelectDirection_Write_Read | read_seu | @Parallel Load |  | 
0 | X | XX | X | X | 0 | 0 | 0 | Read | Reads the 6-bit Shift Register
 |  |  |  |  |  |  | 1 |  | Reads back the 6-bit PCR
 |  |  |  |  |  | 1 | 1 |  | Reads the 6-bit SEU status from the PCR
 |  |  |  |  | 1 | X | X | Write | Writes the 6-bit PCR
1 | 0 | 00 | 0 | 0 | X | X | X | Data acquisition | Valid event generated if TOTmeasured ≥ 0
 |  | 01 |  |  |  |  |  |  | Valid event generated if TOTmeasured ≥ 1
 |  | 10 |  |  |  |  |  |  | Valid event generated if TOTmeasured ≥ 2
 |  | 11 |  |  |  |  |  |  | Valid event generated if TOTmeasured ≥ 3
 | 1 | XX | 0 | 0 |  |  |  | Data debug acquisition | Integrating TOT up to 63 counts
 |  |  |  | 1 |  |  |  |  | Integrating PC up to 63 counts
 |  |  | 1 | 0 |  |  |  |  | 1-hit integrating TOT up to 63 counts
 |  |  |  | 1 |  |  |  |  | 1-hit integrating PC up to 63 counts

## Slide 34: TOT Processor
Tpeak > 25ns
Tw#1< 25ns
Tw#2> 25ns
Event-by-event processor
Very useful to correctly set a hit in BxID
Events below TOT threshold are ignored
Most noise events can also be filtered
TOT threshold | TOT clocks | Fix delay to valid event
00 | 0 | 50ns
01 | 1 | 75ns
10 | 2 | 100ns
11 | 3 | 125ns
Qin=700e-
Off_time
On_time

## Slide 35: Simulation of TOT processor
Peak Time | 25ns
Return to zero | 12.5ns/Ke-
ENC | 60e-rms
Threshold mismatch | 30e-rms
Gain | 35e-/mV
Analog Threshold | ~570e-
Pixels tested | 64
TestPulses per Qin/pixel | 100
Chart: Qin [e]
Threshold | 0 | 1000 | 2000 | 3000 | 4000 | 5000 | 6000 | 7000
Chart: Qin [e]
Threshold | 0 | 1000 | 2000 | 3000 | 4000 | 5000 | 6000 | 7000
Best setting: TOT=01  100% Qindetected ≥ 1.0 Ke-

## Slide 36: T. Poikela

## Slide 37: Double column datapath
FIFO
Node
Node
FIFO
FIFO
Node
Node
FIFO
64 FIFOs
0 1
Register
Arbiter
Previous
Node
FIFO
Data
Size = 2
packets
Up to 20
Mpackets/s
Next
Node
8 digital
Front-ends
8 analog
Front-ends
Sensor
Super pixel
column
Analog pixel
column
Digital pixel
column
Super pixel packet:
hitmap 8b
Addr 6b
ToA 10b
Timestamping @ 40 MHz

## Slide 38: Periphery datapath
64 EoCs
64 EoCs
Left
Data Fabric
Right
Data Fabric
Center
Node
Pixel Matrix
FIFO
19.20
Gbps
Data Node
Router/
Scramblers/
Framing
4.8 Gbps per bus
4 x 5.12 Gbps
GWT link
Data
out
160 Mpackets/s/node
64x up to 20 Mpackets/s
64x up to 20 Mpackets/s

## Slide 39: EoC Data Fabric
Center
Node
64 EoC Blocks (Periphery left side)
0,0
0,15
1,15
2,15
3,15
1,0
2,0
3,0
0,1
1,1
2,1
3,1
data[29:0]
handshake
L0
R0
L1
R1
L2
R2
L3
R3
0
1
2
3
Packet
Router
Periphery
Right side
20.48
Gbps
Channel:
30b @ 160 MHz

## Slide 40: Packet processors (Center node)
Packet
Proc L0
Packet
Proc L1
Packet
Proc L2
Packet
Proc L3
Max latency
Max hits
Time stamp
@ 40 MHz
Filtered
packets
Filtered
packets
Filtered
packets
Filtered
packets
Counter
SP
BX ID
T0
T1
Latency: T1 - T0
10
32-b counter
EoC
data
data
data
data
empty/
read
empty/
read
empty/
read
empty/
read

## Slide 41: EoC Data Fabric Node
FIFO
2x 30 bits
FSM
write/
full
1
0
empty/read
30
Previous Node
EoC Block
Next Node
empty/read
empty/
read
30
30
get_max
count()
Arbiter
4-bit counter
Double column
empty/read
data

## Slide 42
Input 0
VOQ 0,0
VOQ 0,1
VOQ 0,2
VOQ 0,3
Input 1
VOQ 1,0
VOQ 1,1
VOQ 1,2
VOQ 1,3
Input 2
VOQ 2,0
VOQ 2,1
VOQ 2,2
VOQ 2,3
Input 3
VOQ 3,0
VOQ 3,1
VOQ 3,2
VOQ 3,3
Output0
Output
1
Output
2
Output
3
Input ID
0
1
2
3
0
1
2
3
0
1
2
3
0
1
2
3
0
1
2
3
Output ID
Round robin
arbiter
Round robin
arbiter
Round robin
arbiter
Round robin
arbiter
Round robin
arbiter
Round robin
arbiter
Round robin
arbiter
Round robin
arbiter
30b
30b
30b
Packet Router

## Slide 43: Router and output block
Scrambler
Scrambler
Scrambler
Scrambler
Framing
Framing
Framing
Framing
Packet
Router
(4x4
crossbar)
Serializer
Serializer
Serializer
Serializer
f2r[0]
f2r[1]
f2r[2]
f2r[3]
16-to-1b
ser.
160
MHz
160
MHz
40
MHz
160
MHz
4x 30b
4x 30b
4x 128b
320 MHz
DDR
128b registers:
parity + header
Channel
Mask [3:0]
EoC Data Fabric
Output
drivers
Fixed latency

## Slide 44: GWT (NIKHEF)
dc: 13mA ,
ripples: <2mV p-p
MUX
16 phases
Reg
<0:7>
data
8bit @ 320MHz
Reg <8:15>
data
8bit @ 320MHz
Driver
negedge
posedge
CMOS
Multi-phase DLL
Edge-
combiner
ph_0
ph_1
ph_2
ph_3
ph_14
ph_15
VDD_HS
GND_HS
GND_CORE
VDD_CORE
serialized data @ 5.12 Gbps
5Ω
5Ω
3Ω
3Ω
534pF
534pF
dc: 7mA ,
ripples: <2mV p-p
Low power 5.12 Gbps byte-interleaved serializer and wireline transmitter
<30mW/channel
320MHz

## Slide 45: Clock distribution
320 MHz
(TFC)
40 MHz
(system)
PLL
ClkGen
TMR:
40, 160, 320
80 MHz
(slow control)
clk_en
Data domain
Control & config domain
clk_en
bypass
PLL
Registers
Datapath
40
160
320
40
160
320
ECS Master
Master 40
Data in
(sensor)
Data out
(GWT)
Data in/out

## Slide 46: Velopix: TMR/RadHard design
Pixel Matrix (<400Mrad):
  Only NMOS ELT transistors used in analog front-end
  Full TMR in FSM & configuration FFs:
    Pixel configuration  Asynchronous self correcting TMR latch
  Pixel data flip-flops (FF) unprotected
  Custom made HD Standard cell library (74 cells) with HVT NMOS HVT and regular PMOS:
    Standard cell library characterized at 400 Mrad, High-Vt NMOS + Std PMOS
Periphery (<50Mrad):
  Full TMR in FSM & configuration FFs
  Data path unprotected

## Slide 47: VeloPix
14.14 mm
Analog front-end
16 x 55 um²
Double column:
  512 pixels
  64 super pixels
Full matrix:
  128 Double columns
  ~190 Mtransistors
  14.8nF digital decoupling (thick gate)
Pixel & Superpixel Logic (HD Library)
16.6 mm
Active Periphery:
  40, 80, 160 and 320 TMR clocks
  HVT TSMC (tcb013ghphvt library)
  4nF digital decoupling (thin gate)
2.4 mm

## Slide 48: J. Buytaert

## Slide 49: VeloPix test readout system
SPIDR-readout system developed at NIKHEF (NL)
  Based on a Virtex VC707
Similar as Timepix3. -> Large reuse of soft & firmware.
Chipboard developed at Univ. Santiago De Compostella.
Not final LHCb readout.
VeloPix

## Slide 50: VeloPix: General Measurements
Measured power consumption (@nominal settings):
  Analog suuply < 480 mW
  Digital:
    Periphery < 380mW
    Pixel Matrix <350mW (idle, i.e. clock distribution only)
    @High rate ~+300mW (simulated)
  Total= ~1.5W @High rate
Slow and Fast control fully functional.
Pixel matrix Configuration and readout
On-chip biasing DACs (next slide)
Internally measured packet latency (@low rate)
eCDRPLL (CERN) total jitter @320MHz <6psrms

## Slide 51: VeloPix on-chip biasing DACs
VTHR (~15e- LSB)
Range: 15.3 ke-
TestPulse (~31.25e- LSB)
Range : > ~25 ke-
8 bit DAC’s
10 bit DAC
14 bit DAC implemented as
16 x 10bit DAC with offsets)

## Slide 52: VeloPix Serializers (GWT)
A low power (~20mW) 5.12 Gbps byte-interleaved serializer and wireline transmitter (NIKHEF).
Voltage driver, large swing (1.6 V diff)
Measure internally generated PRBS15 signal with scope  Measured BER<10-12
Eye of VeloPix PRBS15 stream

## Slide 53: FE Measurements
Front end characterization measurements done through the slow control readout.

## Slide 54: Front-end gain: (test pulses in one pixel)
Chart: DACCode
Chart: DACCode
| 5668 | 5732.2379999999976 | 5799.5339999999997 | 5884.4070000000002 | 5954.232 | 6036.0959999999995
 | 0 | 1.0625 | 2.1562 | 3.4688 | 4.5 | 5.6562
0.0154 e-/DACLSB* 1 DACLSB = 0.38mV  24.6 mV/ke-
*Assuming Ctest=5fF
Test pulse amplitude in e-(*)
Global threshold level

## Slide 55: Front-end gain uniformity (test pulses in 32 pixels)
Chart: Energy [Ke-]
| 0 | 1.0625 | 2.15625 | 3.46875 | 4.5 | 5.65625
0 | 5648 | 5712.533 | 5780.516 | 5865.814 | 5936.481 | 6018.116
8 | 5674.5 | 5740.032 | 5810.758 | 5898.596 | 5971.845 | 6055.853
16 | 5701 | 5771.428 | 5845.634 | 5938.731 | 6015.09 | 6102.838
24 | 5661 | 5726.605 | 5794.835 | 5879.784 | 5950.423 | 6033.452
32 | 5639.5 | 5706.332 | 5775.894 | 5863.154 | 5935.488 | 6021.32
40 | 5674.5 | 5742.44 | 5812.915 | 5902.869 | 5974.651 | 6060.887
48 | 5689 | 5753.827 | 5821.018 | 5905.595 | 5976.227 | 6057.695
56 | 5668 | 5732.238 | 5799.534 | 5884.407 | 5954.232 | 6036.096
64 | 5668 | 5733.203 | 5801.787 | 5888.526 | 5960.176 | 6043.359
72 | 5675 | 5738.718 | 5806.615 | 5891.527 | 5961.814 | 6044.79
80 | 5655.5 | 5719.201 | 5787.559 | 5871.267 | 5943.016 | 6023.63
88 | 5647 | 5710.955 | 5778.839 | 5865.06 | 5935.788 | 6018.678
64.72 DACLSB/ke-  1 DACLSB = 0.38mV 
24.6 mV/ke- (0.81 mV/ke- rms)  ~3.3% gain variation
~3.3% front-end gain variation

## Slide 56: Pixel ENC. [Threshold scan over noise floor in PC mode]
8 pixels
Typically ~60 enc

## Slide 57: ENC of full matrix.
Chart: Threshold DACCode
µeq= 4.14 (62.9 e-)*
σeq= 0.27 (4.1 e-rms)*
*@gain of 25mV/ke-
Measured ENC of all pixels
Threshold scan over noise floor
All pixels at code 0xF
No trend visible

## Slide 58: Threshold Equalization.
Chart: DAC Code
µ0x0= 1378
σ0x0= 27.06
µ0xF= 1513
σ0xF= 26.19
µeq= 1446
σeq= 2.65 (40.3 e-rms)*
Unequalised Threshold distribution:
All pixels at code 0x0
No trend visible.
*@gain of 25mV/ke-

## Slide 59: Summary of pixel measurements
Pixel gain | ~24.6 mV/Ke-
Pixel to pixel gain variation | ~3.3%
Pixel ENC | 62.9 e-
Pixel to pixel threshold mismatch | 410 e-rms
Pixel to pixel threshold mismatch calibrated (Threq) | 40.3 e-rms
Expected minimum threshold | > 450 e-
Threshold equalization only calculated not measured on chip
All measurements assuming Ctest=5fF

## Slide 60: Bumps on a Velopix chip
SEM images done by S. Vähänen (Advacam)

## Slide 61: Velopix Modules
Bumping to the 200 µm n-p Hamamatsu Si sensor
8 3x1 modules and 6 single sensors available
Not thinned  Experiment requires thinning to ~200 µm

## Slide 62: First source measurements [S40 single sensor]
Fe55 600s Thr ~900 e-
Image taken through the ECS (Slow Control) in photon counting mode

## Slide 63: Beamtest
First VeloPix in testbeam last Sunday
Image taken through 1 GWT link
Hitmap from 1 SPS spill (~5 sec)
Received 2.6 Million Superpixel packets with 3.42 Million pixel hits:
  Numbers consistent with #tracks recorded by the Timepix3 telescope (upstream of VeloPix)
  Linking of VeloPix hits to telescope tracks to be done
Note that this is only 0.1% of the rate capability of VeloPix !

## Slide 64: Still to do…
Obtain calibration with X-ray sources:
  Extract correct FE characterization (ENC, gain…)
Read the chip at full bandwidth with all GWT (20 Gbps):
  where to find such beam?
Whole setup test using: GBT as clock reference, optical drivers, 200m of fibers and miniDaq
TID and SEU irradiation campaigns
Production testing later this year at CERN

## Slide 65: Conclusions
VeloPix ASIC designed in 130nm CMOS for the LS2 VELO upgrade
First results show the chip is alive and eyes open:
  Power < 2 W/ASIC, DACs working, pixels functional
  Pixel: Gain ~25 mV/Ke- and ENC 63 e- (no sensor)
  GWT serializer working
All measurements, so far, indicate that the Velopix chip is fully functional as designed:
  First source images taken in photon counting mode : Slow Control
  First beam test through the binary readout chain @ low rate
  Scheduled TID and SEU campaigns to validate design robustness

## Slide 66: VeloPix Contributors
ASIC designers:
  R.Ballabriga, V.Gromov, X.Llopart, S.Miryala, T.Poikela, J.D.Schipper and W.Wong
IP Blocks:
  eCDRPLL: R. de Oliveira and P. Leitao
  Band-Gap: S. Michelis
Support, readout and Testing:
  J. Alozy, M.van Beuzekom, H.Boterenbrood, B. van der Heijden, J. Buytaert, M.Daldoss and E.Lemos